- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
The cursor details display is located in the right menu bar.
It shows details related to the current position of the cursor.
The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
The cursor-information area (see page 6) in the right menu-area uses this information.
You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
The setup I/O devices contains different areas.
#### Audio-mode selection
Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).

Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use forTrace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- Trace Layout invoke the layout – editor for the live trace. See page Error: Reference source not found for details about how to use the layout – editor.
- Hide current trace hides both the amplitude and phase display. The shortcut is ^ (Left to the number 1 key on your keyboard).
- Show Instant Values as Dots show the non averaged result of the current FFT calculation as dots. In the transfer-function measurement the result of the last calculation is shown as a dot. SATlive supports different settings for FFT and transfer-function calculation. You can change the color of the dots in the color section of the Setup → Display window. See page Error: Reference source not found for details.
- Save this entry contains a popup menu where you can save different traces.
  - Save as SATlive data saves the live trace into a file.
  - Save Trace as Picture saves the display as a picture.
  - The Save F5-Trace to File, which is enabled only if a memory trace is assigned, saves the memory trace to a file.
  - Save as Wave Data saves the impulse response as default windows® wave format. You can import this format in various other audio related software.
- Load F5-Trace from File open a saved trace as the F5 memory trace.
- Apply Window add a window to the signal prior to the calculation of the transfer-function. A window excludes later reverb and longer delays from the measurement.
- Direct Reference opens the direct-reference menu. See page Error: Reference source not found for details. This entry is available during an impulse-response measurement.
- Edit Target Trace opens the editor for the Target Trace. See page Error: Reference source not foundError: Reference source not found for details about the Target-Trace functionality.
- Allow Gain-Setup using Mouse allows you to move the currently selected trace by moving the mouse vertically while you hold the left mouse button down.
##### Additional Setting for the RTA measurement
- Calibrate current EQ calibrates the RTA display of the currently selected RTA. To do so, use a loopback from output to the reference (REF) input, select Ultra for the release time and Average for the display. Select the REF input and start the pink noise.
  After a few seconds the entry will be enabled.
##### Additional settings for the FFT measurements
- Windowing select the window applied prior to the calculation of the FFT from the sub menu.
- THD+N Start a total harmonic distortion measurement. See page Error: Reference source not found for details.
##### Additional settings for the FFT + Spectrograph measurement
- Time Range Spectrograph select the desired time for the spectrograph display from the sub menu. The time selected is needed to fill the spectrograph.
- Run only once if this option is selected, the spectro – graph display (but not the FFT - Display) will freeze when it is completely filled.
##### Additional for the measurement of the Impulse Response
- Windowing select the window applied prior to the calculation.
- Display size set the size of the calculated impulse response. Higher size has higher precision but takes more time to calculate. This value also determines the range of time that the impulse response covers. You can access the sub-menu in the right-menu area (see page Error: Reference source not found) also.
- Averages set the number of averages performed. The the Average All Values option averages all measurements since the last reset. See page Error: Reference source not found for details about averaging.
- Amplitude modification set the gain of impulse response, given as a linear factor, not as a dB value. The area Y-Factor in the right menu-area (see page Error: Reference source not found) shows the value currently selected. To access the sub menu just click on this area.
- IR Zoom Value select the zoom factor for the display of the lower trace. The area is always centered around the time value of zero.
  You can see the current zoom factor in the lower right corner of the display area. Click onto the value to open the sub menu.
- High – Cut choose the desired high – cut frequency from the sub – menu. Use this option to remove interfering high frequencies from the calculation of the impulse – response.
- Low – Cut choose the desired low – cut frequency from the sub – menu.
  Use this option to remove interfering low frequencies from the calculation of the impulse – response.
### Trace layout MAT (simple menu layout)
- To select the parameter to edit, click on the symbol left to the value edit field with the left mouse button and select the parameter from the popup menu showing up.
- Amplitude only this option is visible in the transfer function measurement only. It shows the amplitude window in full size.
- Phase only this option is visible in the transfer function measurement only.
  It shows the phase – display window in full size.
- Upper limit move the trace display's top border. The visible range ranges from upper limit minus Y-range to upper limit. This setting is not active in the transfer function measurement.
- Range Y set the trace display's vertical range.
- Line Gap set the scale's vertical distribution.
- Averages set the number of averages to perform. See page Error: Reference source not found for more information about averaging.
- Display Range for Coherence: select the range for the coherence display from this sub-menu. This setting is available only in the transfer-function measurement.
- Configuration stores or loads layout-settings from this sub-menu. See page Error: Reference source not found for details of the configuration handling.
- Trace Layout opens the Layout Editor for the live-trace. See page Error: Reference source not found for details.
- Show Last – Good Values enables the display of previously calculated values if the current values do not fit the quality – criteria. This entry is available only during the measurement of the transfer – function. See page Error: Reference source not found for details about the Last – Good Values feature
- Hints:
- After the parameter is selected the value editor field shows the parameter's current value. The current parameter's name is displayed left of the value editor field.
- The parameter's value can be changed at the value editor field.
- Every parameter features a popup menu containing some preset values.
- The values for the parameters Low DB, Range Y and Averages may differ for each measurement.
- You can select the input(s) to display below the VU meter when you perform a RTA or a FFT measurement.
- Set the layout parameter in the impulse – response measurement either in the menu Tools menu or in the right menu area.
- In measurements which use a split screen, it is possible to change the position of the split. There are two ways to do this:
- Use the popup menu of the split area to select one of the three predefined positions.
- Put the mouse onto the split area, click and hold down the left mouse-button and move the split area up or down by moving the mouse.
### Select the Input for RTA and FFT measurement
- You can select the input(s) to display in the MAT → FFT (see page Error: Reference source not found) or in the MAT → RTA (see page Error: Reference source not found) measurements.
- Either perform a mouse click on the labels below the VU meter or you use the shortcuts F8 (REF) and F9 (MIC) to toggle the display of each trace.
- Hints:
- You can display the REF input, the MIC input or both inputs.
- A fully colored background indicates that the trace assigned to this button currently shows in the display.
- The current delay will be applied to the REF input.
- The mic correction (see page Error: Reference source not found) will affect the MIC input only.
- You can change the assign of the physical input to the functions in Setup → I/O Devices (see page 9).
### Cursor details (MAT)
- The cursor details display is located in the right menu bar.
- It shows details related to the current position of the cursor.
- The following values are shown (top down):
- The frequency at the position of the cursor is shown at the top.
- The first line shows the Wave length assigned to the frequency. To show a list containing more wavelength related info place the mouse on this value and hold down the left – mouse button.
- The second line shows the Period of the frequency at the cursor's position.
- The next line contains the Phase angle of the live trace at the cursor's position (Displayed only if a transfer function measurement is performed). See page 7 for more details about the calculation of this value. The little square at the right indicates the color of the live – trace.
- The next line shows the Phase angle of the selected trace at the cursor's position (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value. The little square at the right shows the display color assigned to the selected trace used for this calculation.
- The next line indicates the difference between the phase angles of the live trace and the selected trace. (Displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace). See page 7 for more details about the calculation of this value.
- Finally the delay between the selected trace and the live trace calculated using the difference of the phase angles and the frequency of the cursors current position is shown. Which is displayed only if a transfer function measurement is performed and the live trace is not selected as the selected trace. See page 7 for more details about the calculation of this value.
#### Calculation of the Phase Values
- Due to the cyclic nature of the phase, it's range is limited to a range of 360°.
- The phase trace might be shifted by some full cycles, so the phase value might not be the value shown, but in the range between + 360° or +720° and so on. The pure phase value does not contain an information which would enable us to calculate which cycle we're looking at.
- The only way to get the 'real' phase is to look on the slope of the phase-trace. This slope contains the so called group-delay, which indicates the delay of a certain frequency.
- SATlive uses this slope to estimate the 'true' value of the phase. Furthermore you can use the group-delay to check if a polarity-inversion is present at this frequency.
- The cursor-information area (see page 6) in the right menu-area uses this information.
- You can use the setting 'Time using unwrapped phase' in the Setup → Post Process window (see page Error: Reference source not found) to select which value to show.
##### Display when the entry Time using unwrapped phase is checked
- Both phase values are shown as estimated 'true' values. In this case values above 180° and below -180° are possible. If SATlive detects a constant phase-shift, then INV is shown right to the phase value.
- If a constant phase shift is detected for one (and only one) of the two traces, then the value for the difference of the phase values is reduced by 180° to remove the effect of the constant phase shift, which is indicated by the string INV right to the value. The difference in time is also adjusted. In this case you need to apply the time-value and you need to invert the polarity of one of the two signals checked.
##### Display when the entry Time using unwrapped phase is not checked
- Both phase values are clipped to the range from -180° to +180°. If SATlive detects a constant phase – shift, then INV is shown right to the phase value.
- The value of the difference in phase and the value of the difference in time calculated from this difference phase will not be corrected.
### Bottom Menu Bar MAT
- The bottom menu bar is located at the lower right area. It is available in the MAT Module. Not all options are available for all kind of measurements.
- Range open the sub menu Range. See page Error: Reference source not found for details.
- Cursor on Peak moves the cursor to the peak of the live trace. This action unlocks the Lock on Peak setting.
- Lock on Peak if this option is activated then the cursor is moved to the live trace's peak each time the live trace is drawn. To unlock the cursor you can either click on this button again or you can perform another cursor related action like:
  - Click on the trace's display to set the cursor.
  - Move the cursor using your keyboard.
  - Invoke the cursor on peak action.
- Weighting Filter opens a sub menu, where you can choose a weighting filter for the mic input. See the weighting filter section on page Error: Reference source not found of this document for further details.
## Setup
### Setup → I/O devices
- The setup I/O devices contains different areas.
#### Audio-mode selection
- Select the kind of audio driver you want to use from the selection in the upper area.
- Windows wave driver this are the generic drivers supplied by windows® OS for the access to the audio hardware. See page 11 for details.
- Directsound driver this are windows directsound® drivers for the audio I/0. Directsound® is part of DirectX®. See page 12 for details.
- ASIO driver: Use the ASIO interface to access audio device.
  If you select the ASIO driver you have to select the device you want to use before entering the setup. See page Error: Reference source not found for details. In the ASIO mode you can use the Direct-Reference Tool (see page Error: Reference source not found).
- Hints:
- To edit the audio I/O parameters click on the button Edit Settings.
- Use the SAVE button in top menu bar to store your selection.
- ASIO is a trademark and software of Steinberg Media Technologies GmbH.
- Because SATlive can only check drivers and not devices, an ASIO device might appear in the list even if it is currently not connected. The selection of a device that is not present will result in empty fields in the ASIO setup or in an error message.
- Use the shortcut Ctrl + Alt + D to show information about the current driver.
#### Audio debug section
- Number of Buffers to ignore the MLS measurement used in the impulse – response measurement needs some time to reach a stable state. Therefore at least one MLS – sequence needs to be played prior to the measurement. In some cases one sequence might be not enough, in this case you can increase this value if needed.
- Settling time when SATlive starts pink noise or white noise a fade-in of 500ms is performed. This time must also be excluded from the measurement to get stable results. Some drivers need more time to establish a stable audio performance.
- Flip Driver Start flip the start sequence of the audio I/O drivers. This might help some drivers to shorten their settling time.
- Input routing assign the physical inputs to the reference (REF) and measurement (MIC) function. With an ASIO driver use the settings of the ConnectionManager, described on page Error: Reference source not found, to assign a certain input to a function.
- Flip Meter's Assign flips the position of the VU – meter's bars, so that the left meter shows the measurement (MIC) level and the right meter shows the reference (REF) level. This is independent from the physical routing which is set using the input routing selection.
- Enable / Disable Driver select the driver types you want to use. See page Error: Reference source not found for details.
#### Wave I/O setup
- The setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit – 24Bit – 32 Bit resolution of one sample, indicates how exact the analog signal is converted into digital signal.
- Use extended format use the new type of microsofts®'s wave format description for the access to the input driver. You might need to activate this setting to be able to select higher resolution values.
- Latency (ms) specify the size of the data blocks used for transferring the audio data from and to the soundcard. Lower values mean faster response but the data are more likely to be affected by performance issues.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with the volume setting when changing any settings.
- In the extended format windows expands audio data to the selected resolution, even if the hardware does not support that resolution. A selection of a resolution not supported by the hardware creates more data and processor load then necessary.
- The Apply button saves the current settings of the wave I/O and closes the window, but does not select Wave I/O as the current audio driver. The input driver selection must be confirmed using the Save button in the top menu bar.
#### Directsound(R) I/O setup
- This setup is divided into two sections, which contain the same settings, the upper one for the audio input and the lower one for the audio output.
- At the bottom there is the samplerate selection. This is a global setting.
- Soundcard selection select the soundcard you want to use for this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
- Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.
- this function. It is possible to use different soundcards for input and for output.
- 16Bit / 24Bit / 32 Bit resolution of one sample, describes how precise the analog signal is converted.
- Use extended format use new type of the wave-format description.
- Latency (ms) specify the size of the data blocks used for transferring audio data from and to the soundcard. A lower value means faster response but the data is more likely affected by performance issues. You should not change the default value of unless you're facing audio I/O related problems.
Hints:
- After every change of the settings SATlive rebuilds the list of supported samplerates. This may take same seconds.
- To get the best quality you should use the highest samplerate available.
- Different settings may produce different output levels. So please be very careful with volume setting when changing any settings.
- If the extended format is used, then windows expands the audio data to the selected resolution. Therefore selection of a resolution not supported by the hardware just creates more data and processor load as necessary.
- The Apply button saves the current settings of the directSound I/O and closes the window, but does not select the directsound driver as the current audio driver. The audio-driver selection must be confirmed using the Save button in top menu bar.