BX23
Weighing Indicator
Technical Manual

### Electrical Connection Recommendations
1. Always remember that BX23 indicator is very low voltage measuring instrument in the industrial environment. Your proper installation increases reliability and performance of the instrument.
2. The trained person should interface the instrument because of the 230 VAC voltage in the instrument and against malfunction at installation.
3. If the energy condition of your plant is bad, prepare a special power line.
4. The quality of your plant grounding will provide weighing accuracy and the safety of your indicator. If grounding of your plant is bad, prepare a special power line and grounding.
5. The instrument shall be powered off before interfering the instrument.
6. Powered off the instrument before connecting or disconnecting any peripheral instrument.
7. The shielded cable and ground connection of the shield will increase the immunity against electrical disturbances. Shields of cables can be connected to the glands bodies as indicated in the Error: Reference source not found(recommended) or to the grounding pin of the terminal which is cable connected.
8. All required electrical connections should be done as described in the installation section, page 5.
If you have to service the indicator, turn the power off and wait at least 30 seconds before opening housing to interfere it.

Place of the Peripheral Connections

The electrical terminals of the main board are shown in the picture below.

 RS232-1
INT1
Calib.
DIP switch
RS232-2 RS485
INT2 INT3
Power
230 VAC

Load Cell

Modbus
(Option)
Mini USB
Connectors of optional board
Alibi
(Option)
J14, automatic switch on.
(If short circuit)

Figure 4.1 – The main board
## Installation
Follow the steps below carefully to install the instrument.

#### Step 1: Preliminary Preparations

Before starting the installation,
1. Select the weighing terminal location. Refer to recommendations in the previous section.
2. Prepare the PE protective earth cable to grounding the indicator housing. The protective earth should be as good as possible for scale reliability.
3. Prepare the power source near the weighing terminal.
4. Prepare the cabling tray, pipe etc. from platform to the weighing terminal.

#### Step 2: Mechanical Installation

Be sure that the mechanical installation of the instrument is done properly for long time life before starting electrical installation as descripted below. Installation of the desk type and wall type housings are identical due to same kits’ usage as shown below.

If you use the weighing terminal on desk without additional mechanical hardware, place the instrument on to the desk. Go to the next step.

Wall mounting kit to use the instrument on the desk
Install the mounting kit to the rear of the instrument with 4 pcs M5 x 15 mm screws.	Place the weighing terminal on the desk after adjusting the view angle.

Wall mounting
Dismounting the wall mounting kit before installation and follow the steps below.

Install the wall mounting part of the kit to the wall with 3 screws.	Install the indicator holder of the kit to the rear the instrument with 4 pcs M5 x 15 mm screws.	Reassemble the wall mounting kit as seen in the picture.	Adjust the angle of weighing terminal.