Method Investigation Report

## MI Title: Determining the Suitability of the Soil Mehlich 3 Extraction in 200mL PET Bottles

MI Number:		Section:	Soils

MI Leader:	Sukhjeet	Test Method:	SM3

### Outline of Method Investigation

#### Definition of the Problem / Reason for the MI
### With the introduction of ASTRo (Automatic Soil Tapping Robot) in the Soil Lab, there is a need for the Mehlich 3 extraction to be performed in 200mL PET bottles instead of 125mL Erlenmeyer flasks that are currently used for the test. This will allow the Mehlich 3 test to be tapped out robotically, reducing human error and saving time.
### The “spin-off” benefit of carrying out the Mehlich 3 extraction in PET bottles is that there will no longer be boron contamination from borosilicate glass flasks being exposed to an acidic extractant containing fluoride. Plastic extraction gear will allow us to reduce the reported detection limit for boron and give us more confidence in these results.

#### Scope
The scope of this method investigation covers the 5 minute extraction of P, K, S, Ca, Mg, Na, Fe, Mn, Zn, Cu, B, Co and Al in a wide range of agricultural and horticultural soils using the Mehlich 3 extractant in 200mL PET bottles. The extraction ratio used in this was 1:10 (6cc of soil in 60ml extractant). All analytes were determined using the existing Inductively Coupled Plasma – Optical Emission Spectroscopy (ICP-OES) method.

#### Aim(s), Goal(s) or Desired Outcome(s)
The aim of this investigation is to demonstrate that the Mehlich 3 extraction carried out in PET bottles matches the performance of the current extraction carried out in Erlenmeyer flasks.
### Olsen Phosphorus, Soil Exchangeable Cations, Extractable Acidity, Soil Sulphate, Organic Sulphur and Anion Storage Capacity extractions have all recently been validated using 200mL PET Bottles. These validations demonstrated that extractions carried out in 200mL PET bottles compare very well to extractions in 125mL Erlenmeyer flasks.
### MI Plan
- SM3 runs will be carried out in duplicate, one run will be extracted in Erlenmeyer flasks, the other will be extracted in PET bottles. This will be continued until enough data is collected for a comparison scatter plot. Around 10 runs will be required, depending on run length.
- QC data from the runs above will be graphed on control charts to give an indication of the accuracy and long term stability of the proposed method compared to the existing method.
- The boron detection limit will be recalculated in PET bottles. It is unnecessary to check the detection limit of other analytes as data from the previous validations for these vessels have shown that the DL is comparable to Erlenmeyer flasks.
- To extract Hagerstown Silt Loam (a Certified Reference Material for Mehlich 3) and Interlaboratory Comparison Programme samples using PET bottles as an external check on results.
### MI Notebook (Details / Experimental)

Alternate Method Comparison
Twelve Mehlich 3 runs (141 samples) were extracted using the current and proposed extraction vessels. The results have been summarised in the scatter plots below.

Figures 1-13: Alternate method comparisons.

Most of the PET bottle results compare well to the Erlenmeyer flask extraction. On average, the results obtained using PET bottles are about 5% higher. The CRM and ILCP data shown in the following sections suggest that this difference is insignificant and in a few cases, the PET bottles produce results closer to the ILCP mean.

The boron plot is noisy at low levels and highlights the issue with boron extracting from glass Erlenmeyer flasks.

Certified Reference Materials
Hagerstown Silt Loam is a CRM for the Mehlich 3 test. This sample was extracted using both PET bottles and Erlenmeyer flasks. The results have been summarised in the graphs below. The error bars indicate ±2SD provided in the “Hagerstown Silt Loam – Soil Characteristics” document (KBI: 4822).

Figures 14-22: CRM comparison.

Mehlich 3 results produced by both Erlenmeyer flasks and PET bottles are within the 95% confidence intervals of the provided CRM values. The large standard deviations provided for this CRM demonstrates how variable this test is.

The boron result from PET bottles is higher than the Erlenmeyer flasks and closer to the certified value. This is because we are currently over-correcting boron results in Erlenmeyer flasks. This was seen in all the runs throughout this trial. Boron results from a typical run are shown in the table below.

Table 1: Boron blank correction

Table 1 shows that the boron extracted from glass Erlenmeyer flasks is inconsistent across samples and blanks. Relatively more boron is extracted from flasks in sample blanks than in the samples themselves.

Interlaboratory Comparison Programme
As an external check of the accuracy of results, this laboratory participates in ILCP's. The two programmes this laboratory is part of for Mehlich 3 are run by the Australasian Soil and Plant Analysis Council (ASPAC) and Wageningen University. Typically 7-13 labs participate in these two programmes for Mehlich 3.

Two previously submitted samples were retested using PET bottles and Erlenmeyer flasks. Results have been summarised in the graphs below. For the ASPAC sample, error bars represent the provided “displayed limit”. For the Wageningen sample, the “mean” (in yellow) is the Normal Distribution Approximation (NDA) mean and the error bars represent ±2 NDA SD. There was not enough data for Wageningen Al and B results (fewer than 5 labs), so the NDA mean and SD weren't provided. In the case of these two results, the median and Mean Absolute Deviation (MAD) were used instead. Little emphasis should be put on these statistics because of the small data set.

Figures 23-34 ILCP Comparison

Soil Mehlich 3 extractions from PET bottles compare well to ILCP results with most results easily falling within the set limits.

In-house QC
In-house Quality Control data provides an indication of the accuracy and long-term stability of soil tests. QC samples are prepared and limits set according to the Soil and Plant Quality Control Sample Preparation and Limits Setting SOP (KBI: 14530).

On the control charts shown below, the black line represents the mean, ±2 standard deviations are shown as pink lines and ±3 standard deviations are shown in red. In general, QC results that lie within 2SD are deemed acceptable.

Figures 35-60: Soil 41 and Soil 32 Mehlich 3 Control Charts Green lines = PET Bottles, Blue lines = Erlenmeyer Flasks.

The QC results collected for Mehlich 3 analytes using PET bottles compare well to Erlenmeyer flasks. There was a noticeable difference in boron results due to plastic vessels being used instead of glass. These limits need updating when the Soil Lab. begins using PET bottles for this extraction.

Boron Detection Limit
The reported detection limit for boron is currently set to 0.5mg/L. This can now be lowered as we will no longer extract these samples in borosilicate glass.

To calculate the MDL, a low boron sample was extracted across 9 days and the SD was multiplied by 2.896 (Student's t-factor for 8 degrees of freedom).The table below summarises the results used in the calculation of the new boron detection limit.

Table 2: Boron Detection Limit calculation.

The final boron detection limit has been calculated as 0.14 mg/L. It is recommended that boron's MDL gets changed to 0.15mg/L.
### Results Summary, Interpretation and General Discussion
A number of different aspects have been covered in this investigation to assess the suitability of extracting Mehlich 3 analytes in PET bottles. The accuracy and long-term stability were comparable to the current method as demonstrated by in house soil QCs. In addition, the Mehlich 3 analytes extracted using PET bottles compared well to other laboratories that participate in the ASPAC and Wageningen ILCP's.
Changing to 200mL PET bottles means that soil samples requiring the Mehlich 3 extraction can be tapped robotically using ASTRo. This will reduce sample transfer and soil tapping errors. Plastic extraction vessels will also allow us to report a lower boron detection limit.

### Approval of Method Investigation Outcome
Approved by MI Leader:		Date:
Approved by Guardian:		Date:

<End of Document>