# Summary
8 channel resistive bridge transducer or
voltage measurement system
- 24-bit resolution
- max. 102.4kHz sampling
- configurable input architecture
- DSP Postprocessing
- TEDS electronic transducer data sheet

# Block Diagram
one of eight identical channels
# Preliminary Specifications
Analog Inputs	Bridge Mode:
   8 resistive bridge transducers
   120, 350, and 1000 Ohms,
   quarter-, half-, and full-bridge
Voltage Mode:
        8 differential inputs, single-ended with external connection

Connector	8-pin RJ45 for each channel or 9-pin Lemo

RJ45-Pinout:	1 - EXC+
2 - EXC-
3 - SIG-
4 - SHUNT / TEDS-IO
5 - SHUNT / TEDS-GND
6 - SIG+
7 - SNS-
8 – SNS+

Lemo Pinout:	1 - EXC+
2 - SNS+
3 - SIG+
4 - SIG-
5 - SNS-
6 - EXC-
7 - SHUNT / TEDS-IO
8 - SHUNT / TEDS-GND
9 - nc

Resolution	24 bit

Sampling Frequency	max. 102.4kHz, programmable

Bridge Excitation	DC, differential, programmable for each channel 0.. ±6V, 12-bit resolution
short-circuit protected, max. current 42mA

Bridge Completion	half-bridge:	programmable for each channel, 0.05%
quarter-bridge:	programmable for each channel 120, 350, 1000 Ohms, 0.1%,
Input range	Signal Inputs (SIG): ± 100mV, ± 1V, ± 10V, programmable for each channel
Sense Inputs (SNS): ± 10V

Noise	100mV range: SNR = 90dB at 102.4kHz sampling, input terminated with 120 ohms =
 1.4µVRMS/V at ± 1.25V bridge supply
 0.7µVRMS/V at ± 2.5Vbridge supply
 0.35µVRMS/V at ± 5V bridge supply

1V range: SNR = 98dB at 102.4kHz sampling, input terminated with 50 ohms
10V range: SNR = 104dB at 102.4kHz sampling, input terminated with 50 ohms
additional downsampling in DSP will improve SNR by 3dB for each downsampling step

THD	< -90dB

Crosstalk	< -100dB at 1kHz between channels

Phase Error	< 1° at 1kHz between channels

Input Protection	max. ± 25V

CMRR	120dB

Calibration	1. internal programmable shunt calibration for each channel, 55kOhm 0,05%,
   connected from EXC- to SIG+
2. common shunt calibration resistor for all channels, on two dedicated pins,
   shared with TEDS, can be externally connected to any bridge arm
 shunt calibration resistors in sockets
TEDS	8 channels, Class 2

Filter	per channel programmable digital filter:
4th order lowpass , or 4th order highpass, or 2nd order bandpass
lowpass 1Hz .. 100Hz, highpass 10Hz .. 10kHz (max. 0.25*fs)

Signalling	red/green bicolor LED for each channel

DSP	2 x Texas Instruments TMS320C6747 floating-point processor

Memory	64 Mbytes SDRAM
8 Mbytes non-volatile Flash Memory

System Interfaces	McASP serial communication port, 2 out, 1 in, 24MBit/sec per lane
USB 2.0 Port, 480MBit/sec
Trigger, Synchronisation, and GPIO ports

Supply Voltage VCC	DSP	+3.3 V ± 5%
Analog	12V ± 10%

Power Consumption	TBD

Operating Temperature	0 .. +50 °C

Size	220 x 100 mm
# Input Configurations