Introduction
Symbols used in this manual
	Indicates a safety hazard; the user should take the precautionary measures as stated to avoid harm to himself or others.
	Indicates a burn hazard to the patient.
	An alert to high voltages that may constitute a risk of electric shock.
NOTE:	Indicates special operating features.
	Technical details

Package
GD300 comes in a package comprising:
- A footswitch pedal
- Patient skin pads
- An electrosurgical pencil
Front Console

1

2

3

4

4

4

7

7

7

5

5

5

6

6

6

8

9

10

11

1. General Indicator Lights
  1. Power Indicator Light: lights up when power is connected.
  2. Neutral Indicator Light: lights up when patient return electrode (patient skin pad) is connected.
  3. Footswitch Indicator Light: lights up when footswitch is connected.
  4. Alarm Indicator Light: lights up when alarm is activated.
2. Mode Selection Button: switches between monopolar and bipolar modes.
3. Mode Selection Indicator Light: indicates the currently activated mode (monopolar or bipolar).
4. Submode Selection Button: switches between the submodes under each mode.
5. Power Adjustment Buttons: decreases or increases the maximum power supply under each mode.
6. Submode Selection Indicator Light: indicates the currently activated sub-mode.
7. Power LED: indicates the maximum power supply.
8. Socket for active electrode (electrosurgical pencil).
9. Socket for patient return electrode (patient skin pad).
10. Socket for electrosurgical forceps.
11. Main power switch.

Rear Console

1

3

2

4

5

6

9

7

8

1. Power socket (220V / 50Hz)
2. First set of fuses
3. Second set of fuses
4. Socket for separate ground cable
5. Footswitch socket
6. Volume knob
7. Speaker holes
8. Ventilation fan holes
9. Label sticker

Warranty
The warranty period is one year from the date of installation and shall cover repair costs of the equipment. Such warranty shall not cover the equipment if it was used for any purpose other than that which it was sold for or which have been stored for an excessive period or in unsuitable conditions or in respect of which defects have not been reported to the retailer within the period of one year from date of installation.

Operating Instructions
Operation Modes
Mode	Sub-mode	Frequency	Rated Power	Waveform	Vp-p
Monopolar Cut	Pure Cut 400 kHz 0 – 300W Sinusoidal 2500 V / Blend 1 400 kHz 0 – 250W Broken Sinusoidal (50-50) 2600 V / Blend 2 400 kHz 0 – 200W Broken Sinusoidal (40-60) 2700 V
Monopolar Coagulation	Contact 400 kHz 0 – 80W??? Broken Sinusoidal / (ratio?) 2900 V / Spray 400 kHz 0 – 100W Broken Sinusoidal (ratio) 4000 V
Bipolar	Cut 800 kHz 0 – 100W Sinusoidal 1000 V / Coagulation 800 kHz 0 – 70W Broken Sinusoidal (ratio) 700 V

Sterilization
Sterilize the equipment using standard sterilizing procedures.
Pre-operative Safety Procedures
1. Clean up any combustible liquids that could have accumulated under the patient’s body, in the depressions on the body and in orifices. Gauzes and pieces of cloth soaked in volatile liquids also pose a fire hazard.
2. The interference generated from use of the ESU may interfere with the normal functioning of pacemakers. If in doubt, consult a cardiologist. Other medical equipment may also be affected.
Setting Up
NOTE:	Ideally, connect the peripherals in following order to ensure that the relevant alarms and signals are in good working order.

1. Connect the power cable and switch the generator on. In the next 10 seconds, the generator will conduct self-checks.
2. If the self-checks are successful, normal operation will begin. The alarm on the machine will be activated at this time. If there is no sound, turn up the volume using the volume knob.
Ensure that the volume knob is turned up and the audio signals are clearly audible. Using the equipment without the sound signals creates a high risk of injury to the patient.

Paste the skin pad onto the patient. Ensure that the as much of the pad is in good contact with the skin.
Avoid pasting the skin pad on area with poor blood circulation, bony prominences, or irregular body contours. Doing so prevents the heat generated during electrosurgery from being dissipated and puts the patient at high risk of burns.

Connect the skin pad and pencil to the generator with the relevant cables. Connect a footswitch if desired.
Keep the electrosurgical pencil away from the patient as far as possible and keep exposed parts of the patient’s body insulated to avoid accidental burns.

The alarm is switched on at impedance levels greater than 150Ω. Under such a situation, only the mode selection button is operational. Once the machine detects normal impedance levels (40Ω and 150Ω) at the skin pad, the alarm will be switched off.
The skin pad measures impedance on the skin by way of a small electric current. Poor contact with the skin will create high impedance that will prevent the machine from activating the current. Without this self-check mechanism there is a high risk of burns due to the poor contact.
If the alarm is not deactivated even after the skin pad and the pencil have been attached, check the contact at the skin pad again.

Performing Electrosurgery
To activate the machine and start electrosurgery, the following methods are available:
Mode	Pencil Switch	Footswitch
Monopolar Cutting	Yellow button	Yellow pedal
Monopolar Coagulating	Blue button	Blue pedal
Bipolar Cutting	n.a.	Yellow pedal
Bipolar Coagulating	n.a.	Blue pedal

When the machine is activated, the generator will emit a tone. A different tone will be emitted for each of the three modes (monopolar cutting, monopolar coagulating and bipolar). Note the variations in tone to avoid using the wrong modes by accident.
Footswitch
The footswitch can be used in place of the buttons on the pencil. When the footswitch is connected, the footswitch indicator should light up.
Time Limit
For safety, the generator will cut the power after 3 minutes of continuous running. The alarm will be turned on and the equipment will be non-operational for a while.
Safety
- Keep the patient/subject insulated from all electrical grounds. Insulate the patient’s body from his other body parts (hands etc.).
- Position any patient monitors, if applicable, as far from the electrodes as possible to avoid interference.
- Use only the minimum necessary voltage necessary for the job.
- Do not conduct electrosurgery when the connections at the sockets are loose.
- Do not short circuit the equipment.
- During the course of surgery, switch off the ESU when not needed
Troubleshooting
1. If power output drops unexpectedly and/or the pencil stops working normally, the contact at the pad may be loose. Check the pad’s contact with the patient’s skin.
2. If the sound signals are activated but no power is being delivered at the electrodes (possibly together with an alarm), first switch off the equipment and inspect the fuse. Then check the pad and its cabling and make the necessary replacements.
3. If the above measures do not work, contact your supplier.
Design Notes
Classification:	Class I, Type BF
Power Supply:	Single-phase 220V ±20% / 50 Hz
Maximum Power:	750W
Working Mode	Intermittent loading & Continuous running?
Product Number

- Keep working parts isolated from an electrical ground
- This device is protected against defibrillators’ discharge effects
- Use a power source stabilizer?
Environmental Specifications
During Operation
Temperature:	0 to 35 °C
Humidity:	30% to 75% Relative Humidity (RH)
Atmospheric pressure:	70 kPa to 106 kPa

Storage
Temperature:	-10 to 40 °C
Humidity:	<75% Relative Humidity (RH)

Ensure that the grounding in the main power supply is reliable. If not, connect the additional grounding cable. Both the foot pedal switch or/and the switch on the pencil can be used.
Replacement of Parts
There are no user-serviceable parts within the generator. High voltages are present in the equipment. If the machine is faulty, do not disassemble. Contact a technician.

Cables
The electrode cablings are specially designed for high frequency usage. Replace them when damaged or when connections with their sockets are loose. For replacements, contact the distributor.
Pencil, Forceps, Pad
Are these disposable?
Fuses
The fuses are located at the back of the generator. To remove a fuse, use a screwdriver with a slotted tip, depress the button and turn anti-clockwise. Replace the fuse on the button, insert the button into the hole, press and turn clockwise.
Maintenance and Servicing
Carry out monthly runs of no less than 20 hours at room temperature and inspect the equipment. The generator must be serviced at least 4 times annually to clear the dust and to ensure the correct and safe functioning of the parts, peripherals and insulation.
Power-Impedance Charts