DEPARTMENT OF HEALTH AND HUMAN SERVICES

FOOD AND DRUG ADMINISTRATION

CENTER FOR DEVICES AND RADIOLOGICAL HEALTH

GENERAL AND PLASTIC SURGERY DEVICES PANEL

OF THE MEDICAL DEVICES ADVISORY COMMITTEE

OPEN SESSION

60th Meeting

This transcript has not been edited and FDA makes no representation regarding its accuracy

Tuesday, July 9, 2002

8:00 a.m.

Gaithersburg Holiday Inn

Two Montgomery Village Avenue

Gaithersburg, Maryland

PARTICIPANTS

Thomas V. Whalen, M.D., Chairman

David Krause, Ph.D., Executive Secretary

VOTING MEMBERS:

Phyllis Chang, M.D.

Michael A. Choti, M.D.

David L. DeMets , Ph.D.

Robert L. McCauley, M.D.

Michael J. Miller, M.D.

TEMPORARY VOTING MEMBERS:

Nancy A. Dubler, LLB

Amy E. Newburger, M.D.

CONSUMER REPRESENTATIVE:

LeeLee Doyle, Ph.D.

INDUSTRY REPRESENTATIVE:

Debera M. Brown

C O N T E N T S

Conflict of Interest and Opening Remarks,

David Krause, Ph.D.	5

Open Public Comment:

Ms. Tammy Griffiths	16

Ms. Denise King	21

Ms. Cheryl Valput	23

Ms. Melanie Palmer	26

Joseph Bubinak, M.D.	28

Ms. Kim Hoffman	32

Ms. Margaret Volpe, Y-ME 	38

Ms. Mary McDonough 	42

Ms. Cynthia A. Pearson, National Women's Health

Network (read by Toniya Barron)	49

V. Leroy Young, M.D., American Society of

Plastic Surgeons	54

Laurie Casas, M.D., American Society for

Aesthetic Plastic Surgery	61

Ms. Sheila Crigler, WIN Against Breast Cancer	70

Diana Zuckerman, Ph.D., National Center for

Policy Research for Women and Families	74

Jae Hong Lee, M.D., MPH, National Center for

Policy Research for Women and Families	81

Ms. Susan Pope Helman	87

Ms. Sheila Crigler (for Ms. Elizabeth Mullen

WIN Against Breast Cancer	93

Ms. Cherien Dabis (read by Ms. Mary McDonough) 	95

Panel Update Regarding Post-Approval Conditions

of Approval for Saline-Filled Breast Prosthesis:

Mentor Corporation:

Background Information Maher Michael, M.D.	103

Post-Approval Study, Mr. Cliff Kline	106

Focus Group Study, Ms. Donna Crawford	115

Retrieval Study Report, Post-Approval Fatigue

Testing and Post-Approval Real-Time

Shelf-Life Testing, Mr. Ronald Crouther	119

Overall Summary, Maher Michael, M.D.	127

FDA Presentation:

Post-Approval Study, Sahar Dawisha, M.D.	165

Focus, Retrieval, Fatigue Testing and Shelf-Life

Studies, Ms. Samie Allen	175

Panel Discussion 	179

C O N T E N T S (Continued)

Panel Update Regarding Post-Approval Conditions

of Approval for Saline-Filled Breast Prosthesis (continued):

Inamed Corporation:

Introduction, Ronald J. Ehmsen, Sc.D.	202

Post-Approval Study , Audrey J. Weiss, Ph.D.	203

Focus Group Study, Kim Croyle	220

Retrieval, Fatigue and Shelf-Life Studies,

Mr. Thomas Powell	222

FDA Presentation:

Post-Approval Study, Sahar Dawisha, M.D.	241

Focus, Retrieval, Fatigue Testing and Shelf-Life

Studies, Ms. Samie Allen	247

Panel Discussion 	253

P R O C E E D I N G S

Call to Order, Conflict of Interest

and Opening Remarks

DR. KRAUSE:  If everybody would take a seat, I would like to start the meeting.  I would like to say good morning to everybody.  Thank you for coming.  We will try to start the public testimony as close to 8:30 as possible for those people who looked at the schedule to see what time the public comment period would be.  I don't want to rush things but I want to make sure everybody who wants to speak who is on our list gets a chance.

We are ready to continue the 60th meeting of the General and Plastic Surgery Devices Panel.  My name is David Krause.  I am the executive secretary of the panel.  I am also a biologist and a reviewer in the Plastic and Reconstructive Surgery Device Branch, in the Division of General, Restorative and Neurological Devices.

I would like to remind everyone to please sign in on the attendance sheets which are just outside the door.  You can also pick up an agenda, a panel meeting roster and information about today's meeting at the tables just outside the door.  The information provided includes how to find out about future meeting dates through the advisory panel phone line, and also how to obtain meeting minutes or transcripts.

Before turning the meeting over to Dr. Whalen, I would like to read today's conflict of interest statement.  The following announcement addresses conflict of interest issues associated with this meeting, and is made a part of the record to preclude even the appearance of an impropriety.

To determine if any conflict existed, the agency reviewed the submitted agenda and all financial interests reported by the committee participants.  The conflict of interest statutes prohibit special government employees from participating in matters that could affect their or their employers' financial interests.  However, the agency has determined that participation of certain members and consultants, the need for whose services outweighs the potential conflict of interest involved, is in the best interest of the government.

Therefore, a waiver has been granted for Dr. DeMets for his interests in a firm that could potentially be affected by the panel's recommendations.  The waiver allowing him to participate fully in today's deliberations involves a contract to his employer and his consulting services with a competing technology firm.  His employer receives funding between $100,000.00 and $300,000.00 per year on the contract that is unrelated to today's agenda.  For his unrelated consulting service, he receives less than $10,000.00 a year.  Copies of this waiver may be obtained from the agency's Freedom of Information Office, Room 12A-15 of the Parklawn Building.

We would like to note for the record that the agency took into consideration certain matters regarding Dr. DeMets and Dr. Miller.  Dr. DeMets reported an interest with a firm at issue, but in a matter that is not related to today's agenda.  Dr. Miller reported that his institution has a current related involvement with a firm at issue.  The agency has determined that these individuals may participate fully in the panel's deliberations.

In the event that the discussions involve any other products or firms not already on the agenda for which an FDA participant has a financial interest, the participant should excuse him or herself from such involvement and the exclusion will be noted for the record.

With respect to all other participants, we ask in the interest of fairness that all persons making statements or presentations disclose any current or previous financial involvement with any firm whose products they may wish to comment upon.

I would now like to turn the meeting over to the Chairman, Dr. Whalen.

DR. WHALEN:  Thank you, Dr. Krause.  Good morning.  My name is Dr. Thomas Whalen.  I am professor of surgery and pediatrics at Robert Wood Johnson Medical School in New Brunswick, New Jersey.

Today our panel will be making a review and update of previously approved premarket applications which were approved in May of 2000.  Although we are in our second day of business as a panel, I believe we have a considerably different audience than we did yesterday, so we will ask our panel members to reintroduce themselves to the audience, acknowledging that these are panel members who are giving of their time to help the FDA in these matters, and I would also ask the FDA staff at the table to introduce themselves as well.

Each person is asked to introduce himself by name and state specialty and position title, starting on my left with Ms. Brown, please.

MS. BROWN:  I am Debera Brown.  I am the vice president of regulatory affairs for Broncus Technologies.  I am also the industry representative for this panel.

DR. DOYLE:  I am LeeLee Doyle.  I am professor of obstetrics and gynecology and associate dean for continuing medical education and faculty development with the University of Arkansas for Medical Sciences, College of Medicine, and I am the consumer representative on the panel.

DR. MCCAULEY:  Robert McCauley, professor of surgery and pediatrics, University of Texas Medical Branch and chief of plastic surgery services for the Shriner's Burns Hospital in Galveston, Texas.

DR. DUBLER:  I am Nancy Dubler.  I am an attorney and the director of the Division of Bioethics at Montefiore Medical Center, and a professor of epidemiology and social medicine at the Albert Einstein College of Medicine.

DR. CHOTI:  I am Michael Choti, associate professor of surgery and oncology, in the Division of Surgical Oncology at Johns Hopkins Hospital in Baltimore, Maryland.

DR. NEWBURGER:  I am Amy Newburger.  I am a dermatologist in private practice, in New York.  I am an attending physician at White Plains Hospital Medical Center.  I teach at St. Luke's Roosevelt Medical consortium in New York.

DR. DEMETS:  I am David DeMets.  I am professor and chair of the Department of Biostatistics and Medical Informatics at University of Wisconsin, in Madison.  I am a statistician by degree and I have a long history and interest in clinical trials.

DR. CHANG:  I am Phyllis Chang, associate professor in the Division of Plastic Surgery, Department of Surgery, and Division of Hand and Microsurgery, Department of Orthopedic Surgery at the University of Iowa College of Medicine, in Iowa City.

DR. MILLER:  I am Michael Miller, associate professor of plastic surgery at University of Texas, MD Anderson Cancer Center.

DR. WITTEN:  Celia Witten, division director of the Division of General, Restorative and Neurological Devices at the FDA.

DR. WHALEN:  Thank you all.  I would like to note for the record that the voting members present constitute a quorum, as required by 21 CFR Part 14.  At this point, I would like to turn the meeting back to you, Dr. Krause, for a moment.

DR. KRAUSE:  Thank you, Dr. Whalen.  Just for the record, I would like to say that Dr. Miller, Dr. Chang, Dr. DeMets, Dr. Choti and Dr. McCauley are voting members on the panel.  Ms. Dubler and Dr. Newburger are temporary voting members on the panel for today's meeting.

I would like to make some brief remarks before we get into the main body of today's panel discussion.  Today is the last meeting for a number of our distinguished panelists.  It is the last meeting for Dr. Whalen, our Chairman, and for Dr. DeMets and Dr. Chang who are voting members.  I would also like to acknowledge that Dr. Joe Boykin's term has also ended.  He was unable to attend today.  These panelists have served this panel with distinction for the past four years and I would like to thank them for their service.  I would like to ask Dr. Witten to please make a few remarks at this time.

DR. WITTEN:  Thank you.  We are very sorry to be saying goodbye to our panel members who are with us for the last time today.  We rely on our panel members for their expertise in helping us make our regulatory decisions.  Panel members give their time to us to serve FDA and, in doing so, serve public health in helping us make good regulatory decisions.  I am particularly sorry about saying goodbye to the three people at the table today.

I have a letter, singed by Linda Sudan, the senior associate commissioner for communications and constituent relations, which I will read, the same letter to all the panel letters who are leaving:

Dear Dr. Whalen, I would like to express my deepest appreciation for your efforts and guidance during your term as a member and chair of the General and Plastic Surgery Devices Panel of the Medical Devices Advisory Committee.

The success of this committee's work reinforces our conviction that responsible regulation of consumer products depends greatly on the experience, knowledge and varied background and viewpoints that are represent on this committee.  In recognition of your distinguished service to the Food and Drug Administration, I am pleased to present you with this plaque, which is not enclosed but will be coming separately.

I have the same letter for Dr. Whalen, Dr. Chang and Dr. DeMets and also for Dr. Miller.  We will be hoping to see you again as consultants on some projects in the future.

DR. KRAUSE:  Thank you, Dr. Witten.  I would like to now turn the meeting back over to Dr. Whalen.

Open Public Comment

DR. WHALEN:  Thank you.  We will now proceed with the open public comment session of the meeting.  I would like to remind everyone who wishes to address the panel to please speak clearly into the microphone as the transcriptionist is dependent on this means to provide an accurate record of this meeting.  In addition, if anyone has printed copies of their remarks, if they could hand them over to the table on my left, to the transcriptionist, that would be deeply appreciated.

At this time, Dr. Krause has some additional instructions for those who will be testifying before the panel.  Dr. Krause?

DR. KRAUSE:  Thank you, Dr. Whalen.  I have some instructions for those of you who will be testifying to the panel this morning.  We are requesting that all persons making statements during the open public hearing of the meeting disclose whether they have financial interests in any medical device company, or if their trip to this meeting has been paid for by someone else.

Before making your presentation to the panel, in addition to stating your name and affiliation, please address the following questions.  I will read the questions into the record once so that we don't have to do it when every speaker speaks.  The questions are as follows:

Question one, has your travel and/or accommodations been paid for or will they be reimbursed by someone else?  If so, please state who.

Question number two, please indicate whether you have financial ties, including grants or other compensation, with industry or health professional societies.

Question number three, please indicate whether you are a party to or a witness in a pending lawsuit related to breast implants.

Question four, do you derive a portion of your income from surgical procedures using saline-filled implants or from treating patients with complaints that they believe are related to saline-filled implants?

I would like to have the attention of all the individuals who have registered to speak to the panel today.  You have all been put in the proper order by Anne Marie.  Anne Marie will direct you to the podium or, hopefully, she has given you some instructions.  At this point, we are going to limit everyone to five minutes.  At the end of each presentation, please stay at the podium in order to give the panelists a chance to ask you any questions, if they have any questions.

I would now like to return the meeting to Dr. Whalen.

DR. WHALEN:  Thank you, Dr. Krause.  We are required by law to allot one hour for public testimony.  Today, because of the nature of the topic that we are dealing with, we have decided to allot two full hours.  I realize, by the nature of the topic that we are dealing with and from the experience that we had as a committee two years ago in dealing with the PMAs on this, that there are some very deeply held convictions and sentiments about this topic.  So, I will apologize in advance if I cut you off at the five-minute limit, acknowledging the depth of your conviction about what you are speaking about.  I am not trying to be rude, but simply trying to accommodate everyone who wishes to speak on this topic today because of the number of people who have come forward wishing to speak.

Also, I have been given an order of speakers that I will follow.  I understand that there may be some mild confusion about who was to follow whom, but please bear with me on this.  At the end of going through everyone that I have had identified to me being called by name, I will have an open announcement for anyone else wishing to speak as long as we still have more of that two-hour time period for people to address the panel.

Therefore, the first person that I have identified is Ms. Arlene Nicole Cummings, from Implantinfo.com.  Is Ms. Cummings available?

[No response]

Very well.  The second one is Ms. Tammy Griffiths, representing herself.

MS. GRIFFITHS:  Good morning.  I would like to thank the committee for allowing me to speak today.  I come here today of my own accord--

DR. KRAUSE:  Could you excuse me a minute, you weren't here when I read the questions.  There are four questions you need to address before you start.  Let me read them again:  Have your travel or accommodations been paid for, or will they be reimbursed by someone else?  If so, please state who.

Second question, please indicate whether you have financial ties, including grants or other compensation, with industry or health professional societies.

Number three, please indicate whether you are a party to or witness in a pending lawsuit related to breast implants.

Question four, do you derive a portion of your income from surgical products using saline-filled implants or from treating patients with complaints that they believe are related to saline-filled implants?

If you could just answer those questions before you start.  That goes for everybody.  Thank you.

MS. GRIFFITHS:  Sorry about that.  No to all of the above questions.

I come here today of my own accord in support of the continued use of saline breast implants.  My name is Tammy, and I am 38 years old and the mother of four children, ranging in ages 5 to 13 years old.  I have been married for 15 years, and my husband and I own and operate a heating and air conditioning company in Texas.

I received a breast augmentation in January of 1999.  Breast feeding four children had taken a toll on me physically, and the availability of breast implants gave me an option.  To me, this was, indeed, a gift both physically and emotionally that I have never regretted.

Prior to my children, I was very content with my breast size which was proportionate to my small frame.  However, I was through with my childbearing years, and at the age of 34, I was very unhappy with the resulting shape of my breasts.  Nursing four babies had taken quite a toll on my physique.  With each child I would swell from my normal 34B cup to a 36D cup.  Once each child weaned, I would return to the 34B cup but each time with less lift texture than I had before.

After my fourth child was through nursing, my breasts were once again a 34B only they appeared depleted, flat and shapeless, and they no longer felt firm.  It was almost as if I had nothing at all.  I began working at the gym on a regular basis especially concentrating on my pectoral muscles, trying to build them back up, but to no avail.  Then I tried a pair of silicone bran inserts to wear underneath my clothing just so that it would appear that I had breasts.  However, when I looked in the mirror and it was just me, all by myself, I remained unsatisfied with my look and began to explore other options.  The mirror did not lie.

My husband Randy and I began to consider breast augmentation as an option.  He was, and still is, 100 percent supportive.  We both started reading everything we could find on the subject.  We read medical reports; we read opinions both for and against the procedure.  After research, we considered all the pros, cons, risks, rewards, side effects, both physical and emotional.  As with any elective surgery, we desired to make an informed decision and so we did.

After much discussion, Randy told me the final decision was mine to make.  He was happy either way and was just satisfied that the medical procedure was a safe option.  He continued to support me totally.  I made the decision to have surgery and have never regretted it.  I located a doctor in my area and again asked many questions.

Once I was satisfied with the doctor, we set a date and the surgery was completed without incident.  I now have Mentor saline implants.  I am slightly larger than I was prior to having children but, most importantly, I feel better about myself now than I did prior to surgery.

I can tell you that this is one of the best decisions or gifts, if you will, that I could have given myself.  I am now three and a half years postop and very happy with my results.  I have found myself feeling even more confident as a woman when I am in a business suit, a swimsuit or just attending a PTA meeting dressed in casual, conservative attire.  I truly believe I am a wife and a mother who is not only well adjusted in her life, but a wife and mother who is also self-confident in her physical appearance wherever I happen to be.

In closing, I would like to say that I feel it is most important that breast augmentation with saline implants continue to be an option for all women, whether they are cancer survivors, whether they need some other type of reconstruction, or if they just want something different.  Given good information and good medical care, saline implants are not only a viable solution for a variety of circumstances but they are safe as well, in my opinion.  Thank you.

DR. WHALEN:  Thank you, Ms. Griffiths.  Are there members of the panel who have any questions?

[No response]

Thank you very much.  Before the next speaker is called for, just to remind people--I certainly wouldn't remember all those four questions, so if anyone gets up to the podium and needs a prompt for the four questions, just indicate that to me and I will be happy to give you the topics of each of the areas that you need to start off with.  Also, to answer those questions obviously isn't part of your five minutes.  The next identified speaker is Ms. Denise King.

MS. KING:  In answer to all of your questions, the answer is no.  Thank you for the privilege of allowing me to speak today.  My name is Denise.  I am 43 years old and have been married for 22 years.  My husband and I have 6 children, ranging in ages from 17 to 7 years.

After nursing all of my children, some for as long as 18 months, and losing the large amount of weight that I had gained due to six pregnancies, my breasts were left deflated, hollow looking.  They looked as though they belonged to a woman twice my age.  It was both physically and emotionally uncomfortable to have the upper part of my body in this condition because I have worked so hard to maintain my weight.

After a lot of thought, research and extensive consultation with a board-certified plastic surgeon, after several discussions with other women that had had breast augmentation, asking both what they did and did not like about implants, on August 7 of 2001, and with the full support of my husband, I underwent a five-hour surgery to reconstruct my breast with an anchor style lift and under the muscle placement of saline implants.  I received nothing less than excellent care from my physician both before, during and after my surgery.

I have followed my doctor's instructions to the letter.  I see him every 12 weeks so that any possible problems may be detected quickly.  In the past 11 months I have been asked several times if I have ever regretted having the surgery.  My answer is that there are no regrets because this was something that I did for my self-esteem.

In closing, I would like to say that I believe that every woman should have the right to decide for herself whether she wants to pursue changing her shape through breast augmentation.  It is also my belief that women need to have available to them as much information about implants and the surgical procedure as possible, enabling them to make a well-informed decision.  Thank you again for your time.

DR. WHALEN:  Thank you, Ms. King.  Any questions?  	[No response]

Thank you very much.  Next is Ms. Cheryl Valput.

MS. VALPUT:  No to all the four questions.  Good morning.  I would also like to thank the committee today for allowing me to speak.  My name is Cheryl Valput.  I am 45 years of age, reside in Ohio with my husband of 22 years and two daughters, ages 19 and 16.  I am a full-time mom and business owner of a real estate company with my husband.

I had breast augmentation done three years ago, in Cleveland, Ohio by Dr. Richard Dowden.  This was one of the best decisions that I have made in my life.  I am not a model, a dancer or a movie star but I felt that this was something I needed to do for me.  I came to this decision for myself because I had to balance a difference in my breast size.  I felt uneven every time I looked in the mirror.  I was tired of stuffing one side of a bra, making sure I looked normal in a bathing suit.  My self-confidence was low and I knew that I had a choice in the matter to change my appearance.

I did research for over six months and read everything I could about breast implants, the good and the bad.  I finally decided that I was either going to proceed with the surgery or put it out of my mind for good.  I deserved to look my best and called Dr. Dowden.  I have saline anatomical McGhan's shaped implants under the muscle.  I went from an A cup to a C cup and finally they were both the same size.  Looking in the mirror wasn't so bad after all.

After surgery I could wear the clothing I desired and my self-confidence has been at its best.  When you feel so good on the inside, it is hard not to feel beautiful on the outside.  This is a very personal decision I made, along with my husband's support every step of the way.

For me, to stand and speak in front of the committee today took some thought.  I don't talk about this with everyone.  My own parents and sisters haven't been told about my surgery, but I feel very strongly about having the choice of getting breast implants.  It is not because I am trying to hide anything, but it is a decision that had to be made for myself.  Taking that choice away from women could not, and would not be right.  We color our hair; fix our teeth; choose the car we want to drive; and change our eye color.  We have the right to change our breast size and shape if we choose to.

I haven't had any complications and I am fully aware of the risk of implants.  The decision of this must rest on the balance between how strongly you want the benefits versus how worried you are about the risk.  Life is a risk, but I would rather go through life looking and feeling this way than without that.  Thank you.

DR. WHALEN:  Thank you.  Again, no questions?

[No response]

Thank you very much.  Next is Ms. Melanie Palmer.

MS. PALMER:  The answer to all your questions is no.  Good morning.  My name is Melanie Palmer.  I am 46 years old, from Ohio.  I have a husband, Tom, and two children ages 15 and 20.  Thank you for allowing me to speak to day on behalf of women who are seeking the possibility of breast augmentation.

Breast augmentation is not only a medical decision but a personal one.  The perception we have of ourselves defines who we are and how we introduce ourselves to the world.  Breast augmentation can help individuals who feel the need to project a better outward appearance and results in increased self-esteem and personal satisfaction.

As with any surgery, there are always risks as there are with breast augmentation.  When deciding on breast augmentation or any other surgery, you need to weigh the benefits and risks as well as expectations.  You want to make sure you are making this decision for yourself and not for anyone else.  You must realize something foreign is within your breast and that over time complications could happen, as with any surgery.

When I approached my husband about breast augmentation he was totally supportive of my decision, knowing I had thoroughly researched the procedure and having full knowledge of what the pros and cons of the surgery could entail.  I consulted with two certified plastic surgeons, both of whom were equally knowledgeable in their professions, making me feel quite secure in my decision.

After all of my research, I was made aware of the fact that my implants could rupture and may not last forever.  At informational web sites, such as FDA.gov and Implantinfo.com, by Nicole, I learned that I was at risk for capsular contracture, rippling, infection and other complications.  Despite these risks, I decided the benefits outweighed the disadvantages.  This was a personal choice made by me and my husband, not by society governing our decision.  Once all the options and any results were researched it was ultimately my decision.

I still feel the decision of breast augmentation was the right choice for me, even though it has only been one year and ten months ago.  I have no regrets and when asked would I do it again, my answer is a definite yes.  Has it made me feel better about myself?  Yes, in ways only a woman can feel from the inside out.  Thank you.

DR. WHALEN:  Thank you.  Are there any questions?

[No response]

Thank you, ma'am.  The next speaker, and I apologize if I mispronounce your name, is Dr. Joseph Bubinak.

DR. BUBINAK:  Thank you, Mr. Chairman and all members of the advisory panel.  I answer no to all four questions.  I am a board-certified hematologist and medical oncologist.  My undergraduate training led to degrees in mechanical engineering and science.  I retired from active practice two years ago.

One year ago, a former patient with a history of ruptured breast implant and a low platelet count called and said she was told that she was, quote, full of platinum, unquote.  Excessive platinum was found in skin, subcutaneous tissue, blood, urine and subsequently bone marrow.  This interested me because I have treated patients with platinum-based chemotherapy since 1976.  A low platelet count is a frequent side effect.  The first chemotherapy agent was cisplatin, which is also mutagenic, carcinogenic, leukemogenic and teratogenic.  Platinum has been found in the milk of patients treated with platinum compounds.

There is little information in the medical literature regarding an association of platinum with breast implants.  Unfortunately, not everyone has a good result.  Information obtained from many sources, including interviews with patients, physicians and researchers, reveal that some implant patients develop a variety of systemic complaints including malaise, hair loss, peripheral neuropathy which is sometimes disabling, loss of short-term memory, rash and other allergic manifestations, respiratory systems, constipation and anorexia, just to name a few.  In short, these are the same side effects people treated with cisplatin cytotoxic chemotherapy experience.  Increased intensity of systemic complaints commonly follow gross rupture of the implant.

I was astounded to learn that the catalyst used to manufacture the silicone for silicone gel and silicone elastomer for both gel-filled and saline-filled implants was platinum chloride, a highly reactive molecule and precursor to the chemotherapy agent cisplatin.  The chemistry of the polymerization process says that the platinum in ideal proportions is reduced to its inactive elemental form.  This, however, does not correlate with the amount of platinum found in tissues both proximate and distant from the implant site.

Two independent researchers now have found platinum in excessive concentration in tissues.  Capsule formation around the implant, a frequent complicating event, tells the world that this device is not inert.  Even without considering the seepage of low molecular weight silicones, the migration of reactive platinum alone can explain capsule formation.

One package insert states that the literature suggests that radiation therapy may increase the likelihood of capsular contracture, necrosis and extrusion.  I have witnessed this first hand.  In this regard, you should understand that platinum-based chemotherapy is commonly used explicitly to sensitize the target tissue to the effects of radiation therapy.

In conclusion, systemic systems do matter and must be listed as potential side effects in the package insert when patients are expected to give informed consent.  Likewise, reports of symptomatic improvement in patients following explantation must also be included.  Despite almost 40 years of clinical experience, there is not one good, solid, prospective epidemiologic study available.  The largest study that I saw was a retrospective study, just last year, and that study showed increased risk for respiratory and brain cancers, and a non-significant increased risk for leukemia of various types.

Milk from implanted mothers needs to be tested for bound and unbound platinum.  Reliable, generally accepted methodology for determining free and bound levels of platinum in any tissue must be developed with all speed.  Analysis for platinum DNA must be made available if other critical questions are to be answered.  Long-term ex vivo testing of implants, subjected to realistic stresses while immersed in physiologic biologically active media at 37 degrees, are needed.

Last, from a pure engineering perspective, considering the failure and complication rates, I wonder what reasoning could have led to the approval of these devices.  Reasonable assurance of safety and effectiveness--no one would argue against the beneficial psychological effects a positive body image will give.  Safety means unhurt, secure from any harm.

I urge this panel to approve only products that are truly safe and effective for all who desire them.  Thank you.

DR. WHALEN:  Thank you, doctor.  Any questions??

[No response]

Thank you.  Next, Ms. Kim Hoffman.

MS. HOFFMAN:  My name is Kim Hoffman.  I am a party to litigation, and my answer is no to the rest of the questions.

Ladies and gentlemen of the panel, I commend you for your due diligence in participating in this meeting and ensuring that your concerns and recommendations regarding the approval of saline breast implants are being handled appropriately.  I hope my testimony today will outline the importance of your continued diligence and careful scrutiny of information that is provided to you, and perhaps you will consider that important information has been withheld from you.  I am happy to provide you with supporting documents, including my testimony before the United States Congress last November on this issue.

I am here today to get several issues on record and to make the panel aware of information which may affect your opinion of the saline implants that were approved two years ago.  I am deeply concerned that a lot of essential information was withheld from the panel in 2000, and I question the accuracy of the data you did receive.

For example, Dr. Bobby Purkait, Mentor's director of research, testified to the panel about issues involving betadine use with implants.  If you review his testimony and compare it to information in the 1997 FDA 483 inspection report, you will find he misrepresented the facts to you.

Furthermore, at the time this panel reviewed Mentor's PMA in March 2000, Mentor was under an open criminal investigation for allegations of serious irregularities in breast implant studies and other issues involving the integrity of Mentor and their products.  Did the FDA make this panel aware of the investigation and the issues involved in the investigation?  The Chairman of the House Commerce Committee felt this important enough to write a letter to the Commissioner of FDA asking why the FDA would even proceed to panel under such circumstances.

The story broke in USA Today in March, 2000.  Mentor's lead counsel made a public statement denying the issue under investigation.  The false statement made by Mentor's representative caused extreme fluctuations in stock prices.

Given that Mentor lied to the public about the criminal investigation, and the investigation involves allegations of serious irregularities in breast implant studies, should this advisory panel trust the integrity of the data supplied from Mentor?  Why should the FDA trust the integrity of the data?  Especially since the data that was presented was problematic to begin with and because two FDA employees with knowledge of the issues involved in the criminal investigation recommended the application integrity policy.

Numerous informants have reported product defects and coverups by the company resulting in a criminal investigation that was opened nearly four years ago and remains open today.  All of these issues should have been resolved and the criminal investigation concluded before this panel was asked to review the PMA data from this company in 2000.  Since that didn't happen, I wanted to make sure you were notified about this today.

Many of the allegations that were raised in the criminal investigation have implications for the safety and efficacy of breast implants. For example, fraudulent manufacturing records subvert the ability of the FDA to review certain safety issues like platinum or polyurethane contamination, and fraudulent study data subverts the ability to review efficacy.

I would like to briefly address the issue of informed consent.  Unfortunately, patients have not been made aware of the numerous problems and violations of good manufacturing practices found during FDA inspections at Mentor, in addition to the violations alleged by industry insiders.  In fact, FDA has been denying the public's requests for these reports when requested under the Freedom of Information Act.  I hope the FDA has at least provided these inspection reports to this panel, and that in the future this panel will ask to review all FDA 483's and EIR's for a company as part of the PMA review process.  It is my understanding that Mentor has had yet another bad inspection report since the approval of saline breast implants.  Again, I hope you will ask the FDA for this information.

I would like to end by reading some quotes for a deposition of Mentor's former vice president of marketing.  Ms. Altavilla is talking about the difference in contracture rates between smooth and textured implants.  She states that textured implants significantly reduce the capsular contracture rate, from almost 60 percent to almost five percent.  However, she went on to say that the textured implants had another problem, they wrinkled 100 percent of the time with patients.  My understanding is that wrinkling leads to crease-fold failure.

Does the data supplied to you indicate a 60 percent contracture rate with the use of smooth implants?  In another unpublished study, done by Mentor's product evaluations department, it was found that Siltex or textured implants have a five times greater deflation rate than smooth implants.  Does the data you have received reflect that?  Do the consent forms?

Two years ago members of this advisory panel expressed concerns that patients need better informed consent.  Unfortunately, the panel itself did not have all the information they needed and neither do the patients.  I think the panel now has to question whether the information provided by the manufacturers can be trusted, and must ask the FDA what efforts have been made to check whether the data is accurate.  With all due respect, I urge you to carefully scrutinize all information received from Mentor Corporation.

At the last panel meeting, some members marvelled that women wanted breast implants despite their high complication rates.  Did the panel consider that women are not clearly informed about the high complication rates?  Plastic surgeons at the meeting repeatedly stated that the rupture rate was very low, despite clear research data to the contrary.

I hope you will review the material I am supplying in the form of written testimony, and I hope you will give serious consideration to the issues raised and the impact it will have on public health, not just for breast implant recipients but for their family members and the cost of medical care in this country.  Thank you.

DR. WHALEN:  Thank you.  Any questions for Ms. Hoffman?

[No response]

Thank you.  Each of the six individuals we have heard from thus far have been representing themselves.  We now proceed to various individuals who have identified that they are speaking on behalf of certain organizations.  The first is Ms. Margaret Volpe, from the Y-ME National Breast Cancer Organization.  Ms. Volpe?

MS. VOLPE:  As you said, I am representing Y-ME National Breast Cancer Organization.  Y-ME has received a minuscule amount of funding from manufacturers in the past, and the answer to the remainder of your questions is no.

Thank you for allowing me to present a statement to the advisory panel.  Y-ME is committed to providing support and accurate information to empower individuals touched by breast cancer so that they can select, in conjunction with their healthcare provider, the most appropriate options for themselves, including options for breast reconstruction.  Options give cancer patients some sense of control in restoring health and quality of life.  Saline breast implants are an important option for women who face breast cancer.

Y-ME welcomes the additional data from the manufacturers, and hopes it will put to rest the concerns many have expressed about breast implants.  I was diagnosed with breast cancer in 1995 and had a mastectomy and tissue expander.  I received an implant in February, 1996.  Reconstruction meant not worrying about how clothes fit; feeling whole again; not being constantly reminded of breast cancer and getting on with my life.

I have had no problems or complications with my implant since my surgery.  While I was eligible for TRAM reconstruction, I didn't want to endure the major abdominal surgery and painful recovery period.  I wanted to keep those muscles intact.  Many women do not have this option at all.  They are too thin to have the needed tissue for the TRAM reconstruction.  Even the latissimus dorsi or back-flap reconstruction usually requires an implant.

If implants were not an option we would be reminded daily of the mutilation to our breasts.  Each woman who has had a mastectomy must be allowed to pursue the best option for her, including breast implants.  At present, a woman who has had TRAM reconstruction on one breast is unable to have a second TRAM reconstruction at a later date if she should develop cancer in the other breast.  It is imperative that we continue to have a choice, and for many of us implants are the only choice we have.

Y-ME believes the availability of saline implants is very important to women who face breast cancer.  It is the only uncomplicated option left for women who desire an implant as part of their breast reconstruction after the FDA's 1992 restrictions.  It was very difficult for me to get the implant I received in 1996 because of FDA restrictions that required me to be in a clinical trial.

This panel must abide by the science when evaluating saline breast implants.  Do not allow yourself to get diverted and side-tracked by special interests.  The National Academy of Sciences Institute of Medicine report was issued, and the science is clear.  The IOM conducted an exhaustive and definitive review of all the existing research and found that there is no evidence that silicone breast implants cause cancer or disease.  This report also found the same results for saline breast implants.  U.S. Court's National Science Panel and several European government scientific panels, including the U.K.'s independent review group, issued similar findings.

Y-ME emphasizes the need for a wide range of treatment options as each woman must be able to choose the option that best fulfills her needs.  Y-ME worked with FDA to produce accurate information and uses the FDA and IOM breast implant information booklets when counseling women.  When it comes to the implant itself, women should understand that no medical device lasts forever.  Shunts, heart pacemakers, even artificial knees and joints have a limited life span and possible local complications.  Women should be aware of potential rupture and the need for replacement.

Adequate informed consent is a key part of the process.  Doctors should discuss risks and benefits in detail with their patients.  Saline implants do have a silicone shell, but from the exhaustive research on silicone implants, pointed out by the IOM, we also know that there is no convincing evidence that silicone produces an immunologic response.  The IOM states such diseases or conditions are no more common in women with breast implants than in women without breast implants.

In closing, I urge the committee to act based on the science alone.  Breast cancer is a devastating disease.  In the effort to resume our lives, breast cancer survivors have the right to select appropriate and effective medical therapies or devices.  Thank you very much.

DR. WHALEN:  Thank you.  Any questions of Ms. Volpe?

[No response]

Thank you, ma'am.

MS. VOLPE:  Thank you.

DR. WHALEN:  The next identified speaker was to have been Ms. Sybil Goldrich.  I understand it was by a videotape and that has not been able to be arranged.  Is Ms. Goldrich in the audience?  If not, next then, from the Sheridan Group is Ms. Cherien Dabis.

MS. MCDONOUGH:  Ms. Dabis couldn't be here today.

DR. WHALEN:  You are Ms. McDonough?

MS. MCDONOUGH:  I am.  So, I am going to do part of her testimony and mine, if it is okay with you, and try and cover it all.

DR. WHALEN:  It is just that each speaker gets five minutes.

MS. MCDONOUGH:  All right, I will do my best.  The answer to your questions are all no.  I actually am not representing any company, I am here on my own as well so the answer is no to all of those.

Good morning.  My name is Mary McDonough, and for ten years I played Erin on the TV show, "The Waltons."  After the show ended I chose to have breast implants because of the pressure I felt within my industry to look a certain way.

It is a decision I now greatly regret.  Before going through the surgery I did do my homework.  I talked to my friends who had had implants and I talked to my doctors, and I read the very little information that was available at the time.  Years later, I now realize that I made that decision based on misinformation and I have paid for that decision, like thousands of other women, who have faced serious complications, ultimately requiring the removal of a breast implant.  I had my implants in for ten years.

Like most Americans, I believed the FDA and that if a product was FDA approved or under a study, it meant every precaution had been taken to ensure the safety and effectiveness of the product.  This naivete cost me my health, if not nearly my life.  I have been diagnosed with lupus.  But I am not naive anymore.

Two years ago I sat in this very room and I learned that the process of the FDA that they use to approve products is fundamentally flawed.  I was shocked to learn that my testimony, given on a Wednesday, was irrelevant because the decision had been made on the previous Monday.  It was a mockery of the democratic process and, frankly, it made me very angry.  But my anger has now turned to education and that is my purpose for why I am here today because I want to share this with you as the gatekeepers of America's health and safety, what we have learned in the past two years.

I will raise questions about the data that was used by this committee to approve those products and I will point out gross violations of the FDA's own policies that allowed this product to move forward, and I will call into question the integrity of the manufacturers who produced these products.

American women count on the FDA and the advisory committees such as this one, and to this day the women have been poorly served on the issue of breast implants.  I hope that today we can start to rectify that.  Here are my three points:

First, saline implants were approved in 2000 under a cloud.  Mentor Corporation had a history of violations and warning letters regarding good manufacturing practices and medical device reporting, a history that has persisted through today.

McGhan Corporation also has a long history of violations of good manufacturing practice and medical device reporting.  Since 1992, the company has received numerous warning letters from the FDA for manufacturing violations that directly relate to implant safety.  This information is important because it casts doubt on the integrity and legitimacy of the manufacturer's practices, including their ability to conduct valid clinical trials and, since as an advisory panel, you solely rely on data produced by the manufacturers, their credibility and reliability should be critically important to you.  Yet, both of these manufacturers were approved in 2000 for saline implants without discussion of the history of the manufacturers.

The FDA has an application integrity policy.  This policy allows FDA to halt or suspend the approval process when there are significant questions regarding an applicant's data integrity, or if there are concerns regarding a manufacturer's practices.  By invoking this policy, it would have been difficult, if not impossible, for these manufacturers to bring their PMA applications for approval under the FDA guidelines.  I question, and perhaps you do too, why this policy was not invoked for these two manufacturers.

Second, the data regarding the reporting of adverse events is suspect.  I have learned that there are serious discrepancies in Medical Device Reporting, otherwise known as MDR, for both Mentor and McGhan.

According to Mentor's 2000 PMA clearance letter from the FDA, they were required to submit MDRs under the identifier P990075 as part of the conditions of approval and postmarket surveillance.  However, upon checking the reporting history, no MDR has been filed under that identifier.  Mentor has filed MDRs for saline breast implant under other identifiers, however.  In fact, the majority of MDRs filed by Mentor have been under identifier P940039, and that does not correspond to any approved PMA in the FDA system.

We have found similar discrepancies with McGhan's MDR reporting as well.  For example, since 1996 167 MDRs have been filed for McGhan under identifier K8810444446.  This identifier is not a McGhan 510(k) number.  It belongs to US Dental Corporation's Super Pik Massaging Pick for Oral Hygiene.  That was submitted to the FDA in March of 1988 and was cleared in May of 1988.

Under this system of surveillance, how are we, or more importantly you, the FDA advisors, ever to get accurate information on the complications and the failure rates?  Six months prior to the approval of this product, the manufacturers of these devices were given an exemption to report adverse events for saline and silicone implants not under the MDR system, but under the new alternative summary reporting system.  This exemption was made despite the FDA's own policy not to use this alternative for devices that were approved less than two years.  Exemptions like this are curious but they also produce dangerous results.

Looking at the data from the alternative summary reporting system during the years of 1999 to 2001, 34,356 adverse events were reported and filed with the FDA through the ASR system for Mentor and McGhan.  Of those, 30,290 are for saline implants along, 30,000 adverse consequences.  I can't help but think if these were penile implants this matter would have been resolved.

Finally, the postmarketing data on saline implants--

DR. WHALEN:  Ms. McDonough, you need to come to a conclusion, please.

MS. MCDONOUGH:  Well, I guess you are aware of the adverse events but I just wanted to tell you that there were 30,000 adverse events reported and it just doesn't measure up with this data, and I would like you to take a look at that.  It is so important for you to know so women can know all of the information.  If you could read Cherien's testimony, it is very compelling and I urge you to read it.  She was a 19-year old girl who got a saline implant for a defection which was a tumor in her chest.  Through all the manufacturers' information she made this choice and then had a ruptured implant which caused her great pain.  So, please take a look at that and please look at the data, and thank you very much for letting me talk today.

DR. WHALEN:  Thank you.  Are there any questions for Ms. McDonough?  Dr. Miller?

DR. MILLER:  Thank you for your testimony. What is the Sheridan Group?

MS. MCDONOUGH:  I am here on my own today.  Cherien is with the Sheridan Group.  The Sheridan Group is a lobbying firm.

DR. MILLER:  Thank you.

DR. WHALEN:  The next speaker is Ms.  Cynthia Pearson, executive director, National Women's Health Network.

MS. BARON:  Ms. Pearson couldn't be here today.  My name is  Tonia Baron and I will be reading a statement on her behalf.  I would like to start by saying that the answer to all four questions is no.

My name is Cynthia Pearson, and I am executive director of the National Women's Health Network, a non-profit, non-partisan organization that has been dedicated to improving women's health for more than 25 years.

The Network has been examining the safety of breast implants for more than a dozen years, and our primary concern has been the lack of safety information.  When we first became involved in this issue there were no studies of women in the published research literature.  Although almost a million women had breast implants, no breast implant had ever been approved by the FDA.

Today, there are quite a few published studies on the safety of breast implants and saline implants made by two manufacturers have been approved by the FDA.  Nevertheless, the Network remains very concerned because there have been more than 150,000 adverse reaction reports to the FDA for women with breast implants, and there are still no long-term safety studies.

This meeting provides the first opportunity to revisit the FDA's approval of saline breast implants since they were approved in 2000.  The manufacturer studies were strongly criticized by the FDA's advisory committee two years ago because of the poor quality of their recommendation.  One might expect that the manufacturers would have been grateful that their implants were approved despite the high complication rates and poor quality of the data and would have, therefore, made sure that their five-year studies were better designed and analyzed.

Instead, the new studies have many of the same flaws as the previous studies, and the response rate is even worse than it was at the PMA meeting in 2000.  I will not go into details about the statistics because that does not seem necessary.  Anyone who knows anything about research, and many who even know nothing about research, know that you can't lose 95 percent of your sample and still have a meaningful study.

When Mentor Corporation analyzed their five-year follow-up data on only 60 of the more than 1200 women who were enrolled in the augmentation study, they did a disservice to more than 400,000 women who underwent augmentation surgeries since saline implants were approved in 2000, to you, the FDA advisory committee members and to the FDA.  There is no excuse for that kind of shoddy research.  In fact, it does not deserve to be called research.  It is meaningless data.

The fact that Mentor improved their follow-up to 24 percent at the six-year mark shows how little attention went into their five-year study.  But even a 24 percent response rate is much too low to be meaningful.  A response rate under 50 percent raises more questions than it can answer.  Are the other women dead or alive?  Are they healthy and happy or sick and seeking medical care elsewhere?  We don't know the answers to these questions so the data can't really tell us about the safety of implants.

The McGhan data are a little better.  Like the Mentor study, they manipulate the data to make it seem that saline implants are safer than they really are.  Both Mentor and McGhan made note of the women who had both their implants removed, but they are not included in the complication rate data.  Instead, they are apparently excluded from the study as if they never existed or, worse, as if they are satisfied with their implants or have the same complication rate as everyone else.

Even with this highly inappropriate data manipulation, approximately one out of every three McGhan augmentation patients are described as undergoing additional surgical procedures, an average of more than two additional surgeries per patient.

Consumers deserve an FDA that keeps unsafe or ineffective medical devices off the market.  They also deserve accurate information about the long-term safety of medical implants that are intended for long-term use.  In the case of breast implants, the manufacturers have failed to conduct meaningful long-term safety studies and, as a result, consumers are continuing to buy implants that are FDA approved but are not necessarily safe.

All we do know is that breast implants have an extremely high complication rate, higher than any medical product I can think of with the exception of a small number of life-saving products used by patients who have no other choices.  Even the misrepresentation of data does not hide that fact.

On behalf of the National Women's Health Network, I strongly urge this advisory committee to speak on behalf of the millions of consumers who are not able to come to this meeting but who look to the FDA to protect them from harmful products.  If this advisory committee does not vehemently criticize these studies and the manufacturers' misrepresentation of their own research and urge the FDA to take a stronger watchdog role, breast implants will never be improved and accurate and informative research will probably never be conducted.

DR. WHALEN:  Thank you.  Are there any questions?

[No response]

Thank you.  Next, we will hear from Dr. V. Leroy Young, from the American Society of Plastic Surgeons.  Dr. Young?

DR. YOUNG:  In answer to the questions, I am a practicing plastic surgeon and breast implants are part of my research and part of my income.  I also receive funding from several of the manufacturers and have served as a consultant for several of the manufacturers, and my expenses to travel here were paid for by the American Society of Plastic Surgeons.  I am not involved in any litigation.

What I would like to do with the time that I have is present work that the American Society of Plastic Surgeons and the work of the Plastic Surgery Educational Foundation has either conducted or funded to try to understand the performance of breast implants and the risks and benefits that patients derive from them.

I have provided a written copy of the slides which we have here, and I will tell you that we have a lot more information available that we would be glad to share with you but, because of the limited time, we couldn't put all of that into a presentation.  I apologize that this is taking so long.

DR. WHALEN:  You might want to start with your remarks because the clock is ticking away.

DR. YOUNG:  Well, what I want to emphasize about this is that our organizations are primarily interested in patient advocacy, research and education.  We have made a real effort to improve patient advocacy in the last two years by initiating a national breast implant registry.  This was done by Plastic Surgery Educational Foundation because we felt that it was the right thing to do for patients.  We also felt, because there have been lots of concerns, some of which we have heard today about the quality of research, that there should be independent data that is not sponsored by the manufacturers available for patients and for surgeons.  This registry provides a source of that data, and it provides it in an ongoing, timely fashion and we publish it on a regular basis.

This registry, as I said, was founded in July of 2000.  At the present time we have almost 4000 patients enrolled in it.  From this, we are learning things about the incisions that are used to insert the implants, the type of implants that are being used, the position they are placed in but, more importantly, I think we are learning a lot about why reoperations occur and how often they occur.  From that, we are implementing steps to try to decrease the reoperation rate.

In May of this year we also began an international registry that includes most of the countries in Europe and South America, and we feel that this is important because this is going to let us accrue data more rapidly.  It will let us make comparisons of different groups of surgeons, and it will let us learn about different types of implants.

We are also working on informed consent, and we are working on that in two ways, by administering a survey in surgeons' offices and by posting the survey on the implant info web site.  The beauty of this is that it allows women to comment either positively or negatively about their results and they are not under the influence of the surgeon's office when they do that.  It is a detailed questionnaire that goes over preoperative, postoperative and pre-operation data.  From this, just as examples, we are learning how the size of the implant was chosen and what we have seen from that is that today's patients are much better educated.  They are taking an active role in their medical care.  As you can see from the slides I provided you, the surgeon and the patient decide together about size in  81 percent of the cases.

We are also trying to survey how responsive surgeons are to patients.  From that, we have provided you data that 79 percent of the surgeons prepared the patient for what to expect, and 83 percent took time to answer the questions the patient was interested in.

We are also trying to assess did surgeons inform the patients about the potential risks.  We have heard criticism of that today, and you can see that we are getting good data on that and that the majority of surgeons are apparently informing patients of most of the risk, but we need to work harder to get this to be 100 percent rather than 70-80 percent.

We are also getting data on the impact on life and we can see that there are several positive impacts from that.  One of the big issues at the last panel meeting was related to pain.  One of the things we have done is show, in a survey, that a significant number of women who don't have implants have pain, approximately 48 percent.  We have also looked at the pattern of pain and how implants affect it.  From this survey, it appears that implants don't affect the pattern of pain but in a small percentage, 13-15 percent, it is more severe and it occurs more frequently.  So, breast implants do have an effect on that.

Another issue was nursing.  We have seen from the survey that about equal numbers of patients had children and had a history of nursing before and after augmentation.  The problems that occurred with nursing were similar in the two groups of patients but they did not appear to be any worse in the patients who nursed after augmentation. So, we are getting a lot of useful information from this.

One of our main concerns is the report in the literature that one in five women who has an augmentation has a second operation within five years.  We are trying to decrease that and learn if that is really true.  We have done one study, which we have just completed, which shows that if you apply the principles that we have learned as a result of these studies that I am mentioning, we can at least cut this rate to ten percent, and we think we can decrease it even further.

Furthermore, we are funding research all over the world.  We have two studies in Finland now on rupture and local complication, and we are looking at the frequency of etiology of reoperations, as I said.  We have just completed a betadine implant related study, and we have funded a Center for Device Retrieval Analysis which focuses on breast implants.

PSEF and ASPS are committed to the patients and we are committed to understanding these things, and I will be glad to try to answer any questions you may have.

DR. WHALEN:  Thank you, Dr. Young.  Is PSEF separately incorporated?

DR. YOUNG:  Yes.

DR. WHALEN:  Does it receive any grants or funding from the manufacturers of saline breast implants?

DR. YOUNG:  It has in the past, yes, indirect funding.

DR. WHALEN:  Do you have any idea what the dollar amount of that is?

DR. YOUNG:  Early in the '90's I think there was altogether about six million dollars.  That amount has remained constant over the past ten years or so, other than the funding that we received for the betadine study which was about $30,000.

DR. WHALEN:  Thank you.  Are there other questions for Dr. Young?  Yes, Dr. Dubler?

DR. DUBLER:  Thank you very much for your presentation.  I have a question.  Do you and members of your organization advertise about breast enhancement?

DR. YOUNG:  Yes.

DR. DUBLER:  You do?  And what is the goal of that advertisement?

DR. YOUNG:  I think in most instances it is to encourage people to inquire about it.

DR. DUBLER:  To inquire about it, and maybe to go ahead?

DR. YOUNG:  If it is appropriate for them, yes.

DR. DUBLER:  And, what would you say the goal of an informed consent process is?

DR. YOUNG:  To educate the patient sufficiently that they understand the risks and the benefits of the procedure or treatment so that they can make a decision that is appropriate for them.

DR. DUBLER:  So, that sounds like an even-handed process.

DR. YOUNG:  Yes.

DR. DUBLER:  Is advertising and marketing an even-handed process?

DR. YOUNG:  No.

DR. DUBLER:  Can a person who does one do the other equally well?

DR. YOUNG:  I believe so.

DR. DUBLER:  Thank you.

DR. WHALEN:  Other questions of Dr. Young?

[No response]

Thank you.  Next, we are to hear from Dr. Laurie Casas, from the American Society for Aesthetic Plastic Surgery.  Dr. Casas?

DR. CASAS:  Thank you.  First I will answer your four questions.  I am a board-certified plastic surgeon and have been in clinical practice for 12 years.  I have managed hundreds of patients with breast implants and, therefore, a portion of my income is derived from breast implant surgery.  My travel expenses to this hearing were paid for by the American Society for Aesthetic Plastic Surgery.  I am currently chair of the Society's communications commission.  I have no financial ties to any implant manufacturer.  I am neither a witness nor a party in a pending lawsuit related to breast implants.

I am here today on behalf of the many thousands of patients who tell us that breast implants have made a positive difference in their lives.  I am associate professor of surgery at Northwestern University Medical School.  I have co-authored a book on breast surgery, and am currently a participant in a multi-site investigation of patient satisfaction following cosmetic procedures.

Given the proven level of safety and efficacy of breast implants, the American Society for Aesthetic Plastic surgery is in full support of the FDA's approval of these devices, along with the FDA's continued oversight.  Members of ASAPS, who are all board-certified plastic surgeons, welcome this oversight for two reasons:

First, to ensure that the collection of data on breast implants is ongoing and, second, that women considering the procedure have the benefit of the most current information.

Part of ASAP's mission is to provide accurate information to the public. We do this directly through our web site, and also through our members via patient education and informed consent.

Informed consent occurs when patients are provided with all the facts and information necessary to make an educated decision to proceed with a medical treatment or surgical procedure.  Full informed consent is in the best interests of both patients and physicians.

ASAP's members are patient advocates.  We feel that patient education, safety and satisfaction are of primary importance.  We believe strongly that a woman's right to choose breast implants is paralleled by her right to be fully informed about the risks and the benefits.

The Society offers continuing education to its members, providing information to assist them in evaluating and judging the appropriateness of their patients' motivations and expectations.  Therefore, we believe that patients with inappropriate motivation or expectations should be counseled against surgery.

There is a growing body of evidence that suggests that cosmetic surgeries, such as breast augmentation, lead to improvement in at least three areas of psychological functioning: body image, quality of life and depressive systems.

The research arm of the Society, the Aesthetic Surgery Education and Research Foundation, has recently funded a multi-site outcome study on patient satisfaction following aesthetic surgery.  This study will provide additional important data on that critical subject.

An increasing number of women today are choosing breast augmentation to enhance their appearance, over a million women in the past decade alone.  Most women seeking breast augmentation come to plastic surgeons for two principal reasons.

The first is a woman who finds that her once normal breasts have lost considerable volume following pregnancy and lactation and are not in proportion to the rest of her body.  The second reason a woman comes to us for breast augmentation is if she has grown to adulthood with no breast development.  She feels her breasts are out of proportion.

The change in how she feels about herself after surgery is dramatic.  I have seen lives transformed by breast augmentation.  The typical breast augmentation is usually a woman in her 30's, most often she is married and has children.  She is a responsible adult with a family, career and normal life.

Women have to overcome tremendous obstacles to have breast augmentation, from societal prejudice to a wealth of misinformation.  The fact that so many seek it out speaks to its strong desire.  Satisfied patients have created the popularity of this procedure, and research has shown that the vast majority of women who have breast augmentation would make the same choice again.

The high satisfaction rate, and the determination of so many women to undergo a surgery with the knowledge that it is not a perfect operation, suggests just how deeply the benefits are felt.

Members of ASAPS believe it is our responsibility as health professionals to provide accurate and up-to-date information by which a patient can exercise her right to informed consent.  We welcome the opportunity to work with the FDA and within the FDA guidelines to achieve this goal.

The American Society for Aesthetic Plastic Surgery, and its 2800 board-certified plastic surgeon members and candidates, are committed to this process.

DR. WHALEN:  Thank you.  Any questions for Dr. Casas?  Dr. Dubler?

DR. DUBLER:  Thank you very much for your testimony.  Do you advertise?

DR. CASAS:  No.

DR. DUBLER:  Do your colleagues in your organization advertise?

DR. CASAS:  Yes.

DR. DUBLER:  Why don't you?

DR. CASAS:  I have never had the need to.  I have a very busy practice and for me, personally, the value of spending the money to market is not obvious to me; I don't see a value.

DR. DUBLER:  Are there any commonly done plastic surgery interventions which you don't do because you think they are not for the benefit of patients?

DR. CASAS:  You have to be more specific.  I am not sure what you are asking.

DR. DUBLER:  Liposuction?

DR. CASAS:  Could you rephrase the question?

DR. DUBLER:  Are there any commonly done plastic surgery interventions which are not part of your common practice because you think they are not for the benefit of patients?

DR. CASAS:  Interventions typically mean you are intervening with a past procedure.

DR. DUBLER:  I am sorry, surgical interventions.

DR. CASAS:  Well, surgical procedures, are there any surgical procedures that I don't perform?

DR. DUBLER:  Correct.

DR. CASAS:  Why?

DR. DUBLER:  Because you think they generally don't help patients, are not in the interest of most patients; the data don't support their effectiveness.

DR. CASAS:  You are asking a question that can't really be answered.  I mean, there are thousands of different procedures that we can perform as plastic surgeons to mold and reshape the body, and we don't have a list.  You know, there are pros and cons for every different procedure and you individualize with each patient.  So, it is not like general surgery where you have hernia repair, gallbladder--you know, a list.  We have thousands of procedures that we perform, small, large, somewhere in between, and we individualize.

DR. DUBLER:  You said that patients with inappropriate expectations are discouraged from having surgery.  Why?

DR. CASAS:  Because the most important interface that we have with the patient is the process of informed consent.  So, it is our job and our duty to prepare patients for procedures.  The interview process is very straightforward.  The patient presents with a particular body part they are not satisfied with.  We talk to them about their motivation, their present health status, and then provide them with alternatives.  On that list of alternatives is not to do a procedure, and that is part of the informed consent.  So, if in the process of informed consent, as a physician, we feel that the motivation is inappropriate, their health status is not adequate, we would counsel them against the procedure.

DR. DUBLER:  Are you aware that some of the more recent studies on informed consent, the empirical studies, show that patients in general think that the informed consent process is to protect the doctor and the institution, not really to inform them?

DR. CASAS:  I am not aware of those studies.

DR. DUBLER:  Thank you.

DR. CHANG:  Dr. Casas, I have a question that may be a little difficult to answer, but if in the informed consent process, let's say in the patient brochure, or if every plastic surgeon doing the counseling, came right up front and said early data from manufacturers suggests that three to five percent of these implants may fail within three years, or that one out of five patients may require a reoperation, in your opinion would the rate of breast augmentation stay the same, or might it go down?

DR. CASAS:  It is an interesting question because those are the statistics I give patients and they still want breast augmentation and breast reconstruction with implants.  I think women are looking for a choice in changing their body structure, and for the breast we either have implants or your own tissue, or a combination of the two.  So, we are limited by the choices.  So, I think patients evaluate always the pros and cons of all their choices, at least in my office and my practice, and I think ASAPS does that and teaches that in a very, very straightforward way.  There is absolutely no opportunity in our minds that a patient chooses an operation without full informed consent.  It is critical in the process.

DR. CHANG:  Thank you.

DR. WHALEN:  Thank you.  The next two speakers that have been identified to us are both from WIN Against Breast Cancer, Ms. Mullen and Ms. Crigler.  As I understand it, Ms. Mullen is not going to be present and Mr. Crigler was to read her statement.  What I am going to ask is that Ms. Crigler go ahead with her own remarks at this juncture and then, at a later point in time I will call her back, as time allows, to read Ms. Mullen's statement.  Ms. Crigler?

MS. CRIGLER:  Good morning.  My name is Sheila Crigler, and I am truly blessed to be here today to present my perspective on saline-filled breast implants.  I am in no way receiving financial compensation from any implant manufacturer, nor do I have a pending lawsuit.

I am here as an official person for Women's Information Network Against Breast Cancer, the acronym being WIN.  It is a national non-profit organization, founded by president and CEO Betsy Mullen.  Other hats I wear, I co-facilitate Women Reaching Women, a support group for breast cancer survivors.  Not only do I hope to speak as an advocate, but mainly as a breast cancer survivor with saline-filled breast implants.

Most of us have heard that life begins at 50.  A few weeks after turning 50, I received a telephone call from my surgeon at that time regarding my recent mammogram.  To him, it was a routine telephone call, but telling someone you have breast cancer over the telephone is anything but to the person on the other end.  Life begins.

There were tons of information to plow through; decisions to be made.  Would I live or die?  Intraductal carcinoma in situ, the medical jargon; milk ducts clogged with cancer, the layman's jargon.  My only option was a complete mastectomy.  A woman never considers the possibility of losing a breast--more decisions and possible choices.  A realization was that breasts do not define Sheila but breasts do give Sheila definition.

A consultation with a plastic surgeon brought sunshine to a dim horizon.  I was a good candidate for reconstructive surgery with saline-filled implants.  More choices.  For symmetry, I opted for the removal of the remaining healthy breast.  This was a very personal decision and it is not for everyone.  If I could prevent the possibility or probability of breast cancer in the remaining breast, it was worth it to me to give up my other breast.

The operation was a success, and the real blessing came when it was learned there wasn't any involvement in my lymph nodes.  That meant no chemotherapy or radiation.  My breasts were gone but I got to keep my hair.

It was over a year from the cancer surgery before reconstruction would begin.  I was amazed to learn some survivors cannot look at themselves in the mirror after surgery.  Imagine a billboard, as I called it, in place of mounds that once defined womanhood.

There are different types of reconstructive surgery, with saline-filled breast implants being one of the least complicated in regards to length of surgery and recovery.  If a rupture were to occur, the body will absorb the saline and release it out of the body.  The lack of choices is what creates problems.

Women need the opportunity to make informed decisions.  There is no one particular thing that works for everybody.  Choices again.  A few with strong issues against a situation should never be the deciding factor for the masses.  The need is to inform consumers and not remove choices.

My reconstructive surgery was a success and my saline-filled breast implants are truly a part of my body by now.  It has been almost six years since my permanent implants were put in place.  That has created a completion of a circle.  I began with breasts and have ended with breasts.  Now I can smile at a newly diagnosed patient and show them that this is the face of cancer.  I can proudly display my saline-filled breast implants that say that there is hope.

It would definitely be a travesty if this particular choice were removed.  Hope is what a person does not have when she tells you she would rather die than have her breasts removed.  We did not choose to have breast cancer but we can choose to have reconstruction.  A woman's right to choose is what is at stake here today.  I feel so strongly about this issue that I left my mother, who is hospitalized in Los Angeles, to make the trip here to express my concerns regarding the ongoing study of saline-filled breast implants.  Thank you.  Are there any questions?

DR. WHALEN:  Thank you, Ms. Crigler.

MS. CRIGLER:  Thank you.

DR. WHALEN:  Next is Dr. Diana Zuckerman, from the National Center for Policy Research for Women and Families.

DR. ZUCKERMAN:  Thank you.  The answer to the four questions is no, and I donated my time to be here today and traveled in my own car.

My name is Dr. Diana Zuckerman, and I am president of the National Center for Policy Research for Women and Families, which is a non-profit, research-based organization that explains medical and scientific information so that it can be used to improve the health and well being of women, children and families.

Our research center works on a wide range of health issues, with particular attention to the safety of medical products.  Unfortunately, as the new research on hormonal replacement therapy has reminded us, manufacturers' claims about medical products are not always supported by research.  Our goal is to balance the hype by scrutinizing the research and determining the facts, whatever they might be.

In the case of breast implants, the more than 150,000 adverse reactions that have been reported to the FDA are inconsistent with the manufacturers' claims that implants are very safe and that implant patients are so satisfied that even when their implants break all they want to do is replace them with a larger size.  So, we have used our expertise in epidemiology, biostatistics and public health to carefully scrutinize the research.

I was at the FDA advisory committee meeting when saline implants were approved in 2000.  For those of you who weren't here, I want you to know that the advisory committee expressed a great deal of concern about the extremely poor quality of the research and about the exceedingly high complication rates.  And, I was here to hear the shocked gasps in the audience when the advisory committee voted to recommend approval despite their very strongly expressed concerns.

The advisory committee recommended approval with the caveat that long-term safety data and better studies be required of the manufacturers.  I am sorry that many of those members are not here today because I think they would be very disappointed that the manufacturers made many of the same mistakes or misrepresentations in their data this time that they were criticized for two years ago.  In fact, as a former college professor, and I used to teach research methods and statistics courses, I would have to say I would flunk anybody who would provide a statistical analysis like some of the ones that I saw in these studies.

I want to use the words of former members of this advisory committee as well as other statisticians.  Since we haven't had the opportunity to hear the data yet this morning, having our public comment before it is presented, I wanted to use some of their own words to criticize these studies that you are going to be discussing today.  For example, Dr. Brent Blumenstein, who was the statistician on the advisory committee last time, clearly stated that the McGhan presentations did not meet the standards of "good, peer-reviewed articles" and specifically stated that--I love this one--"accuracy is not manifest in the presentation of the data."

The FDA statistician, Telba Irony, specifically criticized using Kaplan-Meier analyses when so many women are lost to follow-up and the researchers don't know if those who dropped out are similar to those who are in the study.

Both these statisticians criticized the large number of patients who were lost to follow-up in the studies two years ago.  Of course, the new data presented today show even more women lost to follow-up.

In the Mentor study two years ago, Dr. Blumenstein criticized the company for taking patients out of the analysis if they had their implants removed.  He said, "this is an extreme limitation and misrepresents the data."  He then concluded by saying, "I cannot accept the accuracy of any of the data here because of the limitations I'm pointing out.  I cannot feel good about any of the data presented with respect to accuracy and giving that information to an individual patient and having that patient understand what the real risks are."

Phyllis Silverman, who was the FDA statistician two years ago, stated that, "because of the approximately 50 percent loss to follow-up with the large, simple trial, the ability to draw meaningful conclusions from this trial is limited."  I fully agree with those statements, and it is obviously even worse for the new Mentor data which actually had 95 percent--95 percent loss to follow-up at five years, and 76 percent loss to follow-up at six years.  Even if the rates continue to get a little higher, and I see that Mentor has made some efforts to do so, a response rate under 50 percent cannot provide useful safety data.

DR. WHALEN:  Dr. Zuckerman, can you come to conclusion?

DR. ZUCKERMAN:  Yes, I will speed up but I guess I would respectfully request that I get as much time--

DR. WHALEN:  I am sorry, we need to have you conclude.  Everybody gets their five minutes.

DR. ZUCKERMAN:  Yes.  I will just address what Dr. Dubler said, wondering why the manufacturers were unable to afford good researchers to do the research when they have, in fact, spent millions of dollars to advertise in every major woman's magazine, including those read by millions of teenagers.

I have here a letter that was written by two congressmen yesterday, asking that members of this committee be given information about two things, the National Cancer Institute studies on breast implants and I happen to be on the advisory committee for those NCI studies, and also ask that you be informed of the criminal investigation and problems with inspection reports.

So, I would just turn to you and say, and I will give you a copy of that letter, as an advisory committee, if you don't insist on better long-term research and on accurate reporting of that research, teenagers and women will continue to incorrectly assume that breast implants are proven safe for long-term use.  I am not questioning women's rights to make choices about their own bodies, but I am questioning whether they can make an informed decision when there is no really well-conducted safety research.  Are there any questions?

DR. WHALEN:  Dr. Dubler?

DR. DUBLER:  Have you reviewed the materials which surgeons use to discuss with patients whether to go ahead with implant surgery?

DR. ZUCKERMAN:  Yes, and of course they vary a great deal.  One of the problems, as you mentioned earlier, is that so many patients think that even very clear warnings are really meant for liability reasons, not to protect them, not to provide information to them.  But we also know that sometimes doctors give a written informed consent but what they say to the patient is, you know, of course, I wouldn't be doing this if I didn't think it was safe for you; my experience is that patients love their implants and most of my patients are very satisfied.  So, I think informed consent, as we know, is a written version and then there is the oral version and they can be quite different.

DR. DUBLER:  Do you think it would be possible, based on your reading of the data that have been produced, to devise an independent, interactive, video-based informed consent process that really permitted a woman to scrutinize the data from the perspective of the skeptical, not the supportive?  Would that be a possible process to design?

DR. ZUCKERMAN:  I think so.  I mean, we know that there are limitations of informed consent but we have to do a better job than we are doing now.  I think that your suggestion is an excellent one because I think if there was an independent one, particularly if it was a videotape, interesting and not these terribly long, boring informed consent things that many people don't read or don't understand or don't know what to do with, I think that could really make a difference.

I have no doubt that plastic surgeons do warn their patients about some of the risks, and I guess my concern is that if they are warning about capsular contracture, but they if are not mentioning last year's National Cancer Institute study showing an increased rate of cancers and deaths from certain cancers and respiratory diseases among women with saline and silicone breast implants, then I really question whether they are getting informed consent currently.

DR. DUBLER:  Thank you.

DR. WHALEN:  Thank you.

DR. ZUCKERMAN:  Yes.

DR. WHALEN:  Next is Dr. Jae Hong Lee, also from the National Center for Policy Research for Women and Families.

DR. LEE:  Good morning.  I am Dr. Jae Hong Lee, a physician and the senior medical policy analyst at the National Center for Policy Research for Women and Families.  My response to all four questions is no.

The Institute of Medicine report on breast implant safety identified reoperations and local complications as the primary safety issue with silicone breast implants.

In light of the serious concern regarding reoperations and local complications, the post-approval studies recently released by Inamed and Mentor are extremely disappointing.  Both studies are poorly designed and executed, and most likely underestimate the cumulative incidence of complications.

I will discuss just a few of the more glaring weaknesses of both studies.  A major problem shared by both the Inamed and Mentor studies is the exclusion of patients who have had all their breast implants removed prior to the three- or five-year study interval.  One can argue that these patients, among the earliest to have all their implants removed, are the most important group of patients to study.  They certainly should be counted as women with complications.

Excluding data from these patients introduces an unacceptable post-entry exclusion bias into both studies.  Since it seems likely that these patients were experiencing complications rates higher than those who did not have all their implants removed, the overall effect of this exclusion bias will be to seriously underestimate the cumulative incidence of complications.

Exclusion of these patients also invalidates the Kaplan-Meier risk estimates reported in both studies.  One essential condition for using the Kaplan-Meier method is the assumption of independence between censoring and outcome.  In other words, one must assume that the rate of local complications and reoperations in the excluded patients was similar to those who remained in the study.  Since the excluded patients had all of their initial implants removed for one reason or another, most likely due to complications, this is not a reasonable assumption to make.  Using the Kaplan-Meier method in this situation clearly results in an underestimation of the cumulative complication rate, a fundamental flaw in both studies.

The problem is compounded when the reports continue to cite the full number of enrolled patients when discussing specific complications.  For example, the Inamed study enrolled 237 reconstruction patients but excluded 97 patients by year five, for an actual follow-up of only 140 patients.  Inamed then reports reoperations in 100 of the 237 patients over the five years of follow-up.  The question is whey were the excluded patients counted in the follow-up cohort?  Is the percentage of patients with reoperations 100 out of 237, or 42 percent, or is it actually 100 out of 140?  	In contrast, Inamed conveniently excludes the breast implant removal patients when calculating their follow-up rate for reconstruction patients.  They calculated an expected follow-up cohort of 175 by excluding 11 patients who had died and 51 patients who had all their initial implants removed.  As a result, their follow-up rate of reconstruction patients was reported as 80 percent, or 140 out of 175 patients.

As mentioned earlier, it is highly inappropriate to exclude those implant-removal patients from the study cohort.  If those patients are left in, the expected follow-up cohort consists of 226 patients, not 175, and the actual follow-up rate for reconstruction patients is then a clearly inadequate 62 percent.

There are also signs that the Inamed study is statistically underpowered. For example, Inamed suggests that because most of the confidence intervals between five-year and three-year risk rates for complications overlap, the true rates may not be higher at five years.  They imply that it may not be necessary to follow complications with overlapping three- and five-year confidence intervals in subsequent years.

A more reasonable interpretation of this data is that the study is too small and underpowered.  Since even a small difference of three percent represents over 6000 breast augmentation patients, the post-approval studies must be large enough to detect even small differences in risk rates.

As bad as the Inamed study is, the Mentor study is even worse.  The Mentor report acknowledges that low response rate and the differences between the responders and the non-responders is a major limitation of their study.  I would go further and state that most of the data presented in the Mentor study is scientifically worthless due to the disturbingly low follow-up rates.

Setting aside for the moment the very serious methodological problems of both studies, we find little comfort in the reported data.  Even taken at face value, the complication rates are unacceptably high.  For instance, Inamed reports five-year reoperation rates of 26 percent for augmentation patients and nearly 45 percent for reconstruction patients.  As a physician, I find these complication rates to be completely unjustifiable for a purely elective cosmetic procedure.

The poor quality of the studies clearly indicates that both Inamed and Mentor have failed to fulfill all the conditions of their premarket approval agreement.  Furthermore, the reported complication rates in both studies reinforce the Institute of Medicine's concern about local complications and reoperations.  These companies seemingly lack the motivation to perform adequate and appropriate safety studies as long as they can market and sell their products as FDA approved.

So, what should be done?  At the very least, the FDA should require a black box warning on all advertisements noting that the manufacturer has not adequately completed required five-year safety studies.  The panel should also recommend that patient consent forms include a statement noting the lack of reliable long-term safety data.  Another option would be to temporarily suspend general distribution of these implants until better safety data become available.

The Inamed and Mentor studies leave many questions unanswered, but do make clear one important fact, right now, today, there are no adequate data demonstrating the safety of saline breast implants.  Thank you.

DR. WHALEN:  Any questions of Dr. Lee?

[No response]

Thank you.  Next, we have Dr. Susan Pope Helman on behalf of herself.

DR. HELMAN:  Those handouts are a little more detailed than what I am going to say.  I am shaking, sorry.  Good morning.  I am Susan Helman.  I live in Vero Beach, Florida.  My reason for being here is to enlighten the FDA's General and Plastic Surgery Devices Panel regarding the use of non-organic platinum in the manufacturing of saline and silicone medical devices and how I have been affected.

I will be brief.

DR. WHALEN:  I am sorry, I don't mean to exacerbate--

MS. HELMAN:  I am sorry, the answer to all those questions is no.

DR. WHALEN:  Thank you.

MS. HELMAN:  Your handouts include the specific terms and means by which my body specimens were tested.  I received breast implants in 1990 and had them removed in 1992, less than two years.  Nine years later, in May of 2001, samples of my hair, fingernails, toenails, urine, blood, sweat, skin cells and my left axilla lymph node were obtained by ExperTox Forensic Toxicology Laboratories in Houston, Texas for analysis.

On August 15th, 2001, a bone marrow biopsy was obtained and sent out to ExperTox for during a surgical procedure to remove a residual silicone capsule remaining in my left breast and lymph nodes in both breasts, thoracic area, neck and under my right arm.  Samples of these tissues were sent to the University of Florida Diagnostic Laboratories.  The findings from the University of Florida were, one, foreign material through all layers of specimens; two, residual silicone adhesive; three, foreign body giant cells and vacuolated histocytes containing silicone and polyurethane foam; and, four, silicone granuloma.

The results from ExperTox Forensic Laboratories are as follows: whole blood platinum 1542 pmo 1/L.  I don't know what these things mean.  I just know that the level of platinum in a human body should be no more than zero.  So, I am just going to give you the numbers, if that is okay.  Whole blood platinum, 1542; urine platinum 0.38; hair platinum, 1.59; nail platinum, 2.88; sweat platinum 3.85; bone marrow platinum 181.6 upper ug/L.  The oxidation states of the excised nodule under the right arm were 2+, equaling 42 percent and 4+ equaling 58 percent.  I just had another bone marrow biopsy taken two weeks ago.  I got the results yesterday and the number is 325.  So, it has gone up from 185.6 to 325.

In a comparison between this form of platinum and the known platinum containing chemotherapeutic agent cisplatin, it is known that cisplatin has a cytotoxic and neurologic action on the body and is an organic form of a platinum molecule whose properties are less toxic than the free form found in my body.

Skin cells from my cheek were also taken for analysis and were exposed to the platinum taken out of my body, my own skin cells were exposed to platinum taken from my body and the results obtained clearly demonstrated acute, significant toxic changes such as vacuolation and cell degeneration, and conclusively established that I have sustained chronic toxic cellular injury with platinum presently disseminated through my entire fluids and tissues.

I can only speak for myself, but I am an example.  This is how I have been affected and how my body has been affected.  My body is basically killing itself.  There are many women, breast implanted women like me in these United States with debilitating illnesses and extreme challenges too numerous to even list, but what we would ask is that the FDA panel be thorough in your understanding of the manufacture and use of saline breast implants since it is my understanding after all this that the bags used in these devices contain platinum.  Thank you so much.

DR. WHALEN:  Dr. Helman, it has been ten years since your implants were removed?  Is that correct?

DR. HELMAN:  That is correct.

DR. WHALEN:  And, you said in your testimony that at two recent points in time there was a rather dramatic increase?

DR. HELMAN:  That is correct.

DR. WHALEN:  Has any healthcare provider that you have interacted with suggested, in view of that data ten years after these were removed, that some alternative source of platinum than the breast implants might be affecting you?

DR. HELMAN:  They said only if I worked in the platinum jewelry manufacturing industry or as a motorway, like a turnpike person that takes money on the turnpike, or pump gasoline would I even have a trace of platinum in my tissues.

DR. WHALEN:  But, I mean, a dramatic increase at two points in time.

DR. HELMAN:  I asked the toxicologist about it yesterday morning when he gave me my new numbers, and he said it is possible that the platinum molecules relocated after the removal of the residual capsule that was left in my body.  He said, because the original bone marrow biopsy was done on August 15 of last year and then, of course, just two weeks ago, he said it may have relocated.  But I am going to be tested again in six months.

DR. WHALEN:  May I ask what your doctorate is in?

DR. HELMAN:  Metaphysics.

DR. WHALEN:  Are there other questions?  Dr. Doyle?

DR. DOYLE:  The question of timing, you said the residual implant was removed when?

DR. HELMAN:  Last August.

DR. DOYLE:  Thank you.

DR. HELMAN:  Yes, ma'am.

DR. WHALEN:  Dr. Miller?

DR. MILLER:  Thanks for your testimony.  Can you tell me about ExperTox?  How did you select ExperTox in Houston as a place to have your toxicology done?

DR. HELMAN:  I read a paper that was produced by Dr. Lakissa, along with a few other physicians, back in '99 or 2000 and I phoned him and asked him if he would consider doing a toxicological study on me because I was sick.  I was really sick and I didn't know why.

DR. MILLER:  Are services like this not readily available where you live?

DR. HELMAN:  No, they aren't.  It is my understanding that mass spectroscopy--I have trouble pronouncing these things--and special equipment is not readily available.  I think there are a few pieces of this equipment throughout the country but they are kind of sparsely located.

DR. MILLER:  Thank you.

DR. WHALEN:  Thank you, Dr. Helman.

DR. HELMAN:  Thank you.

DR. WHALEN:  I would like to ask if Ms. Crigler would like to return to read for the record Ms. Mullen's statement.

MS. CRIGLER:  The statement that I have for Ms. Mullen is the statement from March of 2000, which I am sure anyone here may not remember.  It would be repetitious.  Ms. Mullen was just saying that I could take full ten minutes just for expedience sake.

Basically, her statement is mainly as mine in that the saline implants gave her quality of life.  At that time, Ms. Mullen had immediate reconstructive surgery and she, as I, feels that it is really a quality of life issue for a cancer patient, much different than an augmentation that someone would request for body enhancement which, as with any procedure, people have used to excess as far as the augmentation.  That is my perception.  But for a cancer patient, we offer a different view as this gives us a different perspective on life.

As I stated, when women cannot even look at themselves in the mirror because of the mutilation with breast cancer, Dr. Susan Love has quoted the only cure is slash, burn or poison.  They slash our bodies to remove the breast.  They poison our bodies with chemotherapy and they burn them with radiation.  After a woman has gone through any or all of these devastating treatments you are looking at an end result as to what she did not choose.

The choice comes with an implant for reconstructive surgery.  This gives a woman a different view of life.  She can feel whole again.  I am sure people have noticed, whether you or anyone on the panel has done it, when people observe a female, most times the concentration is on their breast for whatever reason.  Women feel whole with breast.  For some women it is a dire need.  As we explained, most clothing is designed by men for women, but it includes some space for breast, whatever size.

I feel that it would really be unacceptable if the study is not continued to give women this option for whatever quality of life the woman may need at that time.  Just have the choice available.

DR. WHALEN:  Excepting, of course, the travel question, do you have any idea of Ms. Mullen's potential responses to the other areas that are asked of all speakers?  Is there any conflict?  Is she involved in lawsuits?  Is she paid by any company that is interested?

MS. CRIGLER:  No, she is not involved in a lawsuit or witness.  It is in her statement.  She is founder, CEO and president of WIN ABC, which is the Women's Information Network Against Breast Cancer.

DR. WHALEN:  Thank you.  Any other questions?

[N response]

DR. WHALEN:  Thank you.

MS. CRIGLER:  Thank you.

DR. WHALEN:  If Ms. McDonough is still present in the room and if she wishes to read for the record Ms. Dabis' statement, and if she can also begin, if you know it, with what Ms. Dabis' answers would be with lobbying companies she works for, and the like.

MS. MCDONOUGH:  She doesn't work for a lobbying company.

DR. WHALEN:  Oh, I am sorry, I thought you said that in response to a question.

MS. MCDONOUGH:  No, my response to the question was that I came here by myself and I was reading my own testimony, as is true with Cherien.  If she had been here this morning, she would have been here on her own and by herself.

DR. WHALEN:  So, the Sheridan Group has nothing to do with either of you?

MS. MCDONOUGH:  That is correct.

DR. WHALEN:  Thank you.

MS. MCDONOUGH:  That is correct.  Thank you for clarifying that.  To the best of my knowledge, all of the answers to the questions, the four of them, are no.

My name is Cherien Dabis.  I wanted to be here today to present to you information for women who have receive and will ever receive saline breast implants.

I was born with cystic hygroma, a rare benign tumor of the skin consisting of a collection of abnormal lymph vessels.  At birth, a tumor the size of a grapefruit was perched on the left side of my neck and interwoven with the delicate nerves and blood vessels in my neck, chest, arm and underarm.  A series of surgeries at birth removed the growth, but also necessitated the removal of surrounding tissue, skin and muscle, leaving me with excess scar tissue and half a pectoral muscle.  The scar tissue hindered the range of motion of my left arm.  As I developed, the asymmetry of my chest became more and more apparent.  My left breast was significantly smaller than my right.  What I saw as a deformity led me to hate my body.  So, I decided that when I was old enough I would pursue reconstructive surgery.

At age 18 I underwent tissue expansion, followed by stage-two reconstruction with breast implants at Christ Hospital in Cincinnati.  My plastic surgeon recommended silicone gel implants, but I had done my homework.  I read about the problems women had with silicone implants which, in 1992, led the FDA to restrict their use, ironically, making silicone implants only available to women who were born with defects or those having undergone mastectomies.

I assured my doctor that I did not want silicone implants.  He discounted my concern, but told me that the saline implants were a safer option.  If it leaks, he said, it is only salt water.  It will dissolve, leave your body and you will be unharmed.  But chances are that it won't leak.  It will last forever unless you suffer some major trauma to your chest, such as a car accident.  My doctor tried to convince me to have both of my breasts implanted.  I refused.  I simply wanted to correct my defect.

In Mary of '96, my chest was implanted with a McGhan style 168 saline-filled breast implant.  Another custom-made implant was inserted in my arm to fill the cavity that resulted from my birth defect.  I thought at the time this implant was also saline-filled.

Nearly four years later I began to experience periodic pain and burning in my chest and arm.  The burning and pain eventually worsened and my left arm became more and more immobile.  I knew something was seriously wrong.  I was experiencing painful capsular contracture.  My doctor thought I would never be able to improve the range of motion in my left arm.  Her recommendation was to surgically remove the implant.

During the few weeks following my initial appointment, the pain and burning in my chest worsened.  I could not sleep at night because I could feel the implant moving around, shifting and painfully poking me.  Then I had the ultimate breast implant nightmare.  I stepped out of the shower on June 1, 2000.  I knew it had happened to me.  My breast was gone.

One week later I was in an outpatient operation room, terrified.  After five hours of surgery to remove the implants, I woke up to find out that Dr. Feng had removed one deflated saline breast implant and the other device in my body which had been a solid silicone block.

My insurance did approve the procedure, but I still had to pay 20 percent up front, which amounted to $2000, and I was required to stay in a hotel room for five days so that the doctor could monitor my progress and remove my drains.  The total cost of my trip was $3000.  I had to take out a loan in order to pay for what the insurance had not covered.

Had I known the additional physical and emotional consequences of receiving breast implants, I would have made a different decision.  When I opted for reconstructive surgery I did my homework.  I read what little research was out there.  Only now do I know that the research that was available was conducted by the manufacturers.

As a young woman and a consumer, I ask you to require and enforce the most stringent guidelines on the integrity of applications and of the manufacturers submitting applications for device marketing approval and clinical trials.  I ask that long-term follow-up on the adverse events be demanded of all manufacturers, with close oversight by FDA.  Only through this process will consumers ever have access to accurate surveillance data on complications and failure rates of breast implant devices.

Until there is incentive for manufacturers to make better, safer and more effective devices, women will continue to risk their health, their future, and possibly their lives by choosing these FDA approved medical devices.  Thank you.

DR. WHALEN:  Thank you.  I realize that was a reading of a statement but does anybody have any questions?

[No response]

Thank you.

MS. MCDONOUGH:  Thank you very much.

DR. WHALEN:  We have a brief amount of remaining time so if there is anyone, who has not yet addressed the panel, who wishes to address us in this public session, would you please so indicate by raising your hand right now?  I see none.

This is a convenient time for us to take a break.  I have 10:07.  We will reconvene here at 10:25 and resume our activities.

[Brief recess]

DR. WHALEN:  I would like to remind the public observers at this meeting that while this portion of the meeting is, of course, open to public observation, public attendees may not participate unless specifically requested by a member of the panel.

Before we begin the first presentation from industry, I would like to ask Dr. Celia Witten to make some brief remarks about what the conduct of our day's activities is going to be.

DR. WITTEN:  Thank you.  I just want to describe what we are going to be doing for the rest of the day.  I will start with just mentioning something that I think everybody here probably already knows, which is that at the time that we approved the two saline breast implant PMAs several years ago, we asked the sponsors to continue to study their product in the form of several conditions of approval, which have been outlined in the panel packet.

The purpose of the meeting today is to give the sponsors an opportunity to describe their progress in those studies to date, and there will also be an FDA presentation on each of those studies.  Following that, there is an opportunity for the panel to comment on the data and the studies as presented.

We think it is very important that the panel get an opportunity to see what the data is that the sponsors have generated to date, and also for the public to get an opportunity to see this data.  This is important information and will be incorporated into the sponsors' labeling as part of an ongoing effort to make sure all the information is available to physicians and patients.

I would like to thank everybody here on the panel in advance for your discussion of the data, and I would like to thank the members of the public, who gave presentations this morning already, for their input as well.

DR. WHALEN:  Thank you, Dr. Witten.  So, as we proceed, as long as the day flows as we anticipate, we will have one of the sponsors give their presentation before lunch.  There will be three speakers, and I would ask the panel members to make note of any questions they have for any of those speakers.  Then, following all three speakers we will ask them to entertain our questions.  Following that question and answer period and comments, we will have FDA's presentation for that particular sponsor's presentation, followed, in a similar fashion, by questions, answers and comments.  We will then have a general discussion on that particular status.  Then, we hope to break for lunch and then duplicate that for the other sponsor following lunch.

I would ask that we begin with Mentor Corporation's presentation, with the three speakers who are going to address us.

Panel Update Regarding Post-Approval Conditions of

Approval for Saline-Filled Breast Prosthesis

Mentor Corporation

DR. MICHAEL:  Ladies and gentlemen, good morning.  My name is Maher Michael.  I am the medical director and vice president of clinical and regulatory affairs for Mentor Corporation.

We are here today to update the panel members and FDA staff on the status of the conditions of the post-PMA approval for Mentor saline-filled and Spectrum mammary prostheses.

First, I would like to give you an overview of Mentor's saline prospective study which constitutes the basis of our approved PMA.  There were 1680 patients enrolled in this study.  That study was designed for a three-year patient follow-up.  The devices that were used were 30 percent smooth and 70 percent textured.  The last patient was enrolled in September of 1995.  The PMA was approved by FDA on May 10 of 2000, with some conditions.

The first condition was the post-approval study, and the purpose for that study was to extend the patient follow-up from three years to ten years to collect longer-term safety data.  That study was not designed really to collect any patient satisfaction data.  Today, five-year data will be presented by Mr. Cliff Kline, director of clinical programs.

The second condition was the focus group study, and the purpose for that study was to evaluate the effectiveness of the patient brochure in communicating information to prospective patients about risks and benefits of breast implants.  The status of that study will be presented by Ms. Donna Crawford, director of corporate regulatory affairs.

The third condition was the retrieval study, and the objective of that study was to better understand causes of deflation and further analyze the failure modes.

The fourth condition was the fatigue testing study, and the purpose of that study was to characterize fatigue resistance of the devices using methods, as requested, in the breast implant guidance document.

The fifth, and last, condition was the real-time shelf-life testing, and the purpose of that study was to support our four-year shelf-life data submitted in our PMA, and to extend the real-time shelf-life testing to five years.

The status of the last three conditions will be presented this morning by Mr. Ron Crouther, vice president of advanced development.

Now I would like to present Mr. Cliff Kline who will present and discuss the post-approval study.  Thank you.

MR. KLINE:  Thank you, Dr. Michael.  Good morning.  As Dr. Michael said, I will be presenting the five-year clinical results for Mentor's post-approval study, or PAS, on our saline-filled and Spectrum mammary prostheses.

I will first briefly discuss study design and the chronology of our efforts to contact and follow-up on the majority of patients.  Then, I will provide five-year complication data.

The objective of the post-approval study was to assess the long-term ten-year safety for Mentor saline-filled and Spectrum breast implants by assessing the cumulative incidences of capsular contraction, deflation, breast pain, reoperations and explantations.

Patients were included in this study if they had participated in the saline prospective study and they consented to participate in this post-approval study.  They were excluded if they had died, had all their implants removed or discontinued by choice.

Patients could either complete a mail-in questionnaire or they could elect to be seen by their physician.  The questionnaires are sent to each patient once every year around the anniversary date of her surgery.

Mentor has conducted extensive and varied efforts to contact and follow-up with all potential patients, and I would like to share these efforts with you now.

We received PAS protocol approval in May of 2000.  During July we contacted all saline prospective study investigators and confirmed with them that Mentor would contact their patients directly.  Three physicians did deny us direct access to their patients so those patients were not contacted about participating in the study.

Prior to the first mailing, we worked with the participating sites to confirm the patients' addresses.  We also used the National Change of Address database, NCOA, to update this information.  We then initiated a patient mailing of informed consents and questionnaires, and at the end of 2000 we did a second certified mailing to those patients who had not responded, the non-responders.

You will hear me using those two phrases throughout the presentation, responders and non-responders.  Responders are those patients on whom we have data, whereas noon-responders are those patients who we have not yet had success in contacting.

In 2001, we continued to collect an analyze the data as well as send out annual questionnaires.  We increased our efforts to contact non-responders by using the nationwide 411 telephone directory.

This year, we continued to collect and analyze the data as well as mail out annual questionnaires.  We have increased our contact and follow-up rate by approaching investigators who have non-responders and, if they were successful in contacting these patients, we provided financial incentive.  We also began to correspond with non-responders via FedEx, which allows tracking and verification of patient signatures.  We did this by using the ChoicePoint nationwide database to identify all possible addresses for the non-responders, and then we followed up by sending, via FedEx, packets to all these addresses.

These extra efforts have resulted in increased contact and follow-up in 2002.  As noted in FDA's memo to panel, Table 5(a), currently the augmentation patient contact rate is 75 percent and the reconstruction contact rate is 91 percent.

This graph shows the improvement in patient follow-up from March to May of this year.  Augmentation improved from 54 percent to 64 percent and reconstruction from 73 percent to 79 percent.  The rates are different for contact and follow-up because these two are defined differently.  Contact rate is made up of all patients whom we have been able to get a hold of, whereas follow-up rate only includes those patients on whom we have data.  So a patient is counted as contacted whether she says yes or no to participating in this post-approval study, but she is only counted in the follow-up if we have data.  Please note that the complication data in this presentation are from the March, 2002 data set.

Before I address complication rates, I would like to address the issue noted in FDA's memo to the panel regarding responders and non-responders.  When we analyzed the study data we did find that non-responders were significantly different in some demographic and operative characteristics.  Adjustment for these differences showed essentially no change in the cumulative rate of complications at five years.  Therefore, Mentor concludes that responders adequately represent the entire study population at five years.  This analysis is currently being reviewed by FDA.

The complication rates were calculated via Kaplan-Meier analysis.  This is a statistical method used when 100 percent patient follow-up is not available, and it provides an estimated probability of a complication at a given time period.

The remainder of this presentation is divided into the augmentation and reconstruction cohorts.  I first want to talk about the augmentation patients.  They are defined as a patient who is normally healthy and at least 18 years of age or older, and desires breast enlargement.  The average age of augmentation patients in the saline prospective study at the time of surgery was 32 years of age.  Almost half were married.  The remainder, 30 percent, were single, and 22 percent were widowed, divorced or separated, and 80 percent has at least some college education.

This table details both the three- and the five-year cumulative Kaplan-Meier rates, as well as 95 percent confidence intervals for the complications of reoperation, explantation, capsular contracture, implant deflation and breast pain.  For each of the complications at five years the cumulative rate increased.  For example, reoperation went from 13.2 at three years to 20.2 at five years.  Capsular contracture, 9.0 to 10.1.  This is expected as this is a cumulative rate.  That is, occurrences of complications that occurred in years four and five were added to the three-year cumulative rate.

But it is important to note that for reoperation, explantation and breast pain there was no significant change in the complication risk rate per year during this five-year time period.  The risk rate to the patient for capsular contracture, Baker grades 3, 4 and unknown, decreased.  Only for deflation was there an increased risk rate per year.

If we compare this five-year deflation rate to the published literature, we can see that the post-approval study rate is within the range of published literature which is 0-27 percent.  As you can see, the rate for explantation also falls within the published rate.  Reoperation and capsular contracture actually fall slightly under the published rate.  Literature rates for breast pain were not captured.

This table details the top ten reasons for reoperation.  The first column is categorized by number of reoperations by breasts; the second column, reoperation by patient.  So, 52 patients of 198 patients and 98 breasts of the 343 breasts had capsular contracture as a reason for reoperation.  The top three reasons in both columns are patient requested size exchange, leakage/deflation and capsular contracture.

Explants, which are a subset of reoperations, are discussed in this table.  This details the primary reason for explantation at both three and five years if the rate occurred at a rate greater than five percent.  The primary reason at both three and five years was patient request for size exchange.

Now, before I discuss the reconstruction data, I would like to more specifically define a reconstruction patient.  This is a patient undergoing breast reconstruction as a result of breast cancer or congenital deformity.  She could be expected to face a more extensive initial surgery and require additional treatment such as radiation therapy, or chemotherapy.  Skin coverage over the implant, as well as achieving symmetry is more difficult than in augmentation patients.

Saline prospective study demographic data shows that at the time of surgery a woman had an average age of 46 years.  Approximately two-thirds of the women were married, and almost three-quarters of the women had some college education.

This table details the three- and five-year Kaplan-Meier rates, as well as the 95 percent confidence intervals for reoperation, explantation, capsular contracture, breast pain and implant deflation.  Again, this represents a cumulative rate so the five-year rates are higher than those at three years.  Please note that the five-year numbers are updated, whereas the three-year numbers are those as presented at the 2000 panel.

At five years there was no increased risk rate per year for any complication, while there was a decreased risk rate per year for implant removal, explantation, and reoperation, as well as capsular contracture Baker grades 3, 4 and unknown.

If we compare these rates to those in published literature, we can see that three of the five rates were within the range from the literature.  While reoperation did have rates above those in published papers, it is important to note that the patient risk rate per year actually decreased at five years.

We would also like to note that the cumulative three-year rate, as presented in our saline PMA, was 40.2 percent.  So, in the intervening two years the rate has increased less than three percent.  Literature rates for breast pain were not captured.

This table details the top ten reasons for reoperation.  Again, the first column represents the percentages as categorized by number of reops by breasts, and the most commonly reported reasons were capsular contracture, asymmetry and patient request.  The second column, by patient, the top three reasons were capsular contracture, asymmetry and leakage/deflation.

This table details the primary reasons for explantation surgery at three and five years.  As you can see, the primary reason was capsular contracture.

In overview, patient contact and follow-up rates have increased.  The responders adequately represent the entire study population through five years.  The risk rate per year for the complications discussed decreased or stayed the same for four of five complications in the augmentation cohort and all five of the complications in the reconstruction cohort.  For deflation, in the augmentation cohort the rate was within published literature.

In summary, Mentor's saline-filled and Spectrum implants continue to perform in a safe and effective manner.  As presented today, the complication rates are comparable with published rates.  We will continue to follow patients through ten years and seek to continue to increase the total number of responders.

Thank you.  I would now like to present Ms. Donna Crawford.

MS. CRAWFORD:  Thank you, Cliff.  Good morning.  My name is Donna Crawford.  I am director of corporate regulatory affairs for Mentor Corporation.

I will be discussing the focus group study or the patient informed decision brochure.  Mentor conducted this study as one of the post-approval conditions of saline breast implant PMA.

There were four major purposes of the focus group study.  The first was to determine whether the patient brochure effectively communicates information about the risks and benefits associated with breast implants.

Secondly, it was important to assess whether the information in the brochure is presented in an understandable way and is clearly understood by prospective patients.

Thirdly, we wanted to identify any unintended effects of the brochure and also any unanticipated effects of the brochure.

Finally, we wanted to obtain patient suggestions for improvement and identify any additional information needed by the patients.

The focus group study was conducted under an FDA-approved protocol by an independent research group by the name of Communications Sciences Group.  Four focus group discussions were held, two in Dallas and two in San Francisco.  The focus groups consisted of reconstruction patients or patients considering augmentation.  There were eight to ten individuals in each focus group, and each group was balanced across age, employment status, income level and educational level.

Data were collected in two ways.  First, the participants were asked to read the brochure and complete a self-administered survey prior to participating in the focus group interviews.  Secondly, the focus group interviews were led by a moderator who followed a discussion guide which was part of the focus group study protocol.

Some of the key findings of the study are as follows.  In general, the educational and informed decision objectives of the brochure were met.  The majority of women had a good understanding of the risks and benefits associated with breast implants after reading the brochure.  Eighty-eight percent of the respondents reported that they had learned new information about breast implants after reading the brochures, and 85 percent felt better able to ask their doctors questions about breast implants after reading the brochure.

Most respondents felt that the brochure, on the whole, was clear and understandable, with the possible exception of the clinical data tables.  Most respondents rated the brochure highly on comprehension and relevance, and 88 percent felt that the information in the brochure was useful to them.

The study identified only one possible unintended effect of the brochure, and that was that some of the respondents felt that complication rates were not to be taken at face value because they were overstated in order to protect the manufacturer.

Both the focus group discussion and the survey results found that the brochure was effective in conveying information, and 73 percent said that the information in the brochure was not confusing, and only six percent of the respondents felt that the brochure was confusing.  There was some difficulty in understanding the meaning of cumulative risk rates and interpreting the data tables.

The major suggestions for improvement had to do with improving the layout and format of the brochure; adding explanatory information to the data tables; choosing the content order to group augmentation data together and reconstruction data together; and adding a glossary and table of contents.  The additional information amounted to small points of clarification only.

In response to the focus group findings and suggestions from FDA, the following changes were made:  A table of contents and glossary were added.  The clinical study section was revised to separate augmentation and reconstruction data; and to simplify and explain the data tables.  Some minor wording changes were made to improve the clarity.  For example, an introductory paragraph was added to the clinical study section to explain how Mentor's clinical study data may relate to each patient's own experience.  Sentences were added prior to each clinical data table to explain what the numbers in the table mean and how they were calculated.

The brochure is in the process of being revised to add the five-year follow-up data from the post-approval study.  The revised brochure will be available on the Mentor's web site and a printed version will be available in approximately six weeks following FDA approval.  Therefore, this condition of approval has been fulfilled.

I would now like to introduce Mr. Ron Crouther.

MR. CROUTHER:  Good morning.  My name is Ron Crouther.  I am vice president of advanced development for Mentor, and I would like to present interim results on three studies, our explant retrieval study, our post-approval fatigue testing, and also our real-time shelf-time testing.

First on the explant retrieval study, the objective of this study was to retrieve 300 devices that were explanted because of deflation and perform appropriate analyses to determine the mode of failure.  The 300 devices were selected to cover a range of saline device types, that is, smooth and textured devices, various shapes, and devices containing our two valve types, diaphragm valves and Spectrum kink plug valves.

Upon receipt of the explanted devices, we first captured the device descriptive information and clinical information.  This included device type, date of manufacture, time in vivo and time of surgery.  All devices were then visually examined and additional microscopic examination of the surface of the defect area was conducted to better characterize the type of failure mode.  Leak testing was performed, as was necessary, to confirm that all leak sites had been located.  The last phase of our testing was physical and mechanical property testing.

We provided an interim report on 38 devices to the FDA in August of 2001.  We completed our study of 310 devices and submitted the final report in our May, 2002 annual report.  Because we have just submitted the final report and the FDA has not had adequate time to complete its review, the presentation today is based only on the interim report on 38 devices.

This table summarizes the failure modes we saw on the 38 devices for both smooth and textured devices.  The first failure mode is material separation, a term which Mentor has used for many years to describe a tear or split in the device without very significant abrasion or thinning at the site of the defect.  Material separation is the primary failure mode of our Siltex devices.  As supported by a wider study of Siltex devices, most material separation failures do exhibit evidence of folding at the site of the defect, often compound folding in which the outer surface of the implant is in tension causing the split to propagate from the outside in.

The second failure mode is smooth crease-edge opening or crease-fold failures on smooth devices.  These are similarly related to folding and, in contrast to material separation defects, do exhibit very obvious abrasion and thinning at the failure site.

Sharp-edge openings can be duplicated by puncture with a surgical instrument and are very likely indicative of iatrogenic damage.  We also had two reported cases of leaking valves, however, one could not be confirmed and, if you will notice, the numbers in the left-hand column do not add to 19 as we had one Spectrum kink plug valve device in which the tubing was not fully withdrawn per our instructions for use, and when the tubing was withdrawn the device did not leak.

To summarize, the failure modes that we saw in this interim report reflect what we have seen in our clinical study information, as well as our product evaluation or complaint database.  There is evidence of folding present in a majority of the failures that did not exhibit evidence of instrument damage.  Again, Mentor's final report covering the 310 devices is under review by the FDA.  The general findings in that study are similar to what we saw in the preliminary report.

The next topic I will briefly discuss is Mentor's fatigue testing of our saline-filled breast implants.  Fatigue testing involves cyclic compression testing for our implants up to 10 million cycles and utilizes an apparatus that is schematically represented here.  The sample is located here, between two flat platens, one of them fixed and one of them movable.  The movable platen oscillates up and down to apply the compressive force to the implant.  The implant is immersed in a saline bath that is kept at 37 degrees Celsius.

We had two phases and two objectives to our fatigue testing.  The first objective was to create AF/N curves, that is, applied force versus number of cycles to failure for the various device types.  There were four device types chosen for the study.  Along the Y axis we have the applied force; along the X axis are cycles to failure.  As you can see, as you decrease the applied force the device will withstand more cycles until failure.

The second phase of our testing was to conduct long-term fatigue testing and calculate fatigue safety factors for each of the four device types.  The fatigue safety factors are calculated according to the formula you see here.  I am not going to try to point to it, but it is the force to achieve 10 million cycles without failure, what we call our run-out load, divided by the estimated device load during walking, which we conservatively estimate to be two times the device weight.  The safety factor chosen for this, in consultation with the FDA, was greater than or equal to 2.

The protocol was approved by the FDA.  Four device types were chosen for the study.  Those four were the smooth and Siltex round diaphragm valve devices, smooth round kink valve devices and contour tall profile devices.  The tested devices include all sterilization methods that we currently use and that are approved by the PMA.

We chose for the testing the smallest devices of each device type as those typically have our thinnest wall thickness.  We also made special runs of those devices in which we ran at the absolute minimum of our tolerance extreme for shell wall thickness as well as texture layer thickness.  As such, this device configuration represents worst-case physical testing as defined in the FDA's breast implant PMA guidance document.

These are the results to date for three of the four devices completed.  This is the phase one testing, which was the generation of the AF/N curves.  Again, along the Y axis we have load amplitude; along the X axis we have cycles to failure.  It is a logarithmic scale so, again, 10,000, 100,000, a million and so forth.  The three devices are the smooth Spectrum device, the Siltex diaphragm valve device and the smooth diaphragm valve device.  So, for example, for the Siltex diaphragm valve device a load amplitude or force applied of 50 lbs requires approximately one million cycles to generate failure.

Based on this phase one testing, as well as other experimental data, a run-out value of 10 lbs was chosen for our second phase of testing, which was the long-term fatigue testing.  Again, the goal of that was to withstand the 10 million cycles without failure at the 10 laboratory load level.  We have successfully completed testing on three of the devices.  Again, if we calculate a safety factor based on the formula here, the run-out load of 10 lbs is divided by the estimated device load during walking, or two times the device weight.  Since we chose the same size device for each of the three device types, 125 cc device in each case, the device weights were very close to the same thing, so ten divided by two times the device weight for all three devices turns out to be 16.7.  Again, if you recall, our acceptance criterion was a safety factor greater than 2.

To summarize, three of the four device types have been completed.  The fourth is scheduled to begin shortly, probably within about the next month to month and a half.  AF/N curves have been generated, and fatigue safety factors have been calculated for three of the four devices, which far exceeds the protocol requirements, again using worst-case test samples.

The last topic I would like to discuss this morning is our real-time shelf-life testing.  As Dr. Michael mentioned, this was twofold in terms of objectives.  First, to support our current shelf life of four years and then, secondly, to extend that shelf life out to five years.

The testing is being performed under an FDA-approved protocol.  Seven device types are included in the testing in order to cover small and large devices; all packaging types and sizes; all sterilization methods; and all different device and component configurations.

The testing includes mechanical and shell tensile property tests, which include tensile and elongation, tension set, joint strength and valve competency tests.  The packaging seal peel strength, microbial challenge and dye penetration tests are performed to ensure maintenance of sterility.

Our status is that all devices have been tested at the time zero time point, with all devices meeting all specifications, and the four-year testing will be completed in the year 2005.  The five-year testing will be completed in the year 2006.

This concludes my presentation, and I would now like to reintroduce Dr. Michael.

DR. MICHAEL:  The overall summary of our presentation this morning is that we presented five-year data for the post-approval study, and we will continue to follow all our patients through ten years, and we will continue to update the agency in our annual reports.

For the focus group study, the study has been completed.  For the retrieval study, we have completed the study.  It is under FDA review.  Fatigue testing, we tested three out of four styles.  The last style is scheduled to start six weeks from now.  Lastly, the shelf-life testing is ongoing through five years.

Mr. Chairman, panel members, FDA staff, I would like to thank you for your attention and, after FDA's presentation, we would be glad to answer any questions you may have.  Thank you.

DR. WHALEN:  Actually, if you wouldn't mind, we will ask you to answer some questions before FDA's presentation and then perhaps again.

DR. MICHAEL:  That is fine.

DR. WHALEN:  If I could start off, Ms. Crawford, the focus groups had relatively small numbers as most focus groups do, but it was unclear to me what the composition of that group was vis-a-vis thinking about having the implant, having had the implant, having had it and having had it removed.

MS. CRAWFORD:  Yes, the augmentation focus groups consisted of patients who were considering augmentation.  The reconstruction focus groups were primarily patients that had already had reconstruction with breast implants.

DR. WHALEN:  And all of them were still in place?

MS. CRAWFORD:  Yes.

DR. WHALEN:  Dr. Dubler, do you have any questions about the focus group results?

DR. DUBLER:  Yes, I have one question about the focus groups, that is that it seemed to me that one of the findings was that they really didn't understand the data sets and what they said.  Would you agree?

MS. CRAWFORD:  There was some confusion about exactly what the meaning of what a cumulative risk rate was, and there was some confusion in terms of the data tables, which often switched from identifying data by patient, some by implant, some by total patient population.  We have addressed that and made clarifications to the data tables to address that confusion.

DR. DUBLER:  Have you left them in table form?

MS. CRAWFORD:  Yes, the data tables still are included in the brochure, yes.

DR. DUBLER:  Have you translated them into reasonable lay language, or simply left them in the table form?

MS. CRAWFORD:  It is in table form by patient and each table has at least a couple of sentences of introduction to explain what the numbers are.

DR, DUBLER:  I have another question about the presentation.  Would it be appropriate to ask it now?

DR. WHALEN:  Yes.

DR. DUBLER:  I believe it was Dr. Michael who stated that there were three exclusions from your follow-up cohort.

DR. MICHAEL:  It was Cliff Kline in his post-approval study presentation.

DR. DUBLER:  The three exclusions were patients who had died?  What were the three exclusions, please?

MR. KLINE:  The three exclusions were patients who had died, had their implants removed or had withdrawn by choice, discontinued by choice.

DR. DUBLER:  Discontinued what?

MR. KLINE:  They elected not to participate in the post-approval study when we asked for their participation and consent.

DR. DUBLER:  I see, and what percentage of your cohort was represented by those three categories, total cohort?

MR. KLINE:  I would have to look and determine those numbers by those three groups.  I don't have that information available at the moment.

DR. DUBLER:  Why did you decide to exclude patients who had removed the implants?  One would think that they would be an important source of information.

MR. KLINE:  You are exactly right, in that we captured all complications, including removal, and reported those complications today.  After a patient has her breast implants removed, she no longer has a study device in her body and, therefore, she is no longer studied.  But, of course, until the device is removed we do study her and collect and report on all complications.

DR. WHALEN:  Dr. McCauley?

DR. MCCAULEY:  I just have a question related to the focus group study.  Did these groups have group leaders, and if they did, how were they chosen?

MS. CRAWFORD:  There was a moderator that was an employee of Communications Sciences Group that followed the discussion guide that was part of the protocol.  There wasn't a leader of the focus group participants per se, but there was a moderator who conducted the discussion among the focus group participants.

DR. MCCAULEY:  After your separation out of the data, was there significant improvement in the confusion level or decrease in confusion that some of the participants expressed?

MS. CRAWFORD:  Well, since we did the focus group study on the original version of the brochure before the data were separated out, we haven't gotten feedback on the brochure with the data separated.  So, I can't answer that question.  It was primarily so that somebody undergoing reconstruction would not have to sort through the augmentation data that wasn't applicable to them.

DR. WHALEN:  Dr. DeMets, would you have any questions or comments about the statistics that were presented to us, or any particular vantage point on the percent follow-up that they have vis-a-vis other clinical studies in a population similar to this?

DR. DEMETS:  Yes, I actually have two sets of questions, one for Mr. Crouther and one for Mr. Kline.  You said that you recovered 310 implants and those reports have been done and submitted.  Why is that we aren't privileged to see that today?

MR. CROUTHER:  We just submitted those to the FDA with our annual report and it is still under review by the FDA.

DR. DEMETS:  Is there some regulatory reason why we can't see that data?  I mean, you are giving us 38 and there are 310, or something, that are available.

DR. WITTEN:  We usually review data and send it to the panel before it is presented in a panel session.

DR. DEMETS:  All right.  On the time to failure, how many actual samples are tested in that?  Is it one sample from each of the devices?

MR. CROUTHER:  No, for the long-term fatigue testing a minimum of three, and for the AF/N curves three also.

DR. DEMETS:  Did you calculate or compute any time to failure?  You only have three but you presented your data on how many cycles it took, but is there some way you can translate that into failure time?

MR. CROUTHER:  Not directly into failure time.  The 10,000 cycles, again, was agreed to with the FDA and 10,000 cycles represents walking for eight hours a day at a rate of one Hertz or one cycle per second.

DR. DEMETS:  The remainder of my questions are for Mr. Kline.  Can you explain to me in a little more detail the process that you went through to capture information from the patients just to get responses?

MR. KLINE:  Certainly.  In terms of how we got the patients to agree to participate or not?

DR. DEMETS:  Just the process.  I am not understanding the process that you went through because the response rate is low relative to standards that I am used to.  So, I am trying to understand what you did do.

MR. KLINE:  Well, once we did get protocol approval through working with FDA, we then contacted first the investigators to see what information they had on the patients because, as you know, a patient population can move.  So, once we determined first from the investigators that they wanted us to contact the patients directly, we then worked with them to get the addresses.  Once we had the updated addresses from them, we then also checked with NCOA to make sure that the doctor's address was verified by NCOA.  Then we began to correspond via mail with the patients.  The SPS study only had patients consented for the three-year study.  That is why in the first mailing we included a questionnaire as well as a consent to ask them to participate in the study through ten years.  So, the first step, once we had the address was to see if the patient would consent to participate in this additional follow-up.

DR. DEMETS:  Was there a reason that you didn't ask the investigators themselves--not the investigators, the surgeons to collect this information from their patients as opposed to doing it directly from a mailer?

MR. KLINE:  It was an option, and we asked each doctor if they wanted to contact the patients themselves or they wanted Mentor to, and they all elected--except for the three that declined any contact with their patients--us to be their representative and to contact the patient directly.

DR. DEMETS:  And, was the questionnaire such that the patient would be able to fill out all of the items accurately?

MR. KLINE:  Yes, once they got the questionnaire, the form was fairly basic but also had an explanation as to what we were asking them to do.  There was a letter accompanying that as well as, of course, the informed consent in the first mailing.  If they had any questions, they could call us.  There was a phone number for them to call.  Additionally, if there was something that was incorrectly filled out on the form, the clinical research associate working on that study would contact the patient directly and get clarification.

DR. DEMETS:  Do you have any insight as to why the response rate is what it is?

MR. KLINE:  Well, the response rate, as we have heard today, has improved.  We will continue to work to improve the number of responders.  I would say that many of the patients are transient--not transient but have moved from their location where they were being seen by their doctors as part of the SPS study.  The doctors sometimes did not know where they were at, or thought they knew and it turned out that they didn't.  So, some of it was just finding the patients, which is one of the ways that we improved follow-up by using the ChoicePoint database, which is a database which gives more current addresses than even NCOA.  Then, using that as well as FedEx, we were able to track the patients down and determine, if there were three addresses for one patient, which address was the correct one, if any, and then work with the patient.

DR. DEMETS:  Do you have any sense of what you think you can get this response rate to be, and what a target should be?

MR. KLINE:  Well, I don't.  I don't want to give a hypothetical because we don't know, but we are continuing to work even as we speak to improve the contact and follow-up rate.

DR. DEMETS:  Do you have a sense of what a target rate should be?

MR. KLINE:  Well, as in an FDA guidance document, for a two- or three-year study we would hope 80 percent.  Obviously, at ten years I would estimate we would like to have 60 percent.  We are obviously somewhere in between there and are working to improve both contact and follow-up.

DR. DEMETS:  So, you think an 80 percent response rate at five years and a 60 percent at ten years would be adequate for your purposes?

MR. KLINE:  Yes, 80 percent at two to three years and 60 percent at ten years.

DR. DEMETS:  Do you have any sense of what a typical clinical trial expects in follow-up response rates?

MR. KLINE:  Just talking about the breast implant studies today, the breast implant study guidance document indicates 80 percent follow-up at least at two to three years and, of course, we are enrolling adequate patients in other breast studies for adequate follow-up for ten years.

DR. DEMETS:  Well, I will comment later but I worry a lot about the adequacy of even those response rates given the potential biases that exist, and those biases can be very powerful.

I would like to follow-up on a comment or a claim that you made that the responders are the same as the non-responders.  Could you detail that?  That is a strong statement you make and I am trying to understand that.

MR. KLINE:  That is a very good question, and I would like to have Dr. Eugene Poggio, managing vice president and executive director of biostatistics and epidemiology at ACT, address that since he is a biostatistician and can more correctly address your question.  Is that okay?

DR. WHALEN:  Sure.

DR. POGGIO:  My name is Gene Poggio and, as Cliff indicated, I am managing vice president of biostatistics and epidemiology at APT Associates Clinical Trials.  We are under contract to Mentor to do data management and statistical analysis for both the original SPS study and the follow-on post-approval study.

I have no personal financial connection with Mentor, aside from the fact that the firm I am employed by has contractual arrangements.  My travel was paid through that contract by Mentor.  I am not involved in any lawsuits whatsoever, and I certainly derive no income from implant surgery.

With regard to the issue of response/non-response, the way we dealt with that is we took all the information we had on baseline characteristics and operative characteristics and conducted logistic regression.  This was done separately for the augmentation cohort and the reconstruction cohort.  We used logistic regression to identify variables that were significantly related to response/non-response.  I should say that for purposes of that analysis, response/non-response was defined as participating in the PAS versus not participating in the PAS, and didn't deal with loss to follow-up before that.

Through that, we identified some variables that were related.  For example, for augmentation it was age and annual income, and for reconstruction it was several variables, some operative characteristics and some demographics.  Having identified those variables, we then stratified the population by that set of variables, conducted a Kaplan-Meier analysis within each stratum and then computed a weighted average of the Kaplan-Meier estimates with the weights being the initial number of patients in each stratum.  The result was numbers that were remarkably close to the original estimate.  The largest deviation of the adjusted number to the original number was half a percentage point.

DR. DEMETS:  So, what percent of those risk factors that you identified explained the response/non-response rate?

DR. POGGIO:  What was the total percent explained?

DR. DEMETS:  I know it is not easy to answer in a logistic regression but in analysis of variance you could do that.  There are such measures, by the way.

DR. POGGIO:  I don't know the answer to that question.

DR. DEMETS:  So, suppose it didn't explain much, would your analysis or your adjustments be useful?  You are making a claim that these two groups are the same so I am challenging that.

DR. POGGIO:  I guess my statement would be that we did everything we could to adjust for it.  We looked at the information we had about baseline variables, both demographic characteristics and operative characteristics.  We looked at which ones related to response/non-response and then adjusted for those.  It is obviously conceivable there is some other variable that we don't have access to that could explain part of it.

DR. DEMETS:  I am not sure who wants to answer this, but there was a comment that there was no follow-up of those patients in whom the implant was removed.  Somewhere else it was said that you used a Kaplan-Meier methodology to censor that observation.  There are some assumptions that are required to employ the Kaplan-Meier method about censoring.  Can you comment on how that was investigated in the study?

MR. KLINE:  This is Cliff Kline.  Dr. Poggio can explain how the Kaplan-Meier was used in this setting.

DR. POGGIO:  I think you are referring to some of the earlier comments that we excluded patients with explants.

DR. DEMETS:  Let me be specific.  The assumption is that the censoring mechanism is independent of the process that is going on.  So, I am trying to understand how you came to that conclusion.

DR. POGGIO:  Let me make clear at the outset that patients that were explanted were certainly kept in the analysis up until, if you will, the day after they were explanted.  If we had data after that, all that data was reported to the FDA. In order to be conservative, we didn't feel patients were at risk for the complications once the implant was removed.  You are not really at risk for capsular contracture after removal.  So, if we were to include them after that point the estimates would actually go down, and we didn't think it would be appropriate to do that.

In terms of the issue of the censoring, yes, obviously the underlying assumption in Kaplan-Meier is that the people who aren't censored look like the people who are censored--rather, the other way around.  Obviously, the adjustment we made was a refinement on that in that we don't assume that the censored people look like all of the uncensored people.  We assume they look like the uncensored people in the stratum that we defined by the variables that we stratified.

DR. DEMETS:  I am still not sure I fully understand the statement you are making.  The other issues I want to ask about have to do with if you censor the patients at the time of the implant being removed, do I understand that complications that may take place after that point are not captured?   I mean, we have heard this morning's testimony that some of these complications can occur--

DR. POGGIO:  They are captured in the data provided to the FDA, in data listings.  They are not included in the analysis because we don't feel--I mean, it was really my decision.  We didn't keep them in the Kaplan-Meier because we don't feel they are appropriately included predominantly because those patients are not at risk for most of the complications with the device no longer in place.

DR. DEMETS:  That is an assumption, it seems to me.

DR. POGGIO:  I guess if you asked a surgeon if they could have capsular contracture, a new case of capsular contracture after the device was removed, I would assume they would support that.

DR. DEMETS:  But my question was not about that; it was about other complications.

DR. POGGIO:  As I said, we didn't feel they should be included in the analysis because they are certainly not at risk for many of the complications.

DR. DEMETS:  Which is what the data are trying to understand, if they are there.  Aren't you precluding that if you take them out?

DR. POGGIO:  If I include them in I assume that they are at risk and, in fact, I think that would lower the estimated complication rates.  You could do a special analysis to look at that question.  I would be very reluctant to include them in a principal analysis when I no longer think they are at risk.

DR. DEMETS:  It seems to me that you get rid of some of the complications if you don't follow them beyond that censoring point.

DR. WHALEN:  If I could interject, would it be a more graphic example to state that deflation would be rather ridiculous to continue to measure in someone who has had an explant because it is hard to deflate something that isn't there?  Indeed, if they kept those patients in the denominator it would dramatically lower the deflation rate falsely.

DR. DEMETS:  I am thinking about other kinds of complications.

DR. WHALEN:  I am only trying to interject to say that it seems to me that there have to be different sets of data of complications, ones that could continue with the implant in place and ones that would not.

DR. POGGIO:  And, remember that the ones we are looking at are deflation, capsular contracture, explantation which, obviously could only happen if they had a reimplantation.  Obviously, breast pain could.  But my view is it would be reasonable to look at that as a separate issue but I still would be very reluctant to include it in a principal analysis because they are not at risk for some of the complications, and certainly much lower risk for some others.

DR. DEMETS:  My last question is about the table at five years that you compare to three years, the denominators are different.  In fact, they are larger at five years than at three years.  Can you explain that?

DR. WHALEN:  Could you give a number that you are looking at so we can make sure we are looking at the same thing?

DR. POGGIO:  It is certainly true that we did get additional three-year data in the PAS study also.

DR. DEMETS:  It is your slide 19 in your presentation.  At five years you report 211 patients, at three years you report 138.  There are several like that but that is one.

DR. POGGIO:  Sure.  In that table the N under five years, which is 211, indicates the number of implants that had reoperations, whereas at three years the number of implants involved in reoperations was 136.  Therefore, the denominator increases because we had more implants involved in reoperations.

DR. DEMETS:  So, are we looking at comparable groups?

DR. POGGIO:  You are looking at 211 explants at five years and 136 explants at three years.  So, there are obviously more explants as time goes on.

DR. DEMETS:  Thank you.

DR. WHALEN:  Dr. Dubler?

DR. DUBLER:  Could I ask one more question, please, about the focus groups?  I am very concerned about how the resurgery rate is presented and defined.  So, can you tell us, and this may be unreasonable, how that was explained in your first "informed consent" or what I prefer to call disclosure document, and after your focus groups how you might have restated that?  So, if there is a table about the chance of resurgical interventions?

MS. CRAWFORD:  Yes, I am looking at the version of the brochure that we did the focus group study on right now.  There is a section on reoperations.  It does give the three-year risk rate of reoperation by patient and by implant.

DR. DUBLER:  So, there is a table?

MS. CRAWFORD:  There are tables.

DR. DUBLER:  And what do you say about the tables?  Are there declarative sentences?

MS. CRAWFORD:  Prior to the tables, for example, it just indicates that the following are the cumulative risk rates, first occurrence for the following complications.

DR. DUBLER:  So that is the text?  Could you read us the particular language?

MS. CRAWFORD:  Certainly.  I just want to make sure I am looking at the right place here.  The question at the heading of the section is what were the three-year cumulative complication risk rates of first occurrence?  Then the sentence following that says the cumulative risk rate of first occurrence which occurred in at least one percent of the patients are shown in the following tables, including all levels of severity, mild to severe.  Then it lists the complications.  This particular section is talking about augmentation and it lists the complications that were found in augmentation patients.

DR. DUBLER:  Are there any other declarative sentences that surround the table?

MS. CRAWFORD:  Not in the original version.

DR. DUBLER:  And in the version that was modified by the focus groups, is it very different?

MS. CRAWFORD:  Yes, the column by implants was eliminated, and there were some explanatory sentences.  I don't have that right in front of me at this moment, but that was added.  In the introduction of the whole section on cumulative risk there were a couple of sentences explaining how the risk rate can be interpreted and what it means to the patient.

DR. DUBLER:  Let me ask anyone to answer the following question, what is your chance of going back in for surgery, what is your five-year chance of going back in for surgery after receiving a breast implant?  That is my question.  What would you say to me?  What is your percentage chance?

DR. MICHAEL:  Are you talking about augmentation?

DR. DUBLER:  I want to know how you would translate that table.

DR. MICHAEL:  Are you talking about the augmentation group or the reconstruction?

DR. DUBLER:  Either one, take your choice; augmentation.  What is the chance that you are going to have another surgery in the next five years?

DR. MICHAEL:  Well--

DR. DUBLER:  What percentage chance?

DR. MICHAEL:  Let me mention something here for the augmentation group, in one of the tables that was presented this morning 30 percent of reoperations in the augmentation group was the patient's choice for change of size.  In the reconstruction group 16 percent was the patient's choice for a larger size or a different size, and 15 percent in that cohort was expected because that was stage reconstruction.  So, one-third out of that.

DR. DUBLER:  I want to say to you I know I may want to decide to change my size but, given all of the factors that lead to surgery, what is the chance, including my changing my mind--what is the chance that I am going to have another surgery in the next five years?

DR. MICHAEL:  Based on what we presented in our data this morning, the chance, if using the same mix of the product that was used in the SPS study, at five years the probability of having a deflation is 9.7 percent.

DR. DUBLER:  That is not what I asked you.  What is the chance in the next five years, for any reason, that I am going to have to have surgery again?  What do your data show?

MS. CRAWFORD:  It shows that there is a 20 percent risk rate at five years, and the way that is explained in the patient brochure is that 20 out of 100 patients will experience at least one reoperation during five years.  That is how it would be explained.  There is an example given, and to translate that to a reoperation number, that is how it would read.

DR. DUBLER:  So, your statement is that there is a 20 percent chance in the next five years that you will have to have another surgery.

MS. CRAWFORD:  That is how it would be interpreted.  For example, the brochure right now gives an example for a cumulative risk rate of two percent for infection.  That means that approximately two patients out of 100 will experience at least one infection sometime during the first year.  So, since our five-year risk rate was 20.2 percent for reoperation, that is how it would be interpreted.

DR. DUBLER:  Thank you.

DR. WHALEN:  Can I just interject, more in comment, if a third of your reoperations are patient choice for a different size, would it not perhaps be more appropriate to say there is a 20 percent chance I am going to have another operation; there is a 14 percent chance I am going to have to have another operation?  Since the wording you use was I am going to have to have another operation?

DR. MICHAEL:  I would like to take a couple of minutes to introduce Dr. Roger Freedman.  He is a clinical instructor--

DR. WHALEN:  For what purpose?

DR. MICHAEL:  To elaborate on that issue of the percentage of reoperations based on his experience in practice.

DR. WHALEN:  I don't think that is necessary.  Thank you.  Dr. Newburger?

DR. NEWBURGER:  I have a question for Mr. Kline regarding the attempts to contact patients to participate in the post-marketing study.  How long is the questionnaire that they are sent?

MR. KLINE:  One page.

DR. NEWBURGER:  How many questions on that one page?

MR. KLINE:  Of course, we ask them to verify that their name is spelled correctly, but there is a question on capsular contracture, explantation, reoperation.  So, there is just a very small list of questions to specifically ask them about the complications we are collecting.

DR. NEWBURGER:  Is there any incentive to the patient to return the questionnaire?

MR. KLINE:  Yes, there is.

DR. NEWBURGER:  And what would that be?

MR. KLINE:  I don't recall the exact amount.  It is minimal.  It would be somewhere under $30.

DR. NEWBURGER:  And, how do you explain your protection of patient confidentiality?

MR. KLINE:  Well, that is explained both in the original letter they get as well as the informed consent that they sign which, of course, indicates that we will try to protect their confidentiality; it may not be able to be done if, for example, a government agency such as FDA needs to review these data.

DR. WHALEN:  Dr. Doyle?

DR. DOYLE:  I have a question regarding the reoperations.  How are these women included in the further data?  Is that reoperation considered a new start, or is it a continuation of the patient?  Also, I am confused about your three-year and your five-year cohorts of women.  They are not necessarily the same patients, is that correct?

MR. KLINE:  To answer your second question first, all post-approval study patients were originally SPS, saline prospective study patients.

DR. DOYLE:  But the people who answer at five years and the people who answer at three years are not necessarily the same group of patients.  As you might ordinarily expect where you have attrition, that those who were left at five years would have answered at three.  In this you have two separate groups actually who all started together but the five years may not be in the three-year data and the three years may not be in the five-year??

MR. KLINE:  I would agree with that, but most of the patients--as the panel members saw--that were present two years ago, our follow-up rates were a little bit higher so most of the patients that were reported on at the three-year PMA were also included in this study.

DR. DOYLE:  But some of your five years were not included in your three years.

MR. KLINE:  Pardon?

DR. DOYLE:  Some of your five years were not included in your three years.

MR. KLINE:  There were a couple of patients that we were able to collect additional information on that we didn't have three-year information on at the time of the PMA.

DR. DOYLE:  Okay, and what about how you handled the patients with the reimplantations?  Do they start at year one or do they continue?

MR. KLINE:  They continue unless the reoperation--as I said, explants are a subset of reoperations, obviously if they are explanted they are no longer--

DR. DOYLE:  No, these are the ones who were reimplanted.  They start out as year one or continue?

MR. KLINE:  They continue.

DR. WHALEN:  Are there any questions by any of the panel members about the biomaterials presentation?  Dr. Miller?

DR. MILLER:  I just have a couple of questions on that.  Ten million cycles in that machine, how many years of walking does that represent?

MR. CROUTHER:  That is one year of walking eight hours a day, one second per cycle.

DR. MILLER:  How do you feel that system, that test that was devised--I know it was all agreed upon by FDA, but how do you feel that actually simulates what that implant is experiencing, especially considering the fact that so many failures occurred at the site of a fold?  Do you think that testing in this way is a good simulation of what the implant experiences?

MR. CROUTHER:  It is a good simulation of some of what an implant experience is; it is not a good simulation to duplicate a fold failure like a material separation defect.

DR. MILLER:  Which is how most fail.

MR. CROUTHER:  Correct, most of the textured devices that are non-instrument damaged devices fail that way.

DR. MILLER:  The other question I had is that several of the presenters this morning suggested that there are materials released from the device that may be toxic, and platinum was mentioned several times today.  Is it fair for me to ask about your response to the manufacturing methods that employ platinum, and is this a concern that you have about the presence of such trace materials?

MR. CROUTHER:  I am going to ask Phil Yang, who is our corporate vice president of technology submissions, to answer that question.

MR. YANG:  Phil Yang.  We have done a risk analysis based upon what we know.  We do a specific analysis for I believe 20 heavy metals, of which platinum is one.  We do do that.  That is compared against what is in the literature for toxicity data.  So, we do compare that.  That is all part of the PMA.

DR. MILLER:  When you determined those materials are released, do you look at an implant subjested perhaps to a type of environment as described in your mechanical study?

MR. YANG:  We can't do that because probably we would find metals from the platens themselves.  So, we do it on sterilized, finished devices from the package because that is what would go into a patient.

DR. MILLER:  Is it possible that a device subjected to the environment in vivo would have a different profile of release of trace elements?

MR. YANG:  It is possible.  The problem becomes when you try and analyze for them, you then have to somehow correct for what the patient contributed to that device.  That becomes very complicated.  We tried to do that in some cases but the techniques that we use are not designed--things like proteins get in the way.  So, it is hard to do.  It is not impossible to do that; some people have tried to do that but the question is how good is the analysis, and that has always been a question.

DR. MILLER:  Thank you.

DR. WHALEN:  Dr. Choti?

DR. CHOTI:  Just a couple of questions, Mr. Kline.  One, just to clarify again the follow-up of the two different cohorts that you looked at, you showed a five-year follow-up rate in the new analysis of 60 percent, 50-60 percent on average I think, depending on whether it was reconstruction or augmentation.  What was the three-year follow-up given the new follow-up that you have, a better follow-up, and how does that compare to the original PMA data?  In follow-up to that, when you compared the three-year to the five-year, why did you use the new follow-up data rather than the original PMA data?

MR. KLINE:  The follow-up in the PMA was approximately 70 percent for both cohorts.

DR. CHOTI:  Three years?

MR. KLINE:  Yes, at three years, the PMA submission at three years.  The analysis that we did does include all data, whether it is five-year or updated three-year.  So, the analysis is on all the data that we have.  The numbers for the PMA and the numbers that I showed for the three-year columns were just for consistency's sake.  The three-year numbers submitted in the PMA just varied slightly because of some new patient data that we got.

DR. CHOTI:  The other question is regarding the textured versus the smooth.  In the clinical data, did you look at differences in deflation, capsule and reoperation between the two groups?  Also, what were the relative percentages of the two types?

MR. KLINE:  We did look at the difference in deflation rates between smooth and what we call Siltex or textured product.  The rate for smooth products, for all products was approximately 5-6 percent and the rate for Siltex was approximately 11-12 percent.

DR. CHOTI:  As the population as a whole of the two types, what did you have in the group?

MR. KLINE:  We presented that in Dr. Michael's first slide, 30 percent smooth and approximately 70 percent textured.

DR. CHOTI:  And how was that changed near the end of the trial or more currently?

MR. KLINE:  You are exactly right, with the current mix of product, with the product mix that we had in the study those were the rates.  You know, doctor preferences have changed.  At this point it would be appropriate to bring up Dr. Freedman to discuss the current preference mix that he uses in his own practice.  As Dr. Michael was explaining before, Dr. Freedman is a clinical instructor for plastic and reconstructive surgery at George Washington University.  He is a clinical assistant professor for plastic and reconstructive surgery at Georgetown University, and a consultant to the Department of Plastic and Reconstructive Surgery for the NIH.

DR. FREEDMAN:  My name is Roger Freedman.  Approximately 35 percent of my practice is breast surgery

DR. WHALEN:  I am sorry to interrupt, but could you identify with the questions?

DR. FREEDMAN:  Oh, I am sorry.  I am not involved in any suits.  I have provided my own travel today.  I do put in breast implants and I am also involved in the core gel breast implant study provided by Mentor.  So, I do get some compensation for that study.

My practice consists of approximately 35 percent breast surgery, which encompasses all aspects of both cosmetic and reconstructive breast surgery.  My implant usage is pretty consistent with that of the norm for the nation, which is probably 98 percent to 99 percent smooth round, with a rare case of using textured today.

I think people have appreciated that there is more of an incidence of rippling in the textured implant which, therefore, then gives a potential for fold failures which you were addressing earlier and, therefore, by switching over to the smooth round and placing these implants under muscle, that muscle is then providing pressure on the implant which has a tendency to help smooth out that implant even more.  That is consistently the norm today.

The issue then is the issue of filling an implant.  There are recommendations, nominal recommendations by the manufacturer.  It has been appreciated that if you under-fill an implant to maintain a softer implant there is a higher incidence of rippling and, again, a higher incidence of fold failure.  So, that is not the current norm so people typically, including myself, fill them to their nominal value or slightly overfill those implants, but not beyond the recommendations of the manufacturer.  The numbers that I personally experience in my practice are less than those which are quoted in this report.

DR. CHOTI:  So, can you tell us again the difference of smooth versus the textured in the study as far as the complication rates?

MR. KLINE:  We stratified for deflation, and for deflation smooth is approximately 5-6 percent deflation rate, whereas our textured product, Siltex, is 11-12 percent through the five years.

DR. CHOTI:  Other parameters?  Reoperation, capsular formation?

MR. KLINE:  I don't have that information available right at the moment, but we could certainly look at it and provide it to FDA.

DR. WHALEN:  Dr. Chang?

DR. CHANG:  Even though the data was not presented, can anyone give me an answer with regard to the explanted implants that were examined?  Was there any relationship between thickness of the shell in those that were explanted and that did not have surgical sharp instrumentation?

MR. CROUTHER:  Not within a population of a given type.  You know, our textured devices are thicker, for example, than our smooth devices, but I am assuming you are talking just about the textured, did those where we had fractures exhibit lesser wall thickness?  And, there was no evidence of that, nor was there any correlation with physical properties.

DR. CHANG:  My second question is that previous testimony from public comment just mentioned some questions about investigators about good manufacturing practices so I would ask the sponsor if there are any outstanding questions or communication with the FDA regarding improvement in good manufacturing practices that are outstanding at this time.

MR. CROUTHER:  I will ask Clark Sheriff, from Mentor, to discuss that.

MR. SHERIFF:  Good morning.  I am Clark Sheriff, vice president of regulatory compliance for Mentor Corporation.  At this point there are no outstanding issues with the FDA.  The last inspection was this last February.  There was a comprehensive GMP inspection by the agency, in Dallas, and the few issues that they brought up have been all addressed satisfactorily.

DR. CHANG:  Thank you.

DR. WHALEN:  Dr. Dubler?

DR. DUBLER:  If I heard the previous discussion correctly, the textured implants have twice the deflation rate, approximately twice the deflation rate as the smooth.  Is that correct?   Is there any point in the brochure that makes a statement that says our textured implants have two times the deflation rate as the smooth?

MS. CRAWFORD:  No, that is not currently in the brochure.  Most of this information was developed after the brochure was printed.  We can certainly work with FDA in determining what is appropriate to add to the brochure at this point.

DR. DUBLER:  So, there is no statement that reflects those data?

MS. CRAWFORD:  That is correct.

DR. WHALEN:  Thank you.  We will now proceed to the FDA's presentation with Ms. Allen and Dr. Dawisha.

FDA Presentation

MS. ALLEN:  Good afternoon.  FDA will now summarize the status of the conditions of approval for Mentor' saline-filled breast implant PMA.  For your convenience, we have provided you with a hard copy of FDA's slides.

There are five conditions of approval: post-approval study; a focus group study; a retrieval study; fatigue testing; and shelf-life testing.  Dr. Sahar Dawisha will present the status of the post-approval study and I will present the status of the remaining four conditions of approval.  I will now hand it over-expression to Dr. Dawisha.

DR. DAWISHA:  It is still morning so I can say good morning.  I am a medical officer and I had reviewed and presented the breast implant information back in March of 2000.  As you recall, at that time one of the conditions of approval was that the sponsors provide long-term safety information on their products.

To meet this condition, Mentor Corporation has been conducting a post-approval study which is an extended follow-up of the patients in the saline prospective study, which was originally designed as a three-year study, out to ten years postop in an abbreviated protocol.

As you just heard this morning, this protocol consists of annual mailing or a physician visit, and the endpoints of interest in the study include implant deflation, implant removal and reason for removal, additional surgery and reasons for surgery, presence and grade of capsular contracture, and breast pain related to implants.

As just discussed by Mentor, the protocol was approved in May of 2000, which is when the PMAs were approved.  Mentor began contacting their investigators in July of 2000, asking them to contact their patients or allow the sponsor to contact them.  The initial patient mailing began in October of 2000.  The database that we are going to be discussing today was closed in March of 2002.

It is FDA's goal to update the patient labeling, the physician package insert and the summary of safety and effectiveness, which I will refer to as the labeling, every few years to reflect the current complication information.  We plan on updating the labeling to reflect the five-year data that I am going to be discussing in the next few slides.

Before discussing the post-approval study or PAS study results, I would like to briefly review the saline prospective study patient accounting.  Recall that the saline prospective study served as the basis for the PMA.

The number of patients living and with implants by the end of the saline prospective study, as shown on this table, was 1252 for augmentation and 375 for reconstruction.  By the end of the saline prospective study and before the start of the post-approval study in October--so, there is about a five-month period there--there were a few additional patient deaths and implant removals, making 1250 augmentation and 351 reconstruction the number of patients available for participation in the post-approval study.  You will see these numbers in a subsequent table.

With that background, we can now discuss the patient accounting for the post-approval study patients over time, which is shown in this table for augmentation.  Based on the actual follow-up, divided by the expected follow-up where we define expected follow-up as the theoretically due minus deaths and removal of all implants during the interval, the follow-up rate is shown here for five years through ten years postop.  For example, the follow-up rate at five years is 5 percent, 24 percent at 6 years, 45 percent at 7 years, etc.

Because some patients had exceeded their five-year follow-up visit at the time of the start of the PAS and because the sponsor has recently improved their efforts to contact patients, the follow-up rates beyond six years are superior to that at five years.

The bottom row of this table shows the number and percent of patients with any data at any time, where a returned questionnaires are counted for all previous time points.  For example, using this method the follow-up rate at five years is 54 percent at five and six years.

The patient accounting information for the reconstruction patients in the PAS study are shown on this table.  As you can see, the follow-up rate for reconstruction patients is superior to that for augmentation patients in the previous slide.  For example, the rate at five years is 52 percent, at six years 59 percent, and at seven years 54 percent.

Because of the low follow-up rate, FDA has been working with the sponsor to improve patient contact efforts, which are summarized in this table.  Of the patients expected for participation in the PAS, which is shown in row 1, some patients were excluded, in row 2, primarily because a few physicians didn't want to participate in the PASS, which reduced the number of patients to whom packets were mailed, which is shown in row 3.

Taking augmentation as an example, of the original cohort of 1250 patients, approximately half, which is shown in row 12, have agreed to participate in the PAS, and 777 patients have some data reported at least once in the PAS, which is shown in row 15.

Of the 351 reconstruction patients, approximately two-thirds of the original cohort, which is shown in row 12, have agreed to participate in the PAS, and 265 have some data reported at any time, which is shown in row 15.

Rows 6 and 7 are of interest because they reflect the patients whom the sponsor is pursuing to continue to contact.

Because of concerns with missing information and how this would impact the results, the sponsor was asked to determine whether and to what extent there was bias in the results due to missing information from patients who were considered non-responders, and whether and how the complication information should be adjusted to reflect missing information from these patients.

The results of this analysis are currently under review by us, and additional information has been requested and is ongoing to further clarify this issue.

What I can discuss are the results of preliminary analysis to answer these questions in which the baseline demographic and surgical characteristics were compared between patients who were and were not responders.

There were no significant differences with respect to race, ethnicity, marital status, incision size and obesity, defined as a body mass index of greater than or equal to 24.

For augmentation, there was a significant difference with respect to age at the time of implantation, with responders older than non-responders.

For reconstruction, there were significant differences with respect to surgical approach, implant placement and surface texturing.  Again, these analyses are still being developed.

I will now discuss the safety results from the data that we have available from the responders, first for augmentation and then for reconstruction.

Table 5 summarizes the cumulative Kaplan-Meier risk rate of first occurrence and the corresponding 95 percent confidence interval of complications at three and five years for the augmentation patients.  These values represent the cumulative risk of having the first occurrence and do not capture repeat occurrences.

The three-year rates shown in this table reflect the data from the saline prospective study that appear in the current labeling, and the five-year rates are an update based on the post-approval study.

While the cumulative risk rates at five years are higher than at three years for all the complications shown here, it is only for reoperation, implant removal and implant deflation where the confidence intervals are not overlapping, suggesting that the rates at these time points are significantly different.

I would like to point out that we had not asked Mentor to do any statistical analyses comparing whether the rates were different at one time point to another.  However, they have a slightly different interpretation where they found an increased risk only for implant deflation and a decreased risk for capsular contracture.  They reached these conclusions based on a separate analysis where the risk of each complication was estimated over time for patients who had not had the complication, and then linear regression was used to determine whether there was a statistically increased or decreased risk at five years.

Again, we did not ask Mentor to do these analyses and the Kaplan-Meier shown here are what are in the current labeling and will be included in the updated labeling.

The reasons for reoperation in the augmentation patients are shown in this slide based on the number of reoperations.  If more than one reason was reported, then all the reasons are included in this table.  Note that the current labeling reports the types of reoperation procedures, again, because these were physician elicited responses rather than the reasons for reoperation, which is what was elicited on the questionnaire.

The three-year information shown here was provided by the sponsor at our request to show the progression from three to five years.  I have combined the rows for the purposes of projecting the slide, as described in the footnotes below the table.

Through five years there were 343 reoperations or additional procedures reported in 198 patients, involving 312 implants.  On a per reoperation basis, the three most common reasons for reoperation at five years are cosmesis, and this includes wrinkling, ptosis, asymmetry and aesthetic revision, approximately 36 percent; followed by patient request for a size or shape change, 28.6 percent; and followed by leakage/deflation, 19.2 percent.

The primary reason for implant removal at three and five years is shown for augmentation in Table 7.  If more than one reason was reported, the same hierarchy was used as that reported in the current labeling.

Through five years, implant removal was reported for 211 implants in 132 of the augmentation patients.  The four most common primary reasons for implant removal at both three and five years was patient request for a size or shape change, approximately 30 percent, which was equal to the number of leakage/deflation, also approximately 30 percent; followed by asymmetry, wrinkling, ptosis or scarring, 18.5 percent; and then followed by capsular contracture, 14.7 percent.

Moving on to the reconstruction information, the by-patient Kaplan-Meier values are shown in this table, with the three-year rates based on the current labeling and the five-year rates representing the updated information.

While the risk rates are slightly higher at five years compared to three, the confidence intervals are overlapping for these time points, suggesting no differences.  Note that the overlapping confidence interval is least superimbosable for the complication of implant deflation compared to the other complications.

The reoperation information, which excludes planned procedures, is summarized for reconstruction patients in Table 9.  As with augmentation, the three-year information was recently provided by the sponsor and is not in the current labeling.  Through five years there were 232 reoperations reported in 162 patients, and occurring with 196 implants.

On a per reoperation basis through five years, the three most commonly reported reasons are cosmesis, which includes wrinkling, asymmetry, aesthetic revision and ptosis which is 31 percent; followed by capsular contracture, approximately 29 percent; and followed by a scar or wound revision, 25.4 percent.

The implant removal information is shown in this slide for the reconstruction patients.  Through five years, implant removal was reported for 135 implants in 112 of the reconstruction patients.  The three most common primary reasons at both three and five years is capsular contracture, leakage/deflation and infection.

This concludes the post-approval study presentation for Mentor and Ms. Allen will now continue with the focus study results.

MS. ALLEN1:  The ultimate goal of the focus group study was to improve the existing patient brochure.  Mentor already described how an independent study was conducted to obtain feedback regarding their patient brochure.  They also described some of the key findings from that independent study.

FDA considered the independent study reports submitted by both Mentor and Inamed and required the same types of changes for both companies, if applicable.  The primary changes to Mentor's patient brochure were as follows:

They made significant modifications to the lead-in as well as to the content of the safety tables because most women had difficulty in understanding the safety data.  They stratified augmentation and reconstruction information, and they added a table of contents and a glossary.

Mentor incorporated all requested changes into the patient brochure and received FDA approval.  Therefore, FDA considers this condition of approval fulfilled.  Mentor has just submitted a revised patient brochure and package insert that reflect five-year post-approval data.  After FDA review and approval of this supplement, Mentor will finalize them for public and product use.

The purpose of the retrieval study is to determine modes of failure.  This information may lead to changes in manufacturing design specifications, mechanical testing requirements, and/or labeling.

In their 2001 report, Mentor submitted limited data on 38 explants collected over a four-month period.  Mentor provided clinical or physician observations collected at the time of explantation.  They provided laboratory observations, or device failure characteristics, such as smooth and sharp crease-edge opening.  These were noted with respect to whether the device was deflated or non-deflated.

Based on the limited number of retrieved implants, Mentor made no conclusions regarding whether the device failure characteristics were representative of a true failure or the result of an artifact, such as shipment, excessive handling or the method of explantation.  Accordingly, no hypotheses regarding modes of failure were provided in that report.

Mentor submitted a final report of the retrieval study which is under FDA review.  Therefore, FDA considers this condition of approval still open.

The purpose of the fatigue testing is to determine the fatigue strength of Mentor's product line.  These data provide additional information on the expected long-term performance of the device.  T r are 12 styles across the saline-filled and Spectrum families.  Mentor chose styles 1400, 1600, 2600 and 5000PT as representative of their entire product line.

Mentor completed fatigue testing on three of the four styles.  The resulting endurance load limit was 10 lbs at 10 million cycles run-out for those three styles tested, which did meet the acceptance criteria.

As part of the test report, Mentor also supplied the ultimate static rupture results for those three styles.  The results were over 600 lbs for all three styles, which shows that the implants failed at static loads much greater than that expected during mammography, which is about 55 lbs.

FDA expects Mentor to submit fatigue test results for style 5000PT as part of their 2003 annual report.  Accordingly, FDA considers this condition of approval still open.

The purpose of the shelf-life testing is to support a five-year expiration date on the package label.  Mentor's shelf-life protocol involves real-time package integrity and mechanical testing performed at year zero or baseline at years four and five.

In their 2001 annual report, Mentor provided an interim report with year zero data.  The results were adequate.  However, this report did not include style 5000PT data.  An updated report of year zero testing with style 5000PT data has been submitted and is under FDA review.

FDA expects Mentor to submit an updated report of shelf-life testing year-four data in 2005 and year-five data in 2005.  Accordingly, FDA considers this condition of approval still open.

This is an overall summary of Mentor's five conditions of approval.  The post-approval study will remain open until ten-year data are provided.

The focus group study is complete.  Mentor has already revised their patient labeling to reflect the focus group study findings.

The retrieval study is currently open, however, Mentor has submitted the final report and it is under FDA review.

The fatigue testing is complete on three of the four styles.  Testing on the fourth style is expected to be submitted in the 2003 annual report.

The shelf-life testing will remain open until five-year data are provided.

I will now turn it over to the panel for discussion.

Panel Discussion

DR. WHALEN:  Thank you.  Questions of Ms. Allen or Dr. Dawisha from the panel?

DR. DEMETS:  I have a question.

DR. MCCAULEY:  I do too.

DR. WHALEN:  Dr. McCauley, I heard you first.

DR. MCCAULEY:  Based on the focus group study, initially the brochure revealed that one out of four of the focus group participants did not understand the data, or it was confusing, and the brochure was revised.  But this was not taken back to the focus groups to see if they understand that data before it is approved?  Is that not true?

MS. ALLEN:  That is correct.  The protocol was for that first feedback on the existing brochure, approved back in May, 2000.  There are no plans right now to go back and conduct a second focus group study.

DR. MCCAULEY:  Then how do you know if the consumer understands the data?

MS. ALLEN:  Good question.  We worked with ODE and OHIP, Office of Health and Industry, in order to provide input on making it more in layman's terms.  I don't know, maybe Mentor can provide more input on that.

DR. WITTEN:  I would say we think we addressed the issues that were raised by the focus groups.  As Ms. Allen mentioned, we have someone in another office who looks at patient labeling specifically for us to address issues related to it being understandable to the patient or to the consumer.

DR. WHALEN:  Dr. Doyle?

DR. DOYLE:  I am concerned about the follow-up rate, whether the data is even viable given the low follow-up rate.  In addition to that, 20 percent of the patients have reimplantation and they are considered in the group that goes out to five years rather than starting over as a new implant.  Wouldn't that give falsely low complication rates at five years because these women may have only had the implant at one, two or three years and their second implant, yet, I think I understood that they are within the group that has been carried out to five years.

DR. DAWISHA:  If a patient undergoes a revision, gets their implants removed and gets a second set of implants, as far as I know, they are not included in the Kaplan-Meier values because they are considered revision patients and we have reported them separately in the labeling, as a separate group that gets complications following revision.

DR. DOYLE:  Perhaps I misunderstood but I thought when I asked Mentor this question they said they were left back in the data.  Maybe I misunderstood.

DR. DAWISHA:  Well, they continue to follow the patients and they report the study results to us, but they are in a separate table in the labeling.  They are reported as patients who have complications following revision.

DR. WHALEN:  Dr. DeMets?

DR. DEMETS:  Yes, I am bothered by the response rate.  I appreciate the challenges and difficulties of follow-up in any study, any clinical study that is done.  Having said that, I am still troubled I guess as much as my predecessor Dr. Blumenstein was by the response rate, and what you can make of any kind of strong conclusion based on that kind of response rate.

There has been some comment by both the FDA and the sponsor about comparing responder and non-responder at baseline.  That is nice.  Maybe it is necessary, but it is hardly sufficient because the real assumptions have nothing to do with that.  The real assumption has to do with the outcomes.  Are the outcomes in responders the same as in the non-responders?  Suppose it turned out that the patients who didn't respond all had bad experiences and were so angry that they threw the questionnaire in the trash, that is the kind of bias that you worry about and until you can begin to address that you never know with the non-responder is introducing a bias or not.  So, that is why it is such a problem.

That is why in good clinical trials the response rate that you go after is in the higher 90's because nobody believes that non-response is independent of what is going on.  If we did that, we wouldn't pursue mortality trials to have almost no loss to follow-up.  So, I am really troubled by this, much, as I said, as Dr. Blumenstein was.  I think this issue of comparing responders and non-responders at baseline is nice.

The second point about that is when you do any kind of an adjustment procedures, logistic regression or anything else, they add a little bit but, you know, they don't explain much.  If those regression models explained a lot of the outcome, then those adjustments would be meaningful but they don't explain much, I am guessing.  We don't know; we didn't hear that data, but they don't explain much.  Therefore, the adjustments are sort of a very modest correction and, you know, analysis can never correct for flawed designs.  We need good design and a low response rate is one of those kind of design problems we all wrestle with in a trial.

So, I am very troubled by this response rate.  But I guess my question is, given the passion that the patients undergoing this procedure have, it seems like we ought to be able to get 99 percent response rates.  So, I am puzzled by why we are where we are.

DR. DAWISHA:  Well, I would just like to add I think we share your concern.  I am glad you brought up the issue of the responders and non-responders and whether or not they have suffered a complication.  I had indicated that the analyses are under way and we actually asked the sponsor that specific question, whether having a complication at an earlier time predicted whether someone was or wasn't a responder.  I mentioned in my presentation that the analyses are continuing.  That is one of the analyses.  I guess what we do with that data and how we adjust the data or how much sort of remains to be seen.  But we certainly share your concern as well.

DR. WHALEN:  Dr. Choti?

DR. CHOTI:  Dr. Dawisha, it still is not clear to me.  You mentioned certain things that you asked Mentor to provide you or certain analyses.  Did you specify in this what the response rate should be, what your minimal expectation for a follow-up is?  I guess the bottom line is are you happy, are you satisfied with the data that you have gotten?  You say you are currently working with it, reanalyzing it so we don't have the final but is this kind of sufficient to address the concerns of the panel in 2000, the data you got?

DR. DAWISHA:  Well, I think our goal is to update the patient labeling and the labeling information, and we would like the labeling to be the most valid information that we have.  Certainly, you know, we are not happy with the low follow-up rate.  We are sort of stuck between do we put this information in the labeling and update it, or do we not put it in.  I think our position is that we would rather at least put the information that we have in the labeling, explaining what the limitations are, just so that at least that information is available to patients.

DR. CHOTI:  I guess my question though is were there some specific metrics that you requested in this follow-up data set, and were those met, follow-up being one?  I mean, it is possible that they could ramp up the effort even more to get better follow-up data if necessary, if it was specified that this is the information we need at this follow-up time.  Was there some specific information requested of them?

DR. DAWISHA:  There were no specific follow-up target rates that we had asked the company to follow.  The breast implant guidance document has some general guidelines, one of which is that we expect follow-up to be about 60 percent at ten years.  That is based not just on this product but several other types of products.  We are lucky to get 50 percent follow-up at ten years for those types of studies.  So, no, there was no set follow-up rate.  There was no target set for the sponsors.  We certainly would like to see high follow-up rates out to ten years but we may not.

DR. DAWISHA:  Dr. Dubler?

DR. DUBLER:  Back to my current interest, which is this reoperation rate.  On Table 8, which is the Kaplan-Meier risk rate for reconstruction, the five-year risk rate, as I read it, is 43 percent.

DR. DAWISHA:  That is right.

DR. DUBLER:  In your revision of the informed consent document that you were provided, is that rate stated in that document in any way?

DR. DAWISHA:  You mean the patient labeling?

DR. DUBLER:  No, in the brochure, the company brochure.  It is my sense that patients read labels, the really savvy ones, but a lot of them will rely on the company brochure, and after the focus groups you have seen a revision.  Correct?

MS. ALLEN:  The company has already submitted the supplement with the labeling with the focus groups findings and we approved that.  But they have just recently submitted an updated package insert and the patient brochure with the five-year post-approval data.

DR. DUBLER:  And does that say 43 percent of patients who have reconstruction will have surgery in the next five years?

MS. ALLEN:  Yes, it does.

DR. DUBLER:  It does?

MS. ALLEN:  Yes.

DR. WHALEN:  Seeing no other questions at this point, thank you.  We will go into the concluding discussion.  I would like to do this by asking everyone at the table to make their comments about what we have just heard and, since you were last at the microphone, we will start with Dr. Dubler.

DR. DUBLER:  This has been a puzzle to all of us for the last decade.  Clearly, a woman does have a right to choose.  The problem is that she has had the right to choose between procedures that don't appear, by their own data, to be terrific.  So, then the right to choose has got to be supplemented by information that makes that choice a truly informed one.

I am impressed, someone gave us the statistic that the industry had spent six million dollars supporting one of the industry organizations.  We did hear that figure, didn't we?  That is a lot of money.  And, I would think with one fraction of that money and some creative people we could come up with a process that would bring together various independent groups that analyze the data and the company that analyzes its own data, and agree upon the data that a woman had to confront in a form that wasn't the Kaplan-Meier adjusted risk rates, which no lay person, quite frankly, understands, and come up with a script which could then be put in an interactive video format.  We could think of lots of creative ways of making an informed consent process independent of the physician, who is likely to say, "yes, of course, I have to give you this but, hey, I do this all the time; I think it is terrific.  I gave my 16-year old breast implants for her birthday."

There are ways of ensuring that this prospective patient, indeed, focuses in the most constructive circumstance on the data that she has to consider.  Once she has considered that, so be it.  But there are ways of presenting it, of designing it, of agreeing on what has to be presented that I think we are technically capable of doing.  I think that a written brochure, handed out by the surgeon who does this and thinks it is really quite a good process, is a very 18th century way of thinking about what to do.  We could do a lot better, and I would hope that the FDA would work with the companies to devise really creative, new, independent solutions to the problem of helping a woman confront and analyze the data in terms of her own values.

DR. WHALEN:  Dr. McCauley?

DR. MCCAULEY:  I agree.  I think the entire issue we are dealing with here is about choice.  In order to really make a proper choice you have to have informed consent, and that informed consent has to be based on reliable data in terms of complications or problems that the patient may experience.

I think that when we start looking at data, herein lies the problem and I can appreciate the fact that certainly for implant surgery, not only in plastic surgery but even in orthopedics, the follow-up can be not quite as reliable as some patients with cancer problems, as you have mentioned.  However, I think it is imperative that we really try to push the envelope here to try to get as much follow-up as we can in terms of percentage on these patients that have undergone these types of procedures.

The other issue relates to integrity and perception.  I think this has been a problem of some consumer groups that have spoken to us earlier this morning.  I am happy to hear that the issues related to good manufacturing practices have been addressed and this is no longer an issue.

However, I think it is imperative that the FDA continues to work closely with Mentor to try to improve the data collection, and also to improve the statistical analysis.  Certainly this issue of responders versus non-responders is a crucial issue I think, and I am not sure how that can be resolved but I think it is something that we really need to take into account.

The last point I wanted to make really relates to consumer education, and I think that if we take on a project or the FDA takes on a project that states that what you have to do is appropriate for a study, then for closure would be to take that information back to the focus groups to see if they understand truly what the risks are.  It is a matter of concern, having a proper consent and information for the patient to make a proper choice.

DR. WHALEN:  Dr. Doyle?

DR. DOYLE:  I resonate with the idea that the breast implants really represent a choice for women.  I have a sister who went through reconstruction surgery for cancer so I know that this is a very important choice for women to be able to make.  I also believe that the informed consent process needs to be based on accurate information and that the important thing is that the informed consent contain all the information to make a choice.

I am actually not sure an interactive video is going to alter the fact that once a woman understands what the risks are that they are still not going to make this choice.  I don't think we can protect people from wanting to do something once they understand what the risks are.  So, I believe the informed consent process is the key here.

DR. WHALEN:  Dr. Miller?

DR. MILLER:  If I could just try to provide some perspective on this, when I look at what is going on, we have an unusual problem with the breast implant in that we are trying to balance a benefit/risk ratio where the benefit is extremely subjective to the patient, and even within a particular patient it can change over time.  So, getting a handle on that side of the equation is very difficult.  The risk is difficult to get a handle on.

It is difficult to study this.  You can't do it as nice and cleanly as you can a drug where you have a limited number of patients and a limited number of follow-up time.  It is difficult to conduct a study for ten years in patients who are basically healthy, patients who don't want to consider themselves patients, who want to disappear into the landscape as soon as they have their implants.

I mean, these are all challenges to doing this.  I think that we have to be careful in how we look at addressing this in terms of methods that we have become very comfortable with in looking at other problems.  That is one thing.

On the other hand, we have to do the best possible job we can do.  I share everyone's sentiment about the inadequacy of the studies, the follow-up.  Those need to be done in an unassailable fashion so that this is laid to rest, so that we are not discussing this five years from now.  I think we are all weary of discussing this, and I think a study can be designed where we address this.

I don't know what the follow-up rate is for pacemakers but I know it is pretty darned good.  I mean, if you have a patient with a pacemaker that was put in 15 years ago, you call up a number and the person on the other end of the line practically knows the entire life history of that person who had that pacemaker.  Maybe that is an extreme example because of the nature of that device, but I think that something needs to be done to not permit three physicians to say I won't participate in this study.  I mean, if they won't participate maybe they shouldn't have access to the implants.  I mean, if you have the implant you have to, as a responsible citizen of the medical community, participate in confirming how these implants are used and whether they are safe or not and effective.  If you don't want to participate in that, then perhaps you shouldn't place any implants.

I think those sorts of things could possibly be done.  But I want to avoid the tendency to treat this like other medical problems because it is very unique.  I guess that is my main philosophical point.

DR. WHALEN:  Dr. Chang?

DR. CHANG:  I think two years ago the advisory panel and finally the final decision of the FDA was to keep the option of having saline implants available to women for augmentation or reconstruction.  It wasn't a decision made without controversy, and it is certainly on the record that there have been many reservations regarding process, statistics and other questions which have led to the conditions for premarket approval.

Today is an update of presentation of sponsors trying to meet these conditions of approval and they are in process.  I would urge sponsors to continue to examine the physical properties of their product to try to decrease the preliminary studies of safety, to continue to make improvements and weigh what are the benefits of lessening implant failure with perhaps increasing thickness, comparing to the benefits of pliability and cosmetic effect and perception of pliability.  So, that is one question and request that I would make of the sponsor.

The other question, I would second Dr. McCauley's comments regarding the focus group.  Again, follow-up is necessary to be sure that what we presume to be understandable by the lay public is, indeed, understood in terms of reading information provided to them so that they can make that informed decision.

I am astounded, but maybe I shouldn't be, that reported comments by members of the focus group felt that these statistics were not true; that these numbers were merely presented as a disclaimer by the manufacturing and that this really couldn't happen and won't happen as an individual makes a decision of choosing the saline implant.  So, there is some issue of credibility in terms of presenting this data.

However, I feel that statistically speaking in terms of percentage of follow-up being on the far extreme of what normally is acceptable in a scientific clinical study, it is some data to give us an idea of what happens to these implants after three and five years.  So, there is some data and, despite good efforts of sponsor to collect this, perhaps higher remuneration, perhaps a larger bonus for returning the survey might increase the numbers.  Be that as it may, we would say these are scientifically suspect data just because of the numbers of responders, but there are some numbers and I think that they should continue to be made available to consumers.

Finally, my comment is that, yes, continued efforts should be made to carry out to ten years to find out what happens with these implants, but ultimately I don't think we are really going to get an answer until there is a report by an independent registry that has, again, good follow-up on the people who are registered to see what happens to these implants.

DR. DAWISHA:  Dr. DeMets?

DR. DEMETS:  Well, I have already said some of the things I have been thinking about.  I am obviously not going to quarrel with the issue of the right to choice and assess the risk/benefits, but I worry about information that those decisions and thoughts are based on.

I would only urge I guess the FDA and the panel to appreciate how sensitive results can be to issues of non-response, to excluding patients, to censoring patients from follow-up.  I am a professional statistician and I am still amazed and astounded at times at the power of these biases that work in data.  So, I think that we really need to set a higher standard for something that has been so controversial, so much discussed and is so important.

I think that if we don't do something different than we are doing now, at the ten-year mark we will still be arguing just as much as we are arguing now.  Sixty percent as a target at ten years, I guaranty you, is going to bring more controversy and discussion because 40 percent non-response is overwhelmingly an opportunity for bias to enter into any kind of results.

The trouble with numbers, if you produce tables such as we have today and put 95 confidence intervals on them, is that they take on a credibility that they sometimes don't deserve because of the power of the biases that are at work.  So, I am not saying we should not present the data that we have, but there is a danger to it of misleading the perhaps less sophisticated readers of those tables.

But I would go back to where I started, that is, the response rates that we are observing here and perhaps other devices are simply not acceptable if the decisions really depend on the data that comes out of those surveys.  So, we may have to challenge ourselves across the board and certainly in this arena with all the interest and importance this has to do a lot better than the target we set for ourselves.

DR. WHALEN:  Dr. Newburger?

DR. NEWBURGER:  As my colleagues at the panel have expressed their discomfort with what has been presented, and have constructive suggestions on how to ameliorate the situation, I join their concerns.  Augmentation mammoplasty is the second most frequently performed plastic surgical procedure in the U.S.  Something upward of 300,000 were done this last year, 2001.  It is mind-boggling to me that we have so little data available on a procedure that was done in 300,000 individuals in one year.  The rate is going to increase geometrically as it has been.  This is from the American Society of Plastic and Reconstructive Surgery web site, by the way.

So, I would hope that the efforts to get more accurate follow-up on increasing the numbers of patients in this study would be done to the point of maybe sending investigators to really look for the patients.  I am also concerned that the products appear to have a high rate of product failure.  I think that the high reoperation rate wouldn't be acceptable in other type of prosthetic devices.  I question two standards.  Perhaps because this is considered to be cosmetic, it is not held to the same exactitude that other prostheses are.

DR. WHALEN:  Dr. Choti?

DR. CHOTI:  I echo the comments of my fellow panel members.  It was clear in the review two years ago there were concerns when this was approved, and I commend the FDA and the industry on working on the conditions of approval.  I think we have learned a lot based on the data of these conditions.  Yet, it is astounding, particularly the post-approval study, how weak the data is still.  As we heard, in spite of how common this procedure is, we still are faced with a procedure and device that clearly is important, and it is clear I think that we do not have sufficient data or good enough data.

So, I really do echo what has been said, that coming up with better registries, independent data collection so that we really can have the real answers rather than the data we have been working with today, and I am concerned that it is not going to be improving that much with continued follow-up.

It is interesting that these products have problems.  I mean, we are learning that there is device failure.  The performance is not as good as we would like to see.  So, I think there is clearly a mandate to continue to improve this product.  These devices are placed in healthy individuals.  So, I think it is very important to continue the effort to improve the product which is clearly fraught with problems.

So, I think I really do echo the statements that have been made both regarding informed consent and the data that is available.

DR. WHALEN:  Dr. Witten, for FDA's purposes has this discussion been adequate?

DR. WITTEN:  Yes, thank you.

DR. WHALEN:  We, therefore, will stand adjourned for lunch and at 1:30 sharply we will reconvene.

[Whereupon, at 12:40 p.m., the panel was adjourned, to reconvene at 1:30 p.m.]

A F T E R N O O N  P R O C E E D I N G S

DR. WHALEN:  I would like to welcome everyone back.  I would remind the public observers at this meeting that while this portion of the meeting is open to their public observation, public attendees may not participate except at the specific request of the panel.

We are now ready to continue our panel meeting with Inamed Corporations presentation.

Inamed Corporation

DR. EHMSEN:  Good afternoon, Dr. Witten, Dr. Whalen, members of the panel, FDA representatives and all those in the audience.  I am Ron Ehmsen, vice president of clinical and regulatory affairs for Inamed.  I should point out that what had previously been McGhan Medical Corporation is now operating as Inamed Aesthetics, which is a business unit of Inamed Corporation.

My colleagues and I are here today to present an update on several conditions of approval that were associated with PMA number P990074 which covers McGhan's saline-filled breast implants.  That PMA was approved by FDA on May 10, 2000 for breast augmentation in women over 18 years of age and for breast reconstruction.

There were five conditions of approval.  First was a post-approval study; second, a focus group study aimed at clarifying or trying to understand whether the patients themselves had any questions about the brochures that were provided to them to help them make a choice.  Third is a retrieval study.  Fourth is fatigue testing and fifth is shelf-life testing.

We will break up the presentation into several parts.  Dr. Audrey Weiss, our senior manager of clinical research, will present the results of the post-approval study and then we will move on from there with Kim Croyle, our senior regulatory affairs specialist, and Tom Powell, our director of technologies.  Audrey?

DR. WEISS:  Thanks, Ron.  Inamed's first condition of approval was to conduct a post-approval study to identify long-term outcomes associated with McGhan's saline-filled breast implants.

First I would like to review the data that formed the basis of the original PMA that was submitted.  The original PMA included data from three years of follow-up, three-year post-implant information from two five-year clinical studies.  Additionally, women enrolled in these studies also had begun to complete some of their three-year follow-up visits.  So, limited four-year data was also available at that time.

The two clinical trials were the 1995 augmentation study which, for short, I will refer to as the A95 study.  In that study, 901 patients were enrolled between 1995 and 1996 for primary augmentation.  The second study was the 1995 reconstruction study, which I will refer to as the R95 study for short.  That study had a very similar protocol to the A95 study and enrolled 237 patients for primary breast reconstruction.  Almost all of the patients enrolled in the R95 study had had mastectomy following breast cancer, and there was a handful of patients in that study who had had prophylactic mastectomy.  All of the patients in both the A95 and R95 studies had not had previous breast implants prior to enrollment.

The post-approval data collection is being conducted in two phases, and the objective of the study is to obtain long-term safety information through ten years post-implant on the same 1100 women who were enrolled in the A95 and R95 studies.  Again, at the time of the original PMA complete three-year follow-up information was available.  All women had completed three-year follow-up.

So, the post-approval data collection we actually are conducting in two separate phases.  The first phase has been completed, and involved continuing to follow those 1100 women enrolled in the A95 and R95 studies under the same protocols as those studies had been conducted under through three years post-implant.  Specifically, the protocol involved women coming in to see their physician for a follow-up visit in the office.  The four-year and five-year follow-up information that forms the phase one of the post-approval data collection was based on this method of data collection.  Complete five-year information is now available.  All patients have completed the five-year follow-up visit and the five-year data is what I will present today.

The second phase of post-approval data collection is currently in process.  This phase involves continuing to follow those same 1100 women who were enrolled in the A95 and R95 studies using a mail survey protocol that will follow them from six to ten years post-implant.  On the anniversary of their original implant surgery, patients will be sent a mail survey to complete regarding the status of critical safety outcome variables, including reoperation and implant leakage/deflation.  Again, that phase two is in process and we are currently mailing surveys to patients fort he six- to ten-year follow-up information.

The remainder of the presentation will focus on the five-year follow-up information from phase one.  First, I would like to present the follow-up compliance information for patients enrolled in the A95 and R95 studies.

First, what I have done here is actually included the follow-up compliance rate at each of the required follow-up intervals, which was annually through the five years in the A95 and R95 clinical studies.  The data that was presented for the original PMA was three-year data which was based on 83 percent patient compliance for the augmentation cohort and 88 percent follow-up compliance for the reconstruction group.

For the post-approval phase the follow-up compliance rate has remained at 80 percent or higher.  At five years the follow-up compliance rate was 81 percent for augmentation patients and 80 for the reconstruction patients.

Next, I would like to present the information obtained for specific local complications that were assessed in the A95 and R95 studies.  The specific complications that I will present are reoperations, implant replacement/ removal, leakage/deflation, capsular contracture, infection, and a variety of other complications that were included in the protocols and, for completeness, we have included them here today.  These include complications such as surgical-related outcomes including hematoma, seroma, and skin and nipple related complications.

I should point out with this complications that these are not additive.  The same patients who are included in the reoperation rate, for example, may also be included in implant replacement/removal since implant replacement/ removal is a subset of reoperations.  Similarly, a patient may undergo a leakage/deflation and also be included in the risk for implant replacement/removal.  So, they are independent risks and are not additive.

Again, the method used for data collection was a physical examination by a physician at an office visit, according to the original protocol for the A95 and R95 studies.  The analysis method utilized was a cumulative risk based on the Kaplan-Meier product limit method with 95 percent confidence intervals computed.  The cumulative risk that you will see in the following graphs is represented as a failure rate curve which, you will see, increases or stays level over time.  It will never go down because, as we add additional events in over time the risk can only increase.

What you will see in the cumulative risk curves is the summation of all events that occurred up to the particular time point being reported, and each of the time points that were assessed in the study is presented.

What I will do now is go through each of the specific complications, reoperations, implant replacement/ removal, leakage/deflation, capsular contracture and infection and report on the risk information obtained through five years post-implant.

First, this graph represents the cumulative risk of reoperations for the augmentation and reconstruction patients.  Again, the risk is being presented for each of the time points that were assessed in the study, and you can look at this as a cumulative risk through the time point presented in the graph.  The white line represents the cumulative risk for the augmentation patients and the yellow line represents the cumulative risk for the reconstruction patients.

For example, with reoperations the cumulative five-year risk for augmentation patients is approximately 25 percent and the cumulative risk of experiencing at least one reoperation through the five-year time point is approximately 42, 43 percent.

Next, what I would like to do is breakdown what types of reoperation procedures patients underwent.  The reoperations reported here include any type of operative procedure to the breast or chest area, for example, implant replacement/removal and biopsy/lump removal, for example, is included here.

The next graph will break down for all those patients who underwent reoperation what those reoperations were.   This pie chart represents the breakdown of all of the reoperations for augmentation patients.  What you can see here from the red wedge is that the largest proportion, the largest number of reoperations were implant replacement/ removal.  That could be removal with or without replacement of the device.

The second most common type of reoperation performed for the augmentation patients was a capsule procedure, for example, capsulotomy or capsulectomy.

For reconstruction patients, the breakdown of the types of reoperation procedures is as such.  Again, implant replacement/removal is the most common type of reoperation performed, followed by, in the purple wedge, scar revision or wound repair, and then capsule procedures.

What I will do next is drill down into this chart a bit and look specifically at the most common type of reoperation, which is implant replacement/removal, and look specifically at risk through five years of that particular reoperation and then look at reasons why patients undergo implant replacement/removal.

First, this graph represents the cumulative risk through each of the time points indicated, ending with the five-year time point, of implant replacement/removal for augmentation and reconstruction patients.  Through five years, the risk of experiencing an implant replacement/ removal was approximately 11 percent for augmentation patients and approximately 28 percent for reconstruction patients.

The next two graphs will look at these patients who have undergone implant replacement/removal and look specifically at why patients underwent the device replacement/removal.

First for augmentation patients, the most common reason why patients undergo replacement/removal is seen with the red wedge, which is the patient's own choice to change the size or the style of the device.  The second most common reason for device replacement/removal is leakage/deflation of the device.

For reconstruction patients the predominant reason for undergoing implant replacement/removal is capsular contracture, followed by the patient's choice to change the size or style of the device, and then leakage/deflation.

The next slides that I will present break out these particular complications that have been represented in these graphs and look specifically at the risk to patients of experiencing various complications, including leakage/ deflation, capsular contracture and infection.

This graph represents the cumulative risk through five years for augmentation patients and reconstruction patients who experienced a leakage/deflation.  As you can see, the risk is virtually identical for both augmentation and reconstruction patients, and is approximately six percent through the five-year time point.

For capsular contracture, the five-year cumulative risk for augmentation patients is approximately ten percent and approximately 35 percent risk through five years for reconstruction patients.

Next is the risk of infection following implant surgery.  For augmentation patients through five years the risk of experiencing an infection is approximately one percent and for reconstruction patients approximately five to six percent.

The remaining graphs that I will show for the local complications that were assessed in the A95 and R95 studies are actually summaries of the cumulative Kaplan-Meier risk curves that you see here.  We had another approximately 20 or so complications that were assessed in the A95 and R95 studies, and for completeness I have included them here.  Each one of them, you can imagine, has a Kaplan-Meier risk curve, just like those complications presented here.  However, for brevity, what I have done is include them on a bar chart that summarizes only the five-year risk rate, which would be the highest possible risk through five years.

For example, this graph includes six implant-specific complications, and what you see with the white bar is the risk through five years for augmentation patients, and the yellow bar, the risk through five years for reconstruction patients.

For example, the risk of experiencing asymmetry for an augmentation patient through five years is approximately 12 percent, and for reconstruction patients approximately 40 percent.

In discussions with clinicians, they have indicated that this is to be expected given that with reconstruction patients they are trying to match a reconstructed breast that has had mastectomy with a normal non-reconstructed breast on the other side.

Also reported here are cumulative five-year risks for capsule classification, implant extrusion, implant malposition, implant palpability and wrinkling.

The next graph reports the five-year cumulative risk for augmentation and reconstruction patients for various surgical-related complications that were assessed in these studies.  As you can see from this graph, all of these surgical-related risks occurred at well less than ten percent, actually under seven percent for the five-year cumulative risk.  Presented are delayed wound healing, hematoma, scarring, irritation/inflammation, lymphadenopathy, pneumothorax and seroma.

The next graph presents the cumulative five-year risks for various skin and nipple-related complications that were assessed in the A95 and R95 studies.  Presented are the risks for loss of nipple sensation, nipple paresthesia, skin paresthesia, skin rash, tissue/skin necrosis and breast pain.

The next graph that I will present looks very similar to this, however it differs in two very important ways.  Following implant replacement, we continued to follow patients and look at any outcomes following the replacement, and this graph looks at the cumulative risk of some specific complications following device replacement.

The two critical differences in this graph that I would like to point out from the ones that you have seen previously are, first, that the risk presented here is at three hears following replacement.  Patients were able to be revised any time through the five years.  So, limited follow-up information is available for patients, for example, who were explanted at year four.  We would only have one year of information following the device replacement.  So, we were only able to calculate a valid risk with the information available for three years following the replacement in the study.

The second difference in this graph is that, unlike the previous risk information presented which was on a by-patient basis, this analysis is based on a by-device or by-implant basis.  This was selected because patients could have one side revised rather than both sides.  So, it made the most sense to look specifically at an analysis by the replaced device.

The complications presented here are of second replacement removal following replacement removal.  The risk here, you can see, is approximately 18 percent three years following replacement for augmentation patients and approximately 28 percent for reconstruction patients; risk of leakage/deflation following device replacement; risk of capsular contracture and risk of infection following a replacement surgery.

Next, I would like to present information obtained on reports of breast cancer post-implant and connective tissue or autoimmune disease reports.  First for breast cancer, of the 901 enrolled augmentation patients, there was one post-implant report of breast cancer which occurred 27 months after implant surgery.  For reconstruction patients, there were 24 post-implant reports of breast cancer through five years post-implant.  All of these 24 reports occurred in patients who previously had had breast cancer, which was the reason why they enrolled in the study initially.  The cancer may have recurred in the same breast that originally had the breast cancer or in the contralateral side.

Next, connective tissue and autoimmune disease information, the method of data collection for connective tissue and autoimmune diseases was that a patient would self-report to her physician that she had a particular connective tissue or autoimmune disease.  Based on the self-report by the patient, the physician would attempt to contact a diagnosing physician, attempt to obtain a diagnosis by the rheumatologist for example.  If the physician was able to obtain a diagnosing physician's report we term that here a confirmed report.  In other words, there is a physician diagnosing the patient with the particular connective tissue or autoimmune disease.  If the patient self-report was not able to be confirmed by a diagnosing physician's report, we list that here as an unconfirmed report.  It is still a patient's self-report of the disease but the doctor has been unable to obtain a diagnosing physician's report.  That may be due either to the patient never obtained the diagnosis or is not able to be contacted, for example.

Additionally, and I do not report that here, there have been some cases of self-reports by patients that have been found to be false reports, where the patients initially reported a particular connective tissue or autoimmune disease and subsequently indicated that they actually either had a different type of diagnosis or did not have a connective tissue or autoimmune disease at all.

For the augmentation group, there were 7 confirmed reports of connective tissue or autoimmune disease and 13 unconfirmed reports.  Of the 7 confirmed reports, 3 were Grave's disease, 2 were hyperthyroiditis and 2 were chronic fatigue syndrome or fibromyalgia.

Among the reconstruction patients there was one confirmed report of a connective tissue or autoimmune disease and four unconfirmed reports.  The one confirmed report was a diagnosis of Grave's disease.

Last, I would like to present information obtained concerning patient satisfaction with their breast implants.  At each follow-up interval with their physician, patients were asked whether they were satisfied or dissatisfied with their breast implants and breast implant surgery.  The following graph presents the results from each annual follow-up visit so it includes data from the original PMA through three years, as well as the post-approval data at four and five years.  It indicates the percentage of patients who indicated that they were satisfied with their breast implants.

For the augmentation cohort, you can see that the percentage has been at 95 or 96 percent of patients at each annual follow-up visit, including five years post-implant, with 95 percent of augmentation patients indicating they were satisfied.

For reconstruction patients the percent of patients indicating they were satisfied has remained at around approximately 90 percent, with 88 percent indicating they were satisfied at three years post-implant and 89 percent indicating they were satisfied at five years post-implant.

To conclude the information pertaining to the post-approval study, Inamed is conducting its post-approval study in two phases.  The first phase has been completed and involved continuing to follow patients enrolled in the A95 and R95 studies out through five years post-implant, according to those original study protocols which involved physician examination of the patient at an office visit.

The second phase is ongoing, and Inamed is in the process of obtaining mail surveys from patients who will self-report on the status of their breast implants out through ten years post-implant.

Next, I would like to turn the presentation over to Kim Croyle, senior regulatory affairs specialist with Inamed, who will talk about the second condition of approval, which is the focus group study.

MS. CROYLE:  Thank you, Audrey.  Good afternoon, panel.

In order to meet the second PMA condition of approval, Inamed contracted with Kaplan West Qualitative Research Organization to conduct a focus group study in order to obtain women's opinions and assessment of our patient brochure.

The research objectives of the focus group study were to obtain women's feedback regarding the quality of our patient brochure, and to propose qualitative changes to improve the patient brochure, based on the study findings.

There were six focus groups consisting of 8-13 women each, three groups for augmentation, which consisted of two groups of women who had had, or who were considering, or had considered breast augmentation; one group of women who had previously had breast augmentation.  Additionally, we had three reconstruction groups, two groups of women who had considered or were considering breast reconstruction, and a third group of women who had previously had breast reconstruction.

The key findings in discussion with the patients who were participating and the women who were considering these surgeries were that the brochure was informative, and also was helpful to them and answered most of the patients' questions.  So, we had a lot of positive feedback, particularly regarding the fact that the brochure did provide them with potential risks and complications, the surgical procedures.  It proposed questions they could ask their surgeon about their surgery.  The impact the implants might have on mammography, and their other interest was knowing about the style and size options available.

An additional finding of the assessment from all these women was that the brochure was so comprehensive and extensive that it created some confusion.  There was a lot of information for them to have to, you know, review and assess.  Their key comments were difficulty in understanding and interpreting the clinical tables.  Most of the women found it king of daunting to understand the tables.  Finding, within the brochure, relevant sections was important to them.  Also, difficulty in understanding how the brochure was actually organized, where they would find the information they needed and the graphic presentation.

As a result of their comments and feedback, we have incorporated changes to the patient labeling and provided that to FDA.  Part of the changes that we have implemented are revisions to the clinical tables to make them easier to understand and clearer.  We have created separate sections for augmentation and reconstruction because the women who provided feedback on this point wanted to find the surgery that pertained to them, whether it be augmentation or reconstruction.  We also added a table of contents and a glossary of terms, and we modified the graphic presentation for ease of use.  This condition of approval the FDA has determined we have completed.

Now I will pass the microphone to Tom Powell.

MR. POWELL:  Good afternoon.  The third condition of approval involved an effort to determine mode of failure of saline implants.  The objective of this retrieval study was to use reported and observed information to understand and identify possible mode of implant failure.

In the eight months from July of 2000 to March of 2001, over 2000 saline devices were returned as deflated and were evaluated.  This quantity represents in the neighborhood of a half percent of the devices that were sold.

Evaluations included detailed visual examinations, the interpreting physicians' reports and performing appropriate testing, such as shell, material, mechanical tests.

From this effort, failure characteristics were identified and grouped into the following categories: smooth-edge opening, shell openings associated with a crease or fold, which is indicative of a true device failure as this characteristic is nearly always associated with deflation.

Sharp-edge opening, openings in the device shells where there is no associated crease.  The reason for this characteristic is yet undetermined and may be a true failure or an artifact from handling, and are associated with both reported deflated and non-deflated retrieved implants.

Valve delamination, a characteristic where 50 percent or more of the valve bond area is lost.  For devices where valve observations confirmed physician reports, this may be representative of a true device failure.  However, based on lab identification of valve delaminations on non-deflated retrieved implants, this characteristic may be the result of an artifact.

Leaky valve, this characteristic is identified by a device demonstrating leakage upon return evaluation.  The reason for this characteristic is unknown and may be most likely a result of an artifact as it is frequently associated with non-deflated retrieved implants found in the lab to have a leaky valve.

Additionally, another group of devices was identified as returned devices reported as deflated where lab evaluation could not confirm the deflation characteristic.  These devices were determined by lab evaluation to be functional and the devices in this group made up approximately 10-15 percent of returned reported deflated implants.

The next condition of approval was continued activity on fatigue testing, responding to past concerns.  These were testing of minimum thickness products; controlling the compression by load values rather than the displacement technique previously employed; and testing of the individual units rather than simultaneously testing multiple units.

Specific equipment was purchased to address these concerns and a protocol was accepted.  The smallest size was selected as worst case where the load is concentrated over the smallest area.  The test duration was accepted at either device failure or at 6.5 million cycles.  The acceptance criteria was accepted to be all samples passing the anticipated in vivo load of 5 lbs, all samples pass ingredients twice the anticipated in vivo load, and evidence that the anticipated in vivo load is past the inflection point, or elbow point, defined as the intersection of the best-fit log curve with the linear best-fit curve from test values.

Results demonstrated that all criteria were met.  All samples passed at both 5 lbs and 10 lbs, and the endurance limit or the threshold force below which an implant can undergo the run-out number of cycles without failure was determined to be 20 lbs.  The ultimate static force value was used in determining the inflection point and the in vivo load anticipated at 5 lbs was clearly underneath the inflection point of 44 lbs for smooth products and 48.5 lbs for textured implants.

The last condition of approval was to initiate a real-time study to support a five-year shelf life.  Currently, saline implants have approval for a four-year shelf life at Inamed.

To support this five-year dating, all test product was subjected to shipping simulation and testing of both packaging and product is ongoing.  For packaging performance, four test criteria is evaluated, and those are those up on the screen.  The results for the year zero are also up there.  For product performance seven categories are tested and, again, the results for year zero have all passed and testing is ongoing.

I will turn the microphone back to Ron.

MR. EHMSEN:  Thank you, Tom and Kim and Audrey also.  Just to quickly summarize, the five-year follow-up, as part of the post-approval study, has been completed and the years six through ten are in process at this point.

The focus group study has been completed.  The final report for the retrieval study is being prepared and will be submitted to FDA very shortly.

The fatigue testing has been completed, and the baseline or year zero values for shelf life testing have also been completed, and this will continue on for a period of five years.

I would like to open the floor to questions at this time.  We are joined to today by Dr. Scott Spear, who is professor and chief of the Division of Plastic Surgery at Georgetown University, who may be able to address clinical questions if you have any; also, Joanne Kune, our director f regulatory affairs at Inamed.  We also have several key members of our technical staff, Meggy Backstrand who is our senior biostatistician and contributed greatly to the preparation and organization of this data, Farhan Jahab, who is our group leader for device analysis, and Mike Taylor, our process validation group leader.  So, if you have any questions, we would be happy to try to answer them.

DR. WHALEN:  Thank you.  If I could begin, perhaps Dr. Weiss, you described in one of your slides that was labeled implant-specific complications"and I have a couple of questions.  One was labeled implant malposition and just looking at those two words, I would have thought that would have been on the next slide in terms of a surgeon complication rather than the device.  Could you elaborate on what implant malposition is?  In other words, if somebody comes out with breast augmentation at their right knee, that seems to me to be the surgeon and not your device.  How do you explain that?

DR. WEISS:  I am actually going to defer this to Dr. Spear to answer.

DR. SPEAR:  My recollection is that we wanted to be as encompassing as we could in terms of things that could be related to the implant so, for example, an implant could be put in the right position but, because of its properties or characteristics, it might displace itself.  So, although it could be surgeon related in terms of making a space where it shouldn't be made, it could also be device related.  So, I think it is just meant to be as generous as possible.

DR. WHALEN:  So, if it was in the right place when I finished is sort of like it was dry when I closed?

DR. SPEAR:  It could be.

DR. WHALEN:  The same sort of thing?  Thanks.  On that same slide, Dr. Weiss, what is implant palpability?

DR. WEISS:  Implant palpability would be that you could actually feel the device through the skin.

DR. WHALEN:  So, the ones that you have on your slide are ones where you can--

DR. WEISS:  The surgeon noted it as a complication, that they could actually feel the device through the skin.

DR. WHALEN:  So, it is moderately subjective.

DR. WEISS:  It would be based on physician assessment.

DR. WHALEN:  Correct.  Finally, Dr. Weiss, my last question for you, in the patient satisfaction data that you reported, those are percentages of patients with implants remaining in place?

DR. WEISS:  Correct.

DR. WHALEN:  Other questions?  Dr. DeMets?

DR. DEMETS:  As a follow-up of some of the follow-up studies, do I understand that no patients were excluded from your analysis?  I didn't hear you comment on that.  The patients that you started out with, they are all in the Kaplan-Meier curves for example?  There were no patients excluded?

DR. WEISS:  That is correct, no patients were excluded.

DR. DEMETS:  And how did you handle the situation for the issue of censoring that we discussed earlier today for patients who had the implant removed?

DR. WEISS:  There was no specific correction that was taken into account in the Kaplan-Meier analysis.  The follow-up compliance rate was at 80 percent or above.  So, we did not do any type of bias analysis at this point.

DR. DEMETS:  So, any complication that took place in a patient in whom the implant was removed, at least it counts up until that point in time in all the graphs?  For some outcomes it is not relevant; for some it might be.

DR. WEISS:  Correct, if the patient was lost to follow-up or was explanted of all study devices at a particular point, her data up to the time that she either dropped out of the study or had all of her devices explanted was included.  If she was replaced with another study device, the outcomes associated with the replacement continued to be followed and those were reported here.

DR. DEMETS:  While I think that your response rate is certainly pretty good, 80 percent, for reasons I said earlier it may not be good enough because of the kind of attention that this device is drawing.  But do you have any sense of the potential bias that might be in the non-responders?  You can't do a very thorough job of this, but have you looked at this at all?

DR. WEISS:  I don't have anything specifically for this study.  I know that in some of our other studies we have found that some of the reasons why patients have not returned for follow-up visits have included being unable to because they are out of the country, for example, or have been in an accident.  So, we have had some information to suggest some reasons why patients don't return.

DR. DEMETS:  If I understood you correctly, the information up to the five years was obtained through patient visits to their surgeon and/or the clinical site?

DR. WEISS:  That is correct.  It was a visit to their same physician with whom they had enrolled in the A95 or R95 studies.  Or, potentially a follow-up physician if the patient moved to another area, for example, she could see another physician in her area.

DR. DEMETS:  And, do I understand that you are proposing for the five- to ten-year to do more of a questionnaire?

DR. WEISS:  Correct.  The six- to ten-year follow-up is being obtained via a mail survey to those patients.

DR. DEMETS:  Thank you.

DR. WHALEN:  Questions for Ms. Croyle, you reported your focus groups were confused about the clinical tables.  Have you taken some actions to redesign that data presentation?  Confusion is sort of a non-descript word.  Can you elaborate a little further on that?

MS. CROYLE:  Confusion may not be the most accurate term.  A lot of women are not familiar with looking at that type of information.  It is just something they typically won't look at unless they are in academia or they have a job where they are utilizing tables.  I think it was more just an understanding, comprehension issue.

DR. WHALEN:  In that regard, I know each PMA stands on its own and, therefore, every update stands on its own so I don't want to be comparing one to the other, but there was the opinion voiced earlier that when people looked at some of the data for the other sponsor, they just couldn't believe it was that high and they just thought they were covering their own derrieres in that regard.

MS. CROYLE:  We did not have any of that kind of feedback.  That was information we did not receive.

DR. WHALEN:  So, phrased another way, was it the perception of your independent contractor doing your focus groups that women looking at your complication rates grasped what that meant?

MS. CROYLE:  They grasped what we were trying to do; it was just daunting.  Based on the review, they were allowed to read the brochure prior to the interviews, and most of the feedback from many women, not all--I actually sat in and viewed most of these focus groups in the two-way mirror, and most of them just said I don't really need this.  It is not very helpful to me.  It is not the sort of thing I would utilize.  Other women would say I would need to study it further; I am not sure what these tables are telling me.

DR. WHALEN:  Dr. Choti?

DR. CHOTI:  Ms. Croyle, regarding the focus group, you had some comparison groups.  You had some augmentation focus groups prior to and then a group of women after, as well as the reconstruction.  I am curious whether you saw any differences between the different focus groups.

MS. CROYLE:  I would say the main observation of those who had reconstruction or augmentation was that they very much appreciated that this information was available.  Many of them had had surgery quite a few years previously and they felt that if they had all this additional information at that time it would have been helpful.  So, it was mostly positive feedback from those who had had the experience.  There were a few who said, you know, if I had had this knowledge I might have made a different decision, but it was primarily that it was more informative and helpful.

DR. CHOTI:  Another question to Dr. Weiss, it is interesting that when looking at the local complications I would have expected, if anything, the number to go up, that is, the shape of the curve to be different than an asymptotic curve.  At least the reoperation curve was kind of asymptotic, higher number in the first year than the second and so forth, which perhaps may be related to exchange of implants, and so forth.  But even with deflation one might expect that to go up over time, greater in the fourth year, the third, the fourth and the fifth year than in the first year.  If I recall, your curve was quite linear.  Any speculation as to why we are seeing that?

DR. WEISS:  Not really.  I don't have any suggestions as to why that may be occurring.  I will defer to Dr. Spear again.

DR. SPEAR:  I don't claim to be a biomaterials expert, but I think we heard about this issue of the inflection point with these devices, and I think the expectation is that they will wear at a fairly linear rate up to some point, at which point the failure rate we expect will accelerate.  It is just that at five years it doesn't accelerate.  It might be at ten years or 15 years, or 25 or 30 years.  I don't know if Tom has any comment about that.

MR. POWELL:  I think it would be hard to assess a time expectancy of this implant because of the flexibility and the forces of the body counteract that to some extent and fix it in place.  So, it is a very challenging area for investigation, but I would not really have a good response to that question at this time.

DR. CHOTI:  I do want to compliment you on what I thought was very clearly presented.

DR. WHALEN:  Dr. Dubler?

DR. DUBLER:  I have two questions.  First on the focus groups, when you went back to them, you had your focus groups redesign your brochure and then brought it back.  Is that correct?

MS. CROYLE:  No, the focus group studies were conducted to revise the approved labeling for the PMA.

DR. DUBLER:  All right.  And, you said that people found the tables confusing, or whatever word we are now using.

MS. CROYLE:  Right. Difficult.

DR. DUBLER:  Difficult, fine.  Have you removed the tables and substituted--

MS. CROYLE:  No, the tables are still there because they certainly contain key information, but the language, the explanatories introducing the tables have been expanded and we have worked with the FDA quite some time on trying to make that much more comprehensive for the patients.

DR. DUBLER:  So, even if you don't read the table, you will get the information?

MS. CROYLE:  You will certainly get a brief gist of the contents of the table, yes.

DR. DUBLER:  My second question, and I am not sure to whom to address it, is that the data for your R95 have substantially greater complication rates than for your A95.  Speculation why?  Is the reconstruction happening too soon?  It is so startling in all of the data that I wonder if you have begun to think that perhaps ways of changing surgeon practice might reduce those high rates.  I am a little puzzled about why they are so high.

DR. WEISS:  I will ask Dr. Spear to address this.

DR. WHALEN:  If he can figure out how to change surgeon practice, he will get the Nobel Prize.

DR. SPEAR:  I wish we could.  Actually, it is not surprising from the clinical point of view.  It is pretty much expected.  In fact, if you want to get technical, since these are per-patient complications and many reconstructions are unilateral, the data is probably even more disparate because the per-device complication is probably even higher in reconstruction than it is in augmentation because there are two in augmentation and in reconstruction there is only one device.

But, you know, they are two operations.  There is underlying scar tissue.  There is often radiation involved.  It is a much more technically challenging situation, and it is a fluid situation.  Frankly, the standards of practice in 2002 are different than they were in 1999 or 2000 because of changing patterns of treatment of breast cancer.  So, it is to be expected that the complication rate would be higher.

What is very interesting from an academic point view is that the one complication rate which is the same is the failure rate, which is not specific to the underlying environment that is device specific.  The failure rates are actually identical.

DR. WHALEN:  Dr. Newburger?

DR. NEWBURGER:  Have you noticed any difference in the failure rates of the implants related to the positioning of the implants?  In other words, is there less of a leakage rate if it is inframuscular as opposed to supramuscular?  Is there any complication rate that you can relate to positioning?  For example, the skin and nipple complication rates are reasonable.  Are these related to periareolar positioning?

DR. WEISS:  That was not a focus of the study, to look at differences, although there was a secondary analysis that had been conducted at the time of the original PMA that looked at submuscular versus subglandular placement on certain select variables, including leakage/deflation and at the time that data, I believe, three-year rates were available and there was no difference observed at that time.

DR. WHALEN:  Ms. Brown?

MS. BROWN:  First I would like to compliment the company on getting 80 percent follow-up out to five years.  Getting patients back to the doctor's office I think is probably a pretty big challenge so I compliment you on that.

I was intrigued by the satisfaction rate of 90-95 percent in the context of reconstructive contracture rates of 35 percent, if I understood that correctly, and augmentation contracture rates of 10 percent, and leakage/deflation rates a little less than 10 percent in both those populations.  I just find that really interesting, that patients were that satisfied when they are having those kinds of rates of contracture and leakage.  I was curious as to the 5-10 percent who weren't satisfied, if you have some thoughts on why they weren't satisfied, either the contracture patients, the leakage patients.

DR. WEISS:  I don't specifically have that information with me.  We did have the physician ask if the patient was not satisfied, why she wasn't and she would provide a reason, and I don't have a synopsis of that information, but they could list a complication, for example, as a reason.

MS. BROWN:  I also thought it was a good indication that perhaps the informed consent process is working if 90-95 percent of the time patients are saying they are satisfied after five years.

DR. WHALEN:  I would like to thank the sponsor then and ask that Ms. Allen and Dr. Dawisha again come forward for the FDA presentation.

FDA Presentation

MS. ALLEN:  Good afternoon.  FDA will now summarize the status of the conditions of approval for Inamed saline-filled breast implant PMA.  For your convenience, we have provided you with a hard copy of FDA slides.

There are five conditions of approval, a post-approval study, a focus group study, a retrieval study, fatigue testing and shelf-life testing.

Dr. Sahar Dawisha will present the status of the post-approval study and I will present the status of the remaining four conditions of approval.  I will now hand it over to Dr. Dawisha.

DR. DAWISHA:  Good afternoon.  Recall that the A95 and R95 studies, which the PMA was based on, were five-year studies with three-year data presented to the panel back in March of 2000.

The sponsor has now followed this patient cohort for the total of the five years of the study, and they are going to be following the patients in an abbreviated protocol for the remainder of the five years, for ten years. The database was closed for this update in August of 2001, and I am going to be discussing augmentation, followed by reconstruction.

Table 1 shows the patient follow-up at five years for augmentation on an by-patient basis.  The percent follow-up of 81.1 percent, which is defined as the actual follow-up of 686 patients divided by the expected follow-up of 846 patients, is shown here, as well as the reasons for withdrawals.

This slide summarizes the by-patient cumulative Kaplan-Meier risk rates for selected complications with corresponding 95 confidence intervals.  The three-year data in this table is what is currently reported in the approved labeling, and the five-year data is the updated information which is going to be included in the updated labeling.

Compared to three years, the rates at five years are slightly higher, however, the confidence intervals are overlapping.  The exception to this is for the complication of implant removal where the confidence intervals are not overlapping, suggesting a significantly increased cumulative rate at five years compared to three.  I will be discussing implant removal in more detail later in the presentation.  Just to note that infection is not included here.  The five-year Kaplan-Meier rate for infection was one percent.

The number and types of additional surgical procedures performed in the augmentation patients is shown through four years, which is what is currently reported in the labeling, and through five years as an update.

Through five years, there were a total of 463 additional surgical procedures performed at 293 reoperations in 224 of the 901 augmentation patients enrolled in the study.  Of the 224 patients undergoing reoperation, the majority, 82 percent, underwent one reoperation.  Through both four and five years, implant removal for any reason with replacement was the most commonly performed additional surgical procedure, constituting approximately one-third of the procedures.  It was 30.3 percent through four years and 33.7 percent at five years.

This is followed by capsule procedures, specifically capsulotomy, which constituted the majority of the capsule procedures.  There was about three-quarters of the capsule procedures at both four and five years.

Of the 1800 augmentation implants that were enrolled in the A95 study, there were 166 implant removals, or 9.2 percent, through five years for any reason.  On a by-patient basis, there were 10.9 percent of patients who had an implant removed through five years for any reason.

The primary reason for implant removal, using the same hierarchy as in the currently approved labeling, is shown on this slide.  I have combined categories for the purpose of projecting the slide, as noted in the footnotes below the table.  Patient request constitutes approximately less than half of the primary reasons for implant removal through both four and five years.  It is 43.2 percent at four years and 42.2 percent at five years.  The majority of these patient requests for an implant size or shape change.

Of the complications, leakage or deflation constitutes the most frequent primary reason, approximately 33 percent at both four and five years.

For those patients who underwent implant removal with replacement, i.e., who had a revision and had follow-up, selected complications follow-up implant replacement are shown on this table at two years, which is what is currently reported in the labeling and at three years which is the updated information.  Although the Kaplan-Meier rates are higher at three years than at two years, the confidence intervals are overlapping, suggesting that the rates are not significantly different.

I would like to point out that for implant removal and/or replacement the overlap is minimal, suggesting that the rate at three years is approaching the limit of being significantly higher than at two years.

The sponsor provided updated information pertaining to breast disease and connective tissue disease, however, I am not going to be discussing these results.

We now move on to the reconstruction data.  The patient follow-up for the reconstruction patients is shown on this table through five years.  Again, the percent follow-up, which is the actual follow-up divided by the expected is 80 percent.

The next table shows the by-patient cumulative Kaplan-Meier rates of first occurrence through three years, which is what is currently reported in the labeling, and through five years, which is the updated information for selected complications.  Although the rates at five years are slightly higher than at three years, the confidence intervals for all these complications are overlapping, suggesting no significant differences.  Note that the reoperation rate here excludes planned procedures as part of stage reconstruction.

Table 8 summarizes the number and types of additional surgical procedures excluding planned procedures through four years, which is what is reported in the current labeling, and through five years, which is the updated information.  There were a total of 159 additional surgical procedures performed and 126 reoperations in 100 of the 237 percents enrolled in the R95 study.  Of the 100 patients undergoing reoperation, the majority, 81 percent, underwent one reoperation.  Through both four and five years, implant removal for any reason, with replacement, was the most commonly performed additional surgical procedure, constituting approximately 30 percent of the procedures through five years.  This is followed by skin or scar revision or removal, 27.7 percent, and followed by implant removal without replacement, 13.2 percent.

Of the 316 implants in the R95 study, there were 70 implants which were removed, which is 22.2 percent, through five years for any reason.  On a by-patient basis, there were 26 percent of patients who underwent implant removal through five years for any reason.  The primary reason for implant removal in this group, using the same hierarchy as in the currently approved labeling, is shown here, and through both four and five years capsular contracture constitutes the largest primary reason for implant removal, approximately 26 percent at four years and 31 percent at five years.  This is followed by patient request for a size or shape change, approximately 23 percent at four years and 21 percent at five years.  Then, followed by leakage/deflation, 16 percent through four years and 17 percent through five years.

For those patients who underwent implant removal with replacement, i.e., who had a revision and then had follow-up, selected complications are shown here.  Although the Kaplan-Meier rates at three years are higher than at two years, the confidence intervals are again overlapping, suggesting that the rates are not significantly different.

This concludes my presentation and now Ms. Allen will continue with the focus group study results.

MS. ALLEN:  The ultimate goal of the focus group study was to improve their existing patient brochure.  Inamed already described how an independent study was conducted to obtain feedback regarding their patient brochure.  They also described some of the key findings from that independent study.

FDA considered the independent study reports submitted by both Mentor and Inamed and required the same types of changes for both companies, if applicable.  The primary changes to Inamed's patient brochure were as follows:  They made significant modifications to the lead-ins, as well as to the content of the safety tables because the majority of the women found the information confusing.  They stratified the augmentation and reconstruction information.  They added a table of contents and a glossary, and they modified the graphics to read easier.

Inamed incorporated all requested changes into the patient brochure and received FDA approval.  Therefore, we consider this condition of approval fulfilled.  Inamed has just submitted a revised patient brochure and package insert that reflect the five-year post-approval data.  After FDA review and approval, Inamed with finalize them for patient and product use.

The purpose of the retrieval study is to determine modes of failure.  This information may lead to changes in manufacturing design specifications, mechanical testing requirements, and/or labeling.

In their 2001 report, Inamed submitted data on over 2400 explants collected over an eight-month period.  They provided clinical or physician observations collected at the time of explantation.  They provided laboratory observations or device failure characteristics, such as smooth and sharp crease-edge openings.  These were noted with respect to whether the device was deflated or non-deflated.  They also provided material property test data.

Inamed made numerous conclusions regarding whether the device failure characteristics were representative of a true failure or the result of an artifact.  These were summarized by Inamed earlier and provided in your panel memo.

Inamed made on hypothesis regarding a mode of failure.  That is, based on smooth-edge openings being a characteristic found more in smooth shells, failure may be caused by fold flaw and repetitive abrasion of both sides of the shell.

Inamed will submit a final report of the retrieval study in their 2002 annual report.  Therefore, FDA considers this condition of approval still open.

The purpose of the fatigue testing was to determine the fatigue strength of Inamed's product line.  These data provide additional information on the expected long-term performance of the device.

Of the five styles in their product line, Inamed performed fatigue testing on styles 68 and 168 as representative of their entire product line.  The resulting endurance load limit was 20 lbs at 6.5 million cycles run-out for both styles, which met the acceptance criteria.

As part of the test report, Inamed also supplied the ultimate static rupture results for those two styles.  The results were over 1600 lbs for both styles, which shows that the implants failed at static loads much greater than that expected during mammography, which is 55 lbs.  FDA considers this condition of approval fulfilled.

The purpose of the shelf-life testing is to support a five-year expiration date on the package label.  Inamed's shelf-life protocol involves real-time package integrity and mechanical testing performed at year zero, or baseline, and annually through five years.

In their 2001 annual report, Inamed provided an interim report with year zero data.  The results were adequate.  FDA expects Inamed to submit an updated report of shelf-life testing annually until the desired five-year expiration date is supported.  Therefore, FDA considers this condition of approval still open.

This is an overall summary of Inamed's five conditions of approval.  The post-approval study will remain open until five-year data are submitted.  The focus group study is complete.  Inamed has already revised their patient labeling to reflect the focus group study findings.  The retrieval study is currently open, however, Inamed will submit the final report in July, 2002.  The fatigue testing is complete.  The shelf-life testing will remain open until five-year data are submitted.

I will now turn it over to the panel for discussion.

DR. WHALEN:  Thank you both Ms. Allen and Dr. Dawisha.  Let me add, Dr. Dawisha, if I had had someone with your gift of lucidly turning numbers into knowledge in med school, I might not have developed a life-long passionate hatred of biostatistics.  So, thank you very much.

[Laughter]

Are there any questions of the FDA members?  Yes, Dr. DeMets?

DR. DEMETS:  Although the sponsor didn't say this directly, you certainly allude to the fact that because the confidence intervals overlap there was no significant difference, which is technically true but the other side of the story is was the study big enough to have sensitivity to find differences of that size?  So, I am not quarreling with the general gist or your comment, but I think the only caution in pushing that statement too far is, yes, they may not be statistically different but it could be that the difference is there and you just can't see it with the size of the study, or maybe there isn't a difference.  That is just a caution, and I think somewhere in all the discussion we need to have some discussion or comment on the size of the study, the precision that they are able to find for the different outcomes, whether it is failure rate or other kinds of complications.

DR. DEMETS:  Yes, I would agree.  We actually did not ask the sponsor to do statistics for that reason.  That is why I didn't present statistics but we were just sort of trying to make some sort of comparison of time trend analysis.

Panel Discussion

DR. WHALEN:  Other questions?

[No response]

Thank you very much.  We would like now to have discussion by each of the panel members, the last go around the table for this particular presentation.  I think it is only fitting that at our last meeting Dr. Chang has the last words so we will start with Dr. Miller and then work our way around the table.  Dr. Miller?

DR. MILLER:  I guess I would like, first of all, to compliment this afternoon's presentations.  I certainly feel much more satisfied with the data presented this afternoon, and reassured by it.

I think I would like to emphasize as we consider these things that this is a unique set of patients and unique problem, and looking at single things like reoperation, which ordinarily we would consider something undesirable for most surgical procedures, for most procedures like this, reconstruction procedures or esthetic procedures, it is almost expected that there will be other procedures performed to achieve the desired result.  So, looking at that as an isolated category and considering it a complication per se is not really interpreting it properly.  So, I just would like to be cautious about that.

The other thing I would like to point out is that the satisfaction levels being high, despite a 40 percent incidence of reoperations, just points out this sort of elusive side of this problem, the benefit perceived by the patient.  Ordinarily, a reoperation would lower a person's satisfaction with what they are going through perhaps, and it does at the time, but patients perceive a tremendous benefit from having these procedures, even with multiple procedures.  So, we have to remember that as we consider the risk/benefit ratio on the use of the implants.  But this kind of data is very helpful to get an idea about what the risks side of the equation is.

DR. WHALEN:  I should just state if you two are more comfortable not being in the center, since you have already completed, feel free to adjourn, but we are delighted to have you where you are as well.  Ms. Brown?

MS. BROWN:  This is my first panel meeting at which the breast implants have been discussed, so I don't have the benefit of the history of being at the last meetings.  One of the things I am very pleased to see was a requirement and was actually fulfilled is the evaluation of the patient brochures, the focus groups, because adequately informing the patients of the risks as well as the benefits of the procedure seems to be a key to ensuring future satisfaction of patients.  I guess I would take that as an indication that that is probably a reasonably successful process of the patient brochure informing the patients because the satisfaction rates have been pretty high in spite of the complication rate.

So, I would just exhort the companies to continue to keep an eye on that patient brochure process, keeping that information up to date.

DR. WHALEN:  Dr. Doyle?

DR. DOYLE:  I would like to thank the presenters this afternoon for a very clear presentation.  My only concern is that they have gotten such an excellent follow-up rate with the procedure they have been using to date, which is the physician visit, and they are now going to switch to another procedure.  I would wish, if it were at all possible, that they try to continue to go through the physician because the follow-up rate, I fear from other studies, will fall off drastically for the next five years, and it is so important to have those data in the longer-term follow-up.  If it were at all possible, I would hope that they would be able to continue to go through the physician because I think that this kind of data set--this is an amazingly good follow-up for this particular type of group of women.

I think it is interesting, as Dr. Miller noted, that the patients are satisfied even though there is a high complication rate, I think similar to some of the injectable contraceptives where there was 80-90 percent breakthrough bleeding and, yet, the satisfaction was very high.  So, it speaks again to the fact that is an individual's decision and, as long as they have the right information on which to make that decision, I feel that they have the right to make it and I would continue to urge that we give them the correct data.

DR. WHALEN:  Dr. McCauley?

DR. MCCAULEY:  I would like to basically echo some of the comments that have already been made.  I think Inamed is to be congratulated on their presentations and their clarity, and also the follow-up that they have been able to sustain at five years on their patients.

I would only suggest that at least with the focus group study, after revisions are made in those brochures, to actually take those back to those patients for follow-up just to make sure that the changes that they perceived to be resolved or clarified the brochure a little bit better are actually the same as what the patients actually think.  I think that is an important but I think overall it was an excellent presentation.  I think their data actually resolve some of the issues that were of tremendous concern prior to their presentation.

DR. WHALEN:  Dr. Dubler?  I am the only survivor of the '92 hearings and the hearings two years ago and today, and it is heartening.  I thought the PMA data that were presented two years ago were excellent and today's follow-up has been equally excellent, and I thank you for that.

For the FDA, I would suggest, and I don't know how we would do this, but these data and the patient brochure are the basis for an informed consent process but they are not the informed consent process.  That has to go on with the physicians, and I don't know if there is any way of trying to affect, change or help that process to evolve into one that uses these data and truly empowers patients.  Maybe that is beyond what the FDA can consider, but there is such a nice platform now to help women think about this decision that it would be excellent if we could take that to the next step, if the FDA could encourage that process of informed consent to go forward.  So, I thank you for the data.  They are very helpful, and for a clear presentation.

DR. WHALEN:  Dr. Choti?

DR. CHOTI:  Although many of the points have been clarified this afternoon, I still think it is important to emphasize what we have been talking about all day, and that is, this is a common operation being performed with placement of this device with increased frequency, often in healthy women, and there is still a significant paucity in data to help both the physician and the patient make informed decisions.

Although this data certainly was presented quite clearly, I think I would still encourage the increase in quality, follow-up data, preferably independently collected, longer-term follow-up.  Also, I agree with using that information as well as continued focus groups to continue to improve informed consent because it is still not clear whether this information is really being understood by the woman who is getting these implants put in.

DR. WHALEN:  Dr. Newburger?

DR. NEWBURGER:  I appreciate the seriousness with which Inamed has taken the FDA's directives, and the thoroughness with which the follow-up studies have been applied.  I hope that Inamed will apply the same thoroughness to improving product performance to have less leakage developing, less reactivity of the implants because still, when you look at the rate of leakage of 10, 11 percent over five years, this still translates, at the current rate of implantation, into between 20,000 and 30,000 women having this complication and them having to have a reoperation.  So, I hope you will apply the same thoroughness and diligence in further perfecting the product.  But thank you very much.

DR. WHALEN:  Dr. DeMets?

DR. DEMETS:  Well, I would like to echo the sentiments of my previous colleagues speaking.  I was very pleased, for example, that the follow-up was based on clinic evaluations.  I share Dr. Doyle's concern that if you switch procedures it may drop.  Obviously, there are some practical matters that have to be taken into consideration.

I have already said my comment about response rates.  I think this is a substantial improvement in the right direction.  I think because of the sensitivity and the interest of this particular study and this device, my own personal view is that an 80 percent response rate is probably not good enough to get rid of some of the concerns because with 20 percent non-response there is plenty of room for bias, and I can't predict which way the bias will go.  It could go in favor of or against the device because it is very hard to predict those kind of things.  But 20 percent is a lot of room.  If you have 80 percent now at five years, I worry about what it will be at ten years even if you didn't change procedures, and I worry even more if you do.

So, I guess my recommendation or at least suggestion to the sponsor is to work even harder to get those rates up.  I think it will serve you well in the long-run to address controversies and questions that are bound to come.

I guess this is my last meeting so I am allowed to not follow the rules totally, but if I were in the business of breast implants in any way, I think I would demand that some kind of registry--if there are 300,000 of these per year, it would be of benefit of the investigators, patients, sponsors and even the FDA to have some national registry of these kind of patients so we really get the kind of numbers and the kind of follow-up and the conditions that if you get into the registry you have to stay in the registry so we good, complete--otherwise, this controversy will be here five years from now just as sure as it is right now.

That would be a general plea to everybody.  It is an important area, obviously, and everybody cares about this, so that would be my general plea.

Finally, and a personal comment, I just want to thank the panel for allowing a biostatistician to survive four years amongst a group of surgeons, and occasionally pay attention to what I say or at least the points I am trying to make.  So, I thank the panel for tolerating me for four years, and for the FDA allowing me to be a member of this committee.

DR. WHALEN:  Dr. Chang?

DR. CHANG:  In March of 2000 I felt that McGhan Medical Corp. had done its homework in presenting data for its PMA evaluation.  I believe that Inamed has shown due diligence in the follow-up and presenting the data as part of its approval conditions.  So, I want to commend the sponsor for the data that they have presented today.

A continued plea, and I would echo what Dr. Newburger says, is let's not rest on our laurels.  I make an appeal to sponsor to continue to use its energies and resources to improve on the rate of leakage and deflation.  And, whether this includes physician education, a warning not to under-inflate to try to prevent the folding that can occur because of position usage, that may be an important part of decreasing the rate.  Regardless of this, again, there is room for improvement but I do want to congratulate the sponsor for doing an excellent presentation today, and I would like to thank FDA again for the opportunity to participate these past few years on the panel.

DR. WHALEN:  Thank you.  Dr. Witten, is the discussion satisfactory to FDA?

DR. WITTEN:  Yes, it is.

DR. WHALEN:  Thank you.  Let me add a thanks to all of those who testified, particularly this morning during the public session.  I realize many of them took their own time and some of them their own financial resources to bring a very impassioned and heartfelt message to us, and I assure you we did hear what you had to say to us and have taken that into consideration.

I would like to thank all of my fellow panel members, from very solid new additions to the panel to always stalwart temporary panel members, and give special thanks to Drs. Chang and DeMets for their service elbow to elbow with me for the past few years.

Finally, let me give my profound thanks to the FDA.  Between 12 years active and 16 years reserve service in the United States Navy, I have somehow developed a moderately cynical attitude towards the government--hard to believe.  Despite that, the FDA has consistently, at every level--scientific review, leadership, professional, administrative and support--always done an impeccable job, and they have done it in a fishbowl.  It has to be this way; it should be this way, but every watchdog group and every congressional member of the Hill is watching you guys every time you use a pencil eraser.  So, my hat is off to you.

Having completed our business, this meeting of the General and Plastic Surgery Devices Panel is adjourned.

[Whereupon, at 2:55 p.m., the panel was adjourned]

- - -