MET SERVICES

File STN/________

SYNOPTIC OR CLIMATE STATION INSPECTION REPORT

Name of Station:       Keetmanshoop	                  Station No:   68312

Observer’s name:  J. Jacobs		  	   Person interviewed: J. Jacobs

Site and enclosure (any changes in site, exposure, layout of instruments, general condition or alterations since last visit):

First visit, the instruments are well maintained and there are no obstructions.

RAINGAUGES

MANUAL GAUGE

GAUGE A                                                                                                    GAUGE B

| Type | 5 inch copper rain gauge |
| --- | --- |
| General condition | good |
| State of Rim | Slight dent on the rim |
| Is Rim Level? | Yes |

| Type | nil |
| --- | --- |
| General condition |  |
| State of Rim |  |
| Is Rim Level? |  |

Diameter Test                                                                                        Diameter Test

| GAUGE | REPLACEMENT |
| --- | --- |
| 1. | 1. |
| 2. | 2. |
| Mean: | Mean: |

| GAUGE | REPLACEMENT |
| --- | --- |
| 1. 127.04mm | 1. |
| 2. 126.24mm | 2. |
| Mean: 126.64cm | Mean: |
| Inner can tested for leaks? | Leaking | Not Leaking |
| Outer can tested for leaks? |  |  |
| Funnel tested for leaks? |  |  |

| Inner can tested for leaks? Yes | Leaking | Not Leaking x |
| --- | --- | --- |
| Outer can tested for leaks? Yes |  | x |
| Funnel tested for leaks? Yes |  | x |

(For standard 5” gauge tolerances allowed are: mean of any two diameters +/- 0.010” or +/-0.3mm. For any single diameter +/- 0.020” or +/- 0.5mm)

Gauge Replaced:		No		Replacement Type: _no_____________________________

Measuring Glass Condition:  20mm broken		Replaced: 	no

Maintenance/Repairs

Leveling of the rain gauge was done.

## AUTOMATIC RAINGAUGE

| Type | Casella |
| --- | --- |
| Serial No | W5364 |
| General condition | Is in good shape but the funnel is broken |
| State of Rim | Good |
| Is Rim Level? | Yes |
| Height of rim above ground | 123cm |

Is gauge calibrated correctly? (Provide details of measured quantities and attach copy chart used in calibration)

Calibration is done by siphoning.

Maintenance/Repairs carried out:

None.

SCREENS

| Type | standard |
| --- | --- |
| Condition of Screen | Old and the door is broken |
| Condition of Stand | good |

Repairs carried out and action taken:

None.

THERMOMETERS

| Thermometer | Serial No. /Make | Check Therm | Station Therm | Check Therm | Station Therm | Check Therm | Station Therm | Mean difference | Serial no. of replacement |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Inspector’s Therm | GTH 175 PT | 30 |  | 29.6 |  | 29 |  |  |  |
| Maximum | C.W. Price & Co 2238.05 |  | 30.2 |  | 29.2 |  | 29.2 | +-0.67 | No replacement |
| Minimum | C.W. Price & Co 894/91 |  | 30.8 |  | 30.1 |  | 29.7 | +-0.00 | No replacement |
| Dry Bulb | Lambrecht 5325 |  | 30.1 |  | 29.6 |  | 29 | -0.03 |  |
| Wet Bulb | Therm Scherder |  | 29.9 |  | 29.4 |  | 28.8 | +0.16 | No replacement |
| Grass Minimum | None |  |  |  |  |  |  |  |  |

If the inspector’s thermometer is a Digitron 3204SR & PRT enter pre-visit and post visit checks below:        Mean Diff.

| Inspector’s thermometer |  | Date 29/09/06 | Dtron | Insp. | Dtron 22.9 | Insp. 23.1 | Dtron 22.95 | Insp. 23.3 | +0.28 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Inspector’s thermometer |  | Date 16/10/06 | Dtron 24.1 | Insp. 24.1 | Dtron 24.0 | Insp. 24.0 | Dtron 23.9 | Insp. 23.9 | +-0.0 |
| Inspector’s thermometer calibration due date | 27/02/07 |  |  |  |

Condition of Thermometers, wick and reservoir:

The thermometers are in good condition except the maximum which showed high temperature difference as compared to the traveling standard. All this thermometers have never been calibrated since installed

Condition of Earth Thermometers, Tubes and Plot:

They is only one thermometer (120cm) but is broken.

Action taken and work carried out:

None.

BAROMETERS

| TYPE | SERIAL NUMBER | CONDITION |
| --- | --- | --- |
| PA21 & the Fortin Barometer | Mercury 26550 | good |

|  | Inspector Barometer | Station Barometer | Wind Direction | Speed in Knots |
| --- | --- | --- | --- | --- |
| First Check | 899.52 | 899.9 |  |  |
| Second Check | 899.53 | 899.8 |  |  |
| Third Check | 899.55 | 899.9 |  |  |

Comparisons made with inspector’s Barometer?          Yes but they is no wind system.

Correction cards held on Station?                                  Yes,

Are Correction Cards correct for type of observation prepared by Station? Yes

Action taken:

No action was taken.

NON AUTOGRAPHIC INSTRUMENTS

### SUNSHINE RECORDER

Type:	Campbell Stokes			Serial No. not available

Condition of Recorder and Mast and any action taken:

The available mast is that of a Pressure Plate. Its height is approximately 6m. This height is not suitable for wind observation especially for synoptic purposes.

### CUP COUNTER ANEMOMETER

Type: Cassella			Serial No.    Not available.

Condition of Anemometer and Mast and any action (note results of spin test):

The instrument is in good condition.

### WIND VANE

Type:	       				  Orientation:

Action taken:

Not available.

### EVAPORIMETER

Type: Class A Evaporation Pan				Height above ground level: 19cm

Dipstick Number and Condition of instrument: A graduated ruler is being used no number has been assigned to it.

Action Taken, give details of measurements of any two random 0.1mm graduations on dipstick,(use calipers):

CLOUD SEARCHLIGHT:

Type:					Maker &

No.

Location of instrument:

Condition of instrument and any action taken:

Not available.

Additional instruments:

No additional instruments are available.

AUTOGRAPHIC INSTRUMENTS

THERMOGRAPH :

Type:							Serial Number:

Condition of instrument and maintenance carried out:  This instrument is not available.

HYGROGRAPH

Type:							Serial Number:

Condition of instrument and maintenance carried out: This instrument is not available.

________________________________________________________________________________________________

### THERMOHYGROGRAPH

Type:	Daily			 Serial Number: 543077

Condition of instrument and maintenance carried out:

The instrument is working but the pen holder is broken.

### BAROGRAPH

Type:	Daily						 Serial Number: 3.0811.10.000

Condition of instrument and maintenance carried out:  The instrument is in good working condition and is well maintained.

#### ANEMOGRAPH

Type:	Daily type						 Serial Number:

Condition of instrument and maintenance carried out: Instrument in good working condition

### SOLAR RADIATION

Sensor Type:	 not available	                                                   Serial Number: 	_______________________

Sensor Type:	__________________________________	   Serial Number:	_______________________

Sensor Type:  	____________________________________           Serial Number:	_______________________

Recorder Type:	____________________________________           Serial Number:	_______________________

Condition of instrument and maintenance carried out:	_______________________________________________

_______________________________________________________________________________________________

Additional Instruments:

No additional instruments.

VISUAL OBSERVATIONS AND GENERAL

Wind direction (check correct use of the Beaufort Scale, bearings known?)

Yes the bearings are known and books are available.

Visibility (marks known, table on station?)

Hills about 20km south of the station are used as visibility reference point.

Cloud amounts, heights and types: The table is available

(Heights of surrounding hills known, is there a cloud book and cloud chart on station?)

Yes there is a cloud chart and heights of nearby hills are known.

Field books 301s and remarks:

Available and are adequately stalked and observations are well recorded.

Synoptic/Metar Coding(Practical test, comments, action taken. Is more training required?)

Officers are conversant with codes. Training in climate reports preparation. Preventive maintenance and calibration of instrument short course are need.

### Supplies, are there adequate supplies on station?

1 year supply is always ordered in advance.

Instrument care, management and maintenance:

Maintenance of instruments is done once in while.

Sketch, photograph and horizon profile of site:

See attachment.

Are hydrogen and balloon filling procedures being correctly followed

No upper air is being done at the present time however; preparations are being made to start upper air observations next year, probably the first quarter of the year.

Other items and general remarks on impression and ability of observer(s):

The station is well maintained, however there are very few instruments. The observers have the necessary training but they need some training to enable them to fully understand the importance of their work.

RECOMMENDATIONS                                                                               ACTION COMPLETED

| Soil Thermometers |  |
| --- | --- |
| Stevenson Screen, |  |
| Automatic rain gauge |  |
| Complete Wind system |  |
| Refresher Training of observers |  |

Inspected by: D. F. MOLOTSI and E. MAGAGWA			 Date: 12/10/2006