# Estimation of Uncertainties in the DCOPS Primary Reference to CSC Average Active Area Reference Point

# Digital CCD Optical Position Sensor (DCOPS) Sensors

DCOPS sensors are viewed as a collection of four independent vectors, corresponding to the four CCD pixel arrays, referenced to a common calibration slot on the DCOPS hardware assembly (where the CSC calibration pin is to be inserted).  Systematic uncertainties in the location of the origin of these vectors, i.e. the first active pixel position, is generally independent of the uncertainties associated with the pixel array orientation.  Uncertainties in the location of any given pixel in a CCD to the calibration slot can be separated into uncertainties directly associated with the calibration of the first pixel position and those associated with the pixel array orientation and length.

1.  Uncertainties in Direct Calibration of First Pixel Position.

Calibration of first active pixel positions to the reference pins for large numbers of sensors is done on a specially designed test bench at Northeastern University (NEU).  The calibration bench is constructed of a laser diode assembly, a single directional DCOPS sensor which has been previously calibrated under a microscope, and a mount for the DCOPS sensor which is to be calibrated.  In this arrangement, the calibrated DCOPS sensor acts as a mask to calibrate the uncalibrated DCOPS sensor.  NEU indicates typical uncertainties in the first pixel location using this calibration technique will not exceed 40 (m (COPS Sensor Board Calibration, J. Moromisato et al, Oct 2000, unpublished).

2.  Uncertainties in Pixel Array Orientation and Length

For simulation and reconstruction purposes, it is assumed that the four CCD vectors lie in the local DCOPS XY reference plane and run parallel to the X or Y axis.  The error introduced by this assumption manifests itself as a correction to the Sony specified pixel-to-displacement conversion of 14(m per pixel along the array.  The uncertainty in the determination of the position of the charge distribution will scale across the pixel array as the cosine of the angle by which the CCD array deviates from its optimal orientation.

The uncertainty introduced by the error in pixel array length and the misorientation of the pixel arrays can be estimated as the quadrature of the maximal error due to the misalignment of the pixel arrays within the CCD packages, the uncertainty in the length of the pixel array, and the maximal error due to misalignment of the CCD package within the DCOPS window frame.

2.1  Error Inherent in Manufacture of CCD Arrays and Packaging

Direct measurement of the pixel array length was performed on a representative sample of ILX-551 CCDs.  In addition to the 2048 active pixels in each pixel array, one end of the array contains an additional 33 dummy pixels while the other end contains an additional 6 dummy pixels.  The total length of the pixel strip is calculated to be 29.218 mm based on the Sony pixel specification of 14 x 14 µm2 with no manufacturer quote on the uncertainty in pixel size.

Sony has provided estimates of uncertainties for the placement of the first pixel in the array in the CCD package, but does not offer any estimate as to the uncertainty in the pixel array length or orientation.  These uncertainties where estimated as the maximal deviation from Sony specifications found in the small set of CCDs studied.  Systematic errors embedded in the ILX551A CCD packages have been studied on a small sample of unmounted CCDs.  Direct measurements of the pixel array lengths on these samples revealed a maximum discrepancy of (50 ( 10) µm with the Sony specification (COPS Sensor Board Calibration, J. Moromisato et al, Oct 2000).  All measurements of the array lengths yielded a result which was always greater than the Sony specified value.

Uncertainty occurring in the final active pixel position as the result of a misalignment of the pixel array within the CCD package has been determined to be less than 4 (m based on a measured 15 mrad deviation of the array with the package edge (maximum misorientation found in the small sample of measured CCDs) (COPS Sensor Board Calibration, J. Moromisato et al, Oct 2000).

Figure E1 :  SONY ILX-551 CCD Specification and Direct Measurement (mm).  Diagram A shows the dimensional specifications and tolerances for the ILX-551.  Diagram B shows the dimensions of the CCD taken from the small sample of studied ILX-551s which exhibits the greatest deviation of the pixel array (in red) placement from the optimal location.

2.2  Uncertainty of CCD Package Orientation Inside DCOPS Window Frame

Uncertainty occurring in the final active pixel position as the result of the misalignment of the CCD package, and thus the encased pixel array, scales as the cosine of misalignment.  The present DCOPS window frame design incorporates a specially designed polycarbonate mount for the CCDs.  Tolerances for the positioning of each end of the 53.71 mm long CCD mounts in the window frame is (100 µm, contributing an error of less than 1 µm in the determination of the distance along the pixel array.  Tolerances for the placement (orientation) of the CCDs within the polycarbonate mounts are estimated at >?? mrad [waiting for Rich’s Estimate], contributing a maximal uncertainty of less than 10 µm [waiting for Rich’s Estimate] in the determination of the last active pixel in the array.

3.  Final Estimation of Uncertainty in CCD Pixel - DCOPS Reference Pin Calibration

The final estimation of the uncertainty associated with the determination of the location of any given pixel in the DCOPS sensor relative to the CSC reference slot is taken to be the quadrature of all known errors in the determination of the first and last active pixel positions in the pixel array.  In addition to the systematic errors associated with the placement of CCD arrays relative to the primary reference pin, the manner and stability in which the centroid of charge distributions is determined must be considered.  It has been determined that centroids from successive measurements on the CCD are repeatable to (14 µm.  This uncertainty is characterized as the uncertainty associated with a particular measurement and considered separately in the COCOA simulation.  The contributions and final estimate of the systematic uncertainty associated with the placement of pixels is given in Table E1.

| Uncertainty Origin | Magnitude (µm) |
| --- | --- |
| Direct Calibration of First Pixel (NEU Estimate) | 40 |
| Array Misalignment (inside package) | 4 |
| Array Length | 50 |
| Package Misalignment | 10 |

Final Estimation of Uncertainty :					65µm

Table E1 :  Contributions to Final Error in Determination of Pixel-to-CSC Reference Pin Calibration.   Uncertainties due to misalignment and pixel array length are estimated as the worst case errors in the determination of the final active pixel position.

# CSC Active Center - DCOPS Reference Pin Calibration

As previously discussed, Cathode Strip Chambers are described in terms of their ‘active center’, which corresponds to the average of the individual panel active areas projected onto the plane of the first panel (See Section XX for more detail on chamber definition).  Determining the certainty with which this active center can be externally referenced is crucial to the successful simulation and reconstruction of the endcap muon system.  Uncertainties in the determination of the relationship between CSC active cathode strips and DCOPS reference pins can be separated into uncertainties associated with individual panel definitions and manufacture, assembly of multiple panels to form the complete CSC chamber, the mounting assemblies which affix the DCOPS sensors to the surface of the assembled chamber, and deformations in the chamber after installation in CMS.

All estimations of uncertainties addressed here regarding CSC chamber and DCOPS mounting hardware tolerances are typically gathered from the specified tolerances placed on the fabrication of components.  In most cases, particularly CSC panel definitions, adherence to these tolerances has been confirmed by direct measurement on an appropriate number of preproduction samples.  However, there are many components which have not yet been produced on a large scale.  Uncertainties for these components have been estimated from production drawings.  A far more meaningful estimation of uncertainties for such components should be taken from the rms value of deviations found from a sufficient sample of the finished products.

In keeping with the convention established earlier (Section XX), the local chamber coordinate system is taken as right handed with the local Z axis running across the chamber centerline from the narrow end of the chamber to the wider end and the local Y axis running from the bottom layer to the top layer of cathode strips.

1.  CSC Panel Definition

All cathode strip chambers are constructed of a polycarbonate honeycomb panels with 1.5 mm G-10 epoxy fiberglass skins coated with a 34 (m layer of copper.  Individual panels are first drilled with two CSC Alignment Holes (where the CSC Alignment Pin will ultimately be inserted) along the centerline (See Figure E2).  These two holes ((25 (m tolerance on the diameters) establish the reference system from which all other machining on the panel is established.  A high precision router is then used to mill the cathode strips and associated artwork directly into the copper surface of the panel.  Accuracy of the router has been confirmed by direct measurement.  Errors in absolute strip position exhibit accumulative systematics of <15 (m over chamber width and have ((local rms spread) of 30 (m.  The straightness strips have been measured as having a ((rms) of 22 (m

Figure E2 :  Placement of Pins, Holes, and Etchings Required on CSC Chamber.  Diagram A shows the dimensional specifications and tolerances for the ILX-551.  Diagram B shows the dimensions of the C

Additionally, a series of ten slot cuts are placed on each panel for future calibration and quality control as well.  The slot cuts are referenced directly to the center line defined by the two CSC Alignment Holes as well as a particular cathode strip number -- but not necessarily the first strip nor the same strip in all chambers.  Each mark is referenced within (25 (m to the center point of the CSC Alignment Holes and within (25 (m to the particular strip being referenced (the “Reference Strip”).

A series of secondary dowel pins and slots used to mount the DCOPS sensors to the chamber surface are also drilled at this time.  On the wider end of the chamber, these holes position the precision pins and bolts used by the DCOPS mounting brackets to reference the CCDs directly to the cathode strips.  On the narrow end, a series of holes for alignment pins are established which locate the plates for DCOPS mounting hardware to be glued to the panel surface.  The location of all holes are placed with a certainty of (25 (m on the surface of each panel and have their diameters drilled to a certainty of (25 (m.

2.  Assembly of Final Chamber from Multiple Panels

Once all six panels are milled and prepared, they are layered to form the composite six layer CSC chamber.   The precision with which each panel in the chamber is positioned with respect to neighboring panels can be estimated from the uncertainty with which the central CSC Alignment Holes are placed ((25 (m) and the tolerance placed on the diameter of the hole itself.  Since the placement of all six layers are independent of each other, the maximal misalignment of any two given planes (two sets of holes, each with  25(m uncertainty in the diameter and relative location of  their center points) is estimated to be (50 (m.  The uncertainty in the average center point across all six chamber planes is (87 (m.

After the panels are stacked together, a frame is mounted to the edges of the chamber to add rigidity to the structure, protect connections, and form a Faraday shield around the electronics.  Though the DCOPS mounting tower assemblies will mount on top of the frame structure, placement of the tower mounting hardware is completely determined by an alignment pin that sits in the previously drilled alignment holes.  Since the CMS triggering requirements will accommodate relatively large uncertainties in the final determination of the chambers along the CMS Z coordinate, the placement of elements along the local Y axis of the chamber (CMS Z), including the frame, is not tightly controlled during manufacturing.  The certainty with which the placement of the frame above the first strip plane will be known is estimated to be (127 (m.  The frame itself is estimated as being flat to (25 (m  whereas chamber flatness is estimated as no worse than (200 (m across any 60 cm span.  Tolerances on the panel thickness are asymmetric, +508 (m to -254 (m, and determined by the manner in which the panels are manufactured and cut.

3.  Alignment Mounting Hardware

Mounting hardware to attach the DCOPS sensors to the chamber frames is similar across all chamber varieties.  The mounting hardware consists of a series of spacers, shims, and mounting plates which secure the primary DCOPS mounting bar (where CCD-dowel calibration is done).  CSC chambers in both inner and outer rings of ME (2, (3, and (4 disks are placed into two layers.  The chamber configuration for Trigger towers makes it necessary to mount hardware on the back layer for one outer CSC.  Chambers on back layers will have mounting hardware called towers to give the attached DCOPS sensor an elevation in the SLM line which matches the front layer.

As previously noted, sensors affixed to the larger end of the chambers (+Z local end) are mounted directly to the chamber frame while sensors on the smaller end of the chamber (-Z local end) are attached to a special plate which is glued into position with the aid of a special pin and slot set during the fabrication of the chamber.  The precision with which the inner plate can be glued to the surface of the chamber is determined from the tolerance with which the pin and slot are placed on the chamber surface (( 25 (m each) as well as the precision with which the receiving pin and slot are placed on the plate to be glued (( 25 (m each).  Therefore, it is estimated that the DCOPS plates on the narrow end of the chamber can be glued in position within (50 (m of the specified location - approximately the same tolerance with which they can be attached on the +Z local end to the chamber frame.

Although the heights of the towers on the back layers differ by 290mm from the standard brackets, the plates and assemblies used are essentially the same.  The mounting brackets are attached across two standoffs, each with an asymmetric tolerance of +.100 -0.000 .  The standoffs locate a set of dowel pins on a stock Aluminum jig plate (( 125 (m thickness).  The DCOPS mounting bar connects directly to this plate.  Holes and pins used to secure the DCOPS are positioned and drilled to  ( 25 (m, contributing an additional (50(m uncertainty in the referencing of the sensors to the strip positions.

The potential uncertainty in the orientation of the DCOPS sensors about the primary DCOPS dowel pin due to the uncertainty of the standoffs (separation of 88mm) is less than (1(rad.  Likewise, the uncertainties of the chamber frame positioning above the first panel on the chamber contribute an angular uncertainty about the same axis of less than (1(rad as well.  Rotational uncertainties about other axes are found to be only slightly larger, but can be safely neglected by COCOA as each CCD is one dimensional.

4.  Deformation

One final consideration in defining the precision with which individual layers can be externally referenced is the deformation of the chambers when mounted in the detector.  Since the CSC chambers are mounted vertically, the weight of the chambers creates a shearing effect across the six layers of the chamber.  This effect is will be most prevalent when the shearing occurs perpendicular to the direction of the cathode strips.  Thus, the effect is largest for chambers lying along the CMS X axis.  The displacement of the individual layers due to the shearing force increases from the bottom layer (closest to the RF iron mounts) to the top layer.  The maximal displacement of the uppermost layer has been directly measured as less than 25 (m.

5.  Angular/Rotational Uncertiainties

Components in COCOA are specified by their location and orientation making it necessary to estimate the rotational uncertaities with which the DCOPS sensors are placed on the chamber.  Since the CCDs used in the DCOPS sensors are one demensional pixel arrays, rotational uncertainties about the axis normal to the plane of the CCDs (DCOPS local Z axis) are the most important.  Estimations of uncertainties about this axis, where it might be expected the slope and flatness of surfaces will contribute most, can also be estimated from the tolerance and separation on the mounting bracket standoffs, CSC panel/frame uncertainty (127(m), length (>>100cm), and flatness ((200 (m across 60 cm), and other tolerances placed on the components in the mounting bracket (Al plate, DCOPS window mounting bar - both (25-50(m across (90mm). The uncertainties about the DCOPS Z axis are summarized in Table E4.

Uncertainties about the DCOPS X and Y axis are determined from the uncertainties of the two precision pins or holes used to position components on the chamber and mounting brackets and the separation between them.  In all cases, the uncertainties associated with the placement of pins and holes ((25-50(m) and the relatively large span of between them ((90mm) mean individual contributions to the rotational uncertainty in the orientation of the DCOPS sensors is on the order of 1 (rad.  The uncertainties about the DCOPS X and Y axes are summarized in Table E5.

6.  Final Estimation of DCOPS - CSC Active Center Uncertainties

The final uncertainty associated with the DCOPS reference pin - CSC active centers are determined from the quadrature of all estimated uncertainties in the plane of the strips (Table E2) and perpendicular to the plane of the strips (Table E3). The estimated rotational uncertainties about the DCOPS X, Y, and Z axes are summarized in Tables E4 (DCOPS Z axis) and E5 (DCOPS X and Y axes).  The rotational uncertainties have been estimated from the uncertainty and separation of the pins and/or holes which join components.

| Uncertainty Origin | Magnitude (µm) |
| --- | --- |
| Central Alignment Pin - Notched Alignment Marks | 25 |
| Notched Alignment Mark - Numbered Reference Strip | 25 |
| Intrinsic Strip Positioning (from milling) | 30 |
| Averaged Centerline Across 6 Assembled Planes | 87 |
| Positioning of Primary DCOPS Alignment Pins/Holes | 25 |
| Diameter of Primary DCOPS Alignment Pins/Holes | 25 |
| Placement of Mounting Plate On Chamber | 50 |
| Placement of DCOPS Mounting Plate | 50 |
| DCOPS Calibration, Construction (Table E1 Result) | 65 |
| Maximal Shearing Effect | 25 |

Final Estimation of Uncertainty Along X Axis of Chamber:		144 µm

Table E2 :  Estimation of Error Transverse to CSC Chamber Centerline.  This table shows the uncertainties associated with the determination of the displacement between the DCOPS alignment pin and the cathode strips transverse to the chamber centerline (local chamber X axis).

| Uncertainty Origin | Magnitude (µm) |
| --- | --- |
| Panel Thickness (Maximal deviation) | 508 |
| Frame to Panel Placement | 127 |
| Mounting Bracket Chamber-Shim Standoff | 100 |
| Mounting Bracket Al. Plate | 125 |

Final Estimation of Uncertainty in Y Plane of Chamber:			538 µm

Table E3 :  Estimation of Error of DCOPS Positioning Above First Strip Layer (local Y axis).  This table shows the uncertainties associated with the determination of the displacement between the DCOPS mounting plate and the first plane of cathode strips.  Tolerances for CSC construction are not as tight along this direction since the uncertainties will lie along the CMS Z axis.

| Uncertainty Origin | Magnitude (µrad) |
| --- | --- |
| Upper Cathode Panel - Frame Relationship | 1 |
| Frame (or Fwd Glue Plate) - DCOPS Mounting Bracket Base | 1 |
| DCOPS Mounting Bracket Base - Mounting Bracket Shim Plate | 1 |
| Mounting Bracket Shim Plate - DCOPS Mount Bar | 1 |
| Straightness of DCOPS Mount Bar - DCOPS Window Frame | 1 |

Final Estimation of Uncertainty in DCOPS Orientation on Chamber:	2.2 µrad

Table E4 :  Estimation of Error of DCOPS Orientation About DCOPS CCD Plane Normal (local DCOPS Z axis).  This table shows the uncertainties associated with the determination of the orientation between the DCOPS ccds and the first plane of cathode strips.  Most uncertainties were less than 1 µrad and have been rounded up.

| Uncertainty Origin | Magnitude (µrad) |
| --- | --- |
| Upper Cathode Panel Pins/Holes | 1 |
| DCOPS Mounting Bracket Base | 1 |
| Mounting Bracket Shim Plate | 1 |
| DCOPS Mount Bar | 1 |

Final Estimation of Uncertainty in DCOPS Orientation on Chamber:	2 µrad

Table E5 :  Estimation of Error of DCOPS Orientation of DCOPS CCD Plane (local DCOPS X/Y axis).  This table shows the uncertainties associated with the determination of the orientation between the DCOPS ccds and the first plane of cathode strips.  Most uncertainties were less than 1 µrad and have been rounded up. Rotation uncertainties and misorientation are not as important about these axes, as the CCDs are one dimensional.

References :

CMS Technical Proposal, CERN / LHCC 94-38, LHCC/P1, 15 December 1994

CMS Muon System Proposal, CERN / LHCC 97-32, CMS TDR 3, 15 December 1997

SONY ILX551A 2048-pixel CCD Linear Sensor (B/W) Product Specification, Sony Electronics Inc., Document E00439-PS

(http://www.sel.sony.com/semi/PDF/ILX551A.pdf)

COPS Sensor Board Calibration, J. Moromisato et al, unpublished, Oct 2000

(http://www.dac.neu.edu/physics/j.moromisato/emu/alignment/cops_calib.pdf)

Conversation with Oleg Prokofiev, 7 July 2001

Conversation with Nelson Chester (lead engineer), 5 July 2001