Sample collection and data analysis for Single Crystal X-ray diffraction studies at Beamline 11.3.1

Allen Oliver and Sirine Fakra

Last updated, March 2005

These instructions assume some familiarity with goniometers and diffractometers in general.  Many of the instructions assume that you are aware of how to operate a laboratory instrument, select and mount crystals and how to center them.  It is a guide to operating the D8 goniometer with the APEX2 software package.

- Examine your sample under the polarized microscope and mount it on a magnetic head.    Use the Crystal Wand to support the magnetic head while your are mounting your sample.

- To mount and center your sample, connect to the server computer (black background with the word: “Server” displayed in the middle).  To switch between the two Bruker PCs (server and client), double tap the “Scroll Lock” key on the keyboard.

- Make sure that BIS (Bruker Instrument Services) is running on the server before doing anything.  It should be running at all times.  If it is not, ask for assistance.  Launch the APEX2-Server program (double click on the icon on the desktop).

- After the program launches select Instrument ( Connection.  Type in “localhost” in the text box and click on “Connect”.  You will have to wait while the computer communicates with the BIS software to set up the instrument.  This usually takes 15 – 30 seconds. You will be connected when the connection window disappears.

- Click on “Simple Scans” from the side bar. Wait for the window to appear in the main panel of the APEX2 window.

- To center your crystal drive the goniometer to omega = -180 and phi = 0 deg (enter the angle values in the appropriate text boxes on the right hand side of the window and click on the ”Drive” button just below them).  To rotate your crystal 180 deg. Click on the “Phi + 90” button twice then on the drive button.  Center your crystal as well as you can in this orientation, then rotate the crystal to phi = -90 or 90 deg by a single click of the “Phi +90” button.  Center your crystal as well as you can in this orientation.  Reiterate as necessary until you are satisfied that your sample is located along the phi-axis.  Drive phi to 0 (zero) before continuing with the final height adjustment.  With phi at zero, drive omega to -30 deg. And adjust as necessary. DO NOT GO BEYOND OMEGA= -30 AS YOU WILL HAVE COLLISION.

- Go to the Client computer to start your matrix and data collection (double tap the “Scroll Lock” key).

- Double click on the Start Database icon on the desktop.  The database is necessary for APEX-2 to function.  It keeps a record of the projects that have been run.

- Launch APEX2 from the desktop and login as “guest”, the password is the same.

- Connect to the server computer (Instrument ( connection).  The IP address of the server appears in the pop-up window, click on “Connect” and wait. Once the connection window disappears you are connected to BIS and are able to begin data collection.

- Select the “Setup” option (should be active) and click on the “Describe” module and enter the data as appropriate [usually chemical formula, but crystal information (color, habit dimensions) can also be entered].  After you have entered the data click on the lower of the two “X”'s to close this module (DO NOT click on the upper-right-most “X” this will close the client and you will have to restart the whole process).  It is quite important for the current version of APEX-2 to reduce the number of open windows, the scripting language sometimes gets confused.  In addition, you will need to be patient as there are several portions within the program where there is a short wait for information to be loaded and transferred between the sever and client computers.

- Select the “Collect” option and select the “Experiment” module.  You will be presented with a window that has a spread-sheet style layout describing how to collect the data.  The field codes describe the various goniometer positions, collection times, angles that are scanned and so forth.  The fields are similar to those found in SMART.  This module is “tabbed”.  The currently active tab describes the data collection routine (similar to “Edit Hemisphere” in SMART).  The second tab “Monitor Experiment” shows the current frame as the data are collected.

- Uncheck “Correlate Frame” and check “Generate New Dark Images”.  Next set the default collection time (typically a 1 second frame is appropriate for screening samples).  Set the “Frame-Size” to 512 x 512 (512 mode) or 1024 x 1024 (1k mode).  Normally we run in 512 mode, but you may choose to run in 1k mode.  Click on the “Append Matrix Strategy” button (lower-left corner of the main window) and wait while the data is loaded.  Please be patient, it takes a short time for the software to compile and insert the data in the appropriate fields. When the runs have loaded there will be six lines of sample runs.  The first four describe a “standard” matrix. NOTE: After you’ve run your matrix and analyzed it , change the filename from “matrix” to “Sample ID” (whatever you have chosen) before you begin collecting your full dataset. DO NOT OVERWRITE FILE NAMES AS THE APEX2 SOFTWARE WILL HANG UP. To prepare a matrix run, select the number “5” on the left-hand-most column and, while holding the “SHIFT” key down, click on the number “8” below it.  Right-click on the “Active” field title and select “No”.  This will disable the last four blocks (which are the standard set-up for a full dataset; we don't need to run one of those right now).  When you are ready, click on the “Execute/Resume” button on the lower-right-hand side of the window.  This will start the data collection.  There will be a pause while the server and client set-up the scans and data collection will commence.  The pause also includes automatic collection of new dark frames.  These will be collected at the start of each black of data.  The window will automatically switch to the “Monitor Experiment” tab, and your data will be displayed as it comes off the detector.

- When your matrix runs have completed, select the “Evaluate Crystal” option and launch the “Determine Unit Cell” module.  Wait briefly for the computer to load the data and you see on-screen the first frame from your matrix run.  The center of the screen displays the frame image.  The right-hand side has the various functions necessary to “Harvest” (Threshold or Redetermine Unit Cell in SMART lingo) spots from the frames.  The upper right hand corner displays the name of the first frame from the selection that you wish to use.  The frame name format is: filename_XX_YYYY.sfrm  XX = run/block number, YYYY = frame number.  Hence you should see filename_01_0001.sfrm in the top corner.  Below this is a number of frames that you wish to use for the matrix (it defaults to 20, our standard matrix has 30).  Select “Harvest Spots” and you should see some green circles appear around reflections on the frame.  Adjust the sliders to increase or reduce the number of reflections to be included (I/sigma(I) threshold).  You are also able to select a frame to view with the right-hand slider.  Typically 7 marks on the slider appears to give sufficient data to obtain a unit cell.  Click on “OK” to store the reflections.  Change the block number to the next in the sequence and repeat until done.

- Select “Index” to begin indexing reflections.  It is advisable to check the “Spots Must Span Images” box and set the resolution limit to something reasonable depending on the observed diffraction pattern f your sample (0.8 A is good for many samples).  Adjust sliders to your hearts content and click on the “Index” button in the pop-up window.  This will then attempt to index the reflections and give you a reduced unit cell.  If the routine is unable to determine a unit cell it is likely that your crystal is not very good.  Try indexing again with tighter tolerances on the data.  Click “OK” to continue.

- Choose the “Refine” button to refine your unit cell data and to select a Bravais lattice.  Initially the least-squares analysis of your matrix data will be quite intolerant (the slider will be very low down).  This is due to an algorithm behind the refinement program.  The number of reflections included in the refinement is in the middle of the window.  Usually one “click” on the slider will introduce a significantly larger number of reflections to improve the unit cell.  The upper panels display parameters relating to the instrument and crystal position.  The first two check boxes (crystal to detector distance and beam center X and Y) should be unchecked and the remaining three checked.  Click on the “Refine” button to refine the cell data once.

- Now select the “Bravais” button.  The program will give you its best guess as to the lattice type and Laue group of the cell.  Use your best judgement to decide if the program is correct. If you think it is otherwise, select the lattice/Laue group that you think is correct.  Once you are satisfied click on “OK” to return to the cell refinement window.  If at any stage you think you may have made a mistake, click on the “Cancel” button.  This will quit the refinement routine and not save the refined parameters.  If you click on “OK” it will save the newly refined parameters.

- You will now start a quick iterative cycle of cell refinement to improve the detector, goniometer and crystal position parameters.  With the newly selected cell, click on “Refine”.  Check the crystal to detector distance and beam center boxes and uncheck the detector roll, pitch and yaw parameter.  Refine these.  Step through the other options (crystal position and goniometer zeroes) in a similar fashion.  IF your cell refinement fails, the software will apparently lock up and/or you will see a pink panel appear in one or more of the parameter boxes.  Simply click on the “Cancel” button to quit the cell refinement.  Save your project, quit APEX-2 and then restart it.  You will be able to come back to where you were.  This is currently the only way to clear the “Egg-timer” that will appear.

- Cycle through several iterations of cell refinement with each of the detector/goniometer and crystal parameters until finally you can have all of the parameters refined together.  NOTE: Keep an eye on the frame window to the left, it will show the predicted spot locations.  You can decide if the cell is suitable and refining well visually.  That is, watch the green circles and make sure that they over-lap your reflections.  You can scroll through frames and blocks of data by using the arrow keys displayed at the top of the frame window.

- Once you are satisfied that you have obtained the best cell and orientation matrix that you are able to, you can save the P4P file by selecting File ( Export ( P4P File.  You will be prompted for an output file name (it defaults to the project name).  Click on “OK” to save the P4P file (of potential use for post facto data analysis).

- You are now ready to begin data collection.  Close the “Determine Unit Cell” window (click on the X) and return to the “Collect” ( “Experiment” window, “Setup Experiment” tab.  You need to turn off the matrix collection routine and enable the data collection routine.  To do this, click on the numeral “1” on the left-hand side.  While holding down the “SHIFT” key, click on the numeral “4”.  This will highlight the first four blocks of data.  Right-click on the “Active” field name and select “No”.  Repeat this for the blocks 5 through 8, back make them active instead (“Yes”).

- Click on the “Validate” button.  This will determine if the set of runs is viable or if it has potential collisions.  This is imperative to do if you modify the standard set of runs in any way.

- Once you are ready to start your data collection, uncheck the “Generate new Dark Frames” check-box, this will allow the Izero value to be recorded in the frame header. Then click on the “Execute/Resume” button and data collection will begin.  The computer will switch over to the ”Monitor Experiment” tab and display the frames as they are taken.

- Once your data has collected you will need to integrate. Select “Integrate” ( “Integrate Images”.  The window will open with a spread-sheet feel to it.  To import the frames that you collected click on the “Import Runs from Experiment” button.  You will probably need to set the resolution to a lower limit than the software decides (text box in upper right-hand corner).  You will need to choose the appropriate space group for the crystal system that you have (the default space groups for each setting are picked out with an asterisk next to them)

- Click on the “Refinement Options” button; you need to set the box size to 1.5 X 1.5 X 1.5 and uncheck the Enable Box Size Refinement.  Also set the “Frequency (Images)” number to 100 frames. You also need to set the local refinement to refine crystal translations (under the left-hand Refinement Parameters window).  When you are satisfied that these are correct, click on “OK” to accept these parameters.

- In the “Integration Options” you need to decide whether or not you wish to correct for beam decay.  If you do you need to check the “Enable Beam Monitor Normalization” box. Under “More Options” you will need to unselect “Multi-wire Reference Correction”.  Then click on “OK” to accept these changes.

- Once you are ready, click on the “Start Integration” button.  This will launch SAINT in the background and SAINT Chart in the foreground.  SAINT Chart is a graphic display of the integration progress.  Watch it for any unusual features.  It is simply a graphical representation of the usual numeric SAINT screen output.  Once the integration has finished, click on the “X” at the top right of the window to close the SAINT Chart window.  Close the Integrate window by clicking on the lower of the two “X” in the top right.

- Once the integration has completed you will need to scale and apply absorption corrections.  APEX-2 provides a GUI to the SADABS program.  Click on “Scale” ( “Scale”. This will open the GUI for SADABS.

- Check the “Additional Spherical Absorption Correction”.  For routine samples, given their frequently small size, a mu*r of 0.01 or 0.02 should be sufficient. Uncheck the Lambda\2 Correction.  The beamline has light polarized in a different fashion to laboratory sources.  Next click on the “folder” icon at the top right (next to the text box named: Base”).  This will open an additional window, select the first blocks RAW filename, check the “Open all with Base” box and click on “OK”.  Uncheck the “Merged Batches” box; this will then use all of the separate runs for absorption correction.  Finally click on “Next” to go to the “Parameter Refinement” tab.

- The “Parameter Refinement” tab shows data about your unit cell (as obtained from the final cell refinement, the wavelength of the radiation, total number of reflections, number of unique reflections etc … It is simply a graphical alternative to the SADABS interface.

- Depending on your usual lab practices, the mean I/sig(I) threshold should be set to 3 or 5, and Absorption Type to “Medium Absorber” (this applies a [6, 3] spherical harmonic correction to the data, more or less unusually shaped crystals should have a strong [8, 5] or weak absorber [4, 1] applied as necessary).  You can also modify the high angle cut-off if your data were integrated to too high an angle (filenamem.ls output will reveal this).  When you are ready, click on “Refine” to apply the absorption correction.  The graph on the lower left will show the applied correction and R(int) for the data.  It should refine out to a flat value.  To move to the final analysis of the data click on “Next”.

- You will now enter the “Error Model” portion of the program.  The defaults are acceptable.  Ignore the graphic displays when they come up.  They are just graphical interpretations of the incident and diffracted scale factors, R(int)’s for each block of data etc.  The titles are self-explanatory.  When you are ready to determine outlier rejection etc, click on “Determine Error Model”.  After this is completed, you may wish to re-refine the absorption correction parameters with new g-values.  This is not normally needed.  Finally click on “Finish” to write out the HKL file and prepare the usual SADABS graphs.  View the graphs to make sure that nothing untoward has happened to your data (if it has, go back and re-examine the frames for any obvious errors and re-determine the cell/re-integrate if necessary.  To quit the scaling/SADABS interface click on the “Exit AXScale” button.

- The penultimate function is the “Examine Data” ( “Space Group Determination” (XPREP).  This is precisely the same as the laboratory version.  You should be familiar with its functionality.  If not, go to the CHEXRAY web pages and help files found there-in. (http://xray.cchem.berkeley.edu).  NOTE: You will need to change the wavelength of the radiation used in the experiment (at the cell contents portion of the program).  We typically use 0.775 Å radiation (16 keV).

- Ultimately you will solve your structure.  You will need to solve and refine your structure using SHELX, as it is capable of incorporating scattering factors of elements for radiation other than molybdenum, copper etc.  You will need to use the program CROMER to determine the scattering factors for the elements present in your sample at the wavelength used (usually lambda = 0.775 Å).