# Epicel( (cultured epidermal autografts)

## Background

Although the skin is often not thought of as an organ, it is the largest organ system in the body and its specific and vital functions clearly define it as such.  These include:

- Protective

- Immunologic

- Fluid, protein, and electrolyte homeostasis

- Thermoregulation

- Neurosensory

- Metabolism

Skin is comprised of an epithelium (the epidermis), and a connective tissue matrix (the dermis).  The skin’s primary function is to act as a protective barrier between the body and the external environment.  The skin blocks penetration of microorganisms, absorbs and blocks radiation, and inhibits the loss of water from the body.  The skin also contains several types of sensory receptors (pain, touch, hot and cold), is involved in thermoregulation (prevents heat loss and allows for rapid cooling through evaporation of sweat), and functions in immunologic surveillance by assisting in the presentation of antigens to immune cells.

The epidermis forms the outermost portion of the skin and is approximately 0.04 to 1.5 mm in thickness compared with the 1.5 to 4 mm of full thickness skin.  It is a continuously regenerating, multilayered (stratified), squamous epithelium that keratinizes.  The epidermis is composed of primarily epithelial cells or keratinocytes, as well as other cell types, including melanocytes (pigment producing cells), Langerhans cells and Merkel cells which serve as mechanoreceptors.  The epidermis is organized into several layers named for either a structural property or a functional activity of the included cells.

The basal layer of keratinocytes contains young cells that are mitotically active, providing generation of outwardly migrating epidermal cell layers.  From the basal layer, cells move outward and mature in a process called differentiation.  As keratinocytes differentiate, mitosis no longer occurs and protein synthesis, predominantly of keratins, becomes the dominant process.  As the cells move outward, they lose their nuclei, flatten and evolve into a dead layer called the stratum corneum.  This layer, is a compact, relatively impervious layer that eventually desquamates.  The stratum corneum layer provides the majority of the skin’s barrier function.

The dermis is a relatively thick layer comprised mainly of fibrous connective tissue.  The primary cell type is the fibroblast, a spindle shaped cell which does not frequently replicate but is very active in producing extracelluar proteins, primarily collagen and elastin.  The dermis makes up the bulk of the skin and provides its structural strength.  The dermis protects the body from mechanical injury, binds water to serve as a water storage organ; it functions in thermal regulation and as a receptor of sensory stimuli.

## Burn Injuries

Burn injuries can compromise all of these essential functions and are classified in terms of degrees or depth of injury, whether or not they involve additional trauma or smoke inhalation.  First degree burns and superficial second degree burns involve only the epidermal layer and, perhaps, superficial portions of the dermal layer.  Healing will generally occur rapidly and completely within one to two weeks through migration to the surface of epithelial cells, which survive in deeper portions.

In deep partial thickness and full thickness burns (second and third degree burns), most or all of the dermis is destroyed.  Due to the depth from which migrating epithelial cells must travel, reepithelialization is greatly retarded or impossible.  These wounds are routinely treated with excision and grafting.

Fourth degree injury refers to situations where heat damage extends to deep structures, such as muscle, tendon, bone, etc.  Treatment may require elaborate debridement or even amputation.

## Complications of the Burn Injury

The burn wound may be complicated by additional trauma sustained in the course of the accident (e.g., fractures, lacerations from falls or lung damage from smoke inhalation).  Also an individual’s medical history will impact survival and morbidity.  Ryan et. al. reports smoke inhalation along with burn wound injury increases a patient’s risk factors for mortality,.

Permanent wound coverage must be achieved in a timely fashion to avoid the many complications of the burn wound injury.  Large open wounds cause fluid level imbalances, increased susceptibility to infection and reduced immune response. Lack of wound coverage and its sequelae can create an environment of sepsis, which can lead to multi-system organ failure and death.

## Existing Treatment Options

Generally a patient is admitted to a specialized burn treatment center if they meet criteria which has been established by the American Burn Association.  The criteria include:

- Second or third degree burns of 10% total body surface area (TBSA) in children <10 years old or adults >50 years old;

- Second degree burns of ( 20% TBSA at any age;

- Third degree burns (10% TBSA at any age;

- chemical burns that threaten functional or cosmetic impairment of face, hands, feet, genitalia, or perineum;

- Burns with inhalation injury.

Once the patient is admitted to a burn institution, an accurate estimate of the total body surface area burned is calculated and fluid resuscitation measures begin.  The wounds are treated with topical antimicrobial therapy, which in most cases is silver sulfadiazene cream.  The depth of the injury will determine if the wound will heal on its own with conservative measures (i.e. daily dressing changes) or whether surgical intervention is required.

The conventional treatment for burn wound closure is excision of the eschar or removal of the necrotic material and placement of a split thickness skin graft (STSG).  A STSG is harvested from an area of the individual’s unburned skin.  It includes the epidermal layer, which regenerates and a very thin portion of the dermal layer, which does not regenerate.  This donor skin is placed upon the freshly excised burn wound and closure ensues.  This may occur fairly rapidly in patients with small TBSA burns where there is abundant unburned skin to harvest.  The donor site may be used again after it heals which can range from 6 - 10 days, depending on the thickness of the harvested graft.

In burn injuries, sparing autograft dermis is a therapeutic goal.  Harvested with conventional methods, autograft has a total thickness of approximately 0.010-0.016 inches and contains a significant amount of dermis with the epidermal layer.  This is in contrast to thin epidermal autograft which is approximately one half that of conventional autograft and consists of the epidermis and a very small amount of dermis.  The total number of harvests of conventional autografts, which can be obtained from unburned parts of a patient’s body, is limited to approximately one to three procedures.  Limited availability of autograft, especially in large burns, can be managed by thin autograft harvesting techniques, which reduce the amount of dermal tissue harvested with the autograft.

In a larger TBSA burn, usually >30%, the donor skin must be meshed and expanded in an attempt to cover a larger surface area with a smaller amount of available skin.  If the patient does not have enough donor skin to cover the wound, the physician may need to consider alternative means of permanent wound closure in order to save the patient’ s life,.

The epidermal barrier, which is essential for life, is a central mechanism in preventing dehydration and microbial infection.  Traumatic injuries, such as burns, can result in the permanent loss of epidermal tissue through destruction of the proliferative (basal) keratinocyte layer.

## Epicel (cultured epidermal autografts)

Epicel is manufactured from epidermal cells harvested from uninjured skin, cultivated in vitro to create individual grafts which are then supported by a moisture-retaining petrolatum impregnated gauze dressing.  The grafts are thin, diaphanous sheets of epidermal cells that, when transplanted onto the wound bed, will lead to the reconstitution of a functional epidermal layer.

Epicel( is for use in patients who have sustained a deep dermal or full thickness burn covering 30% or more of their total body surface area.  Epicel provides a means to close full-thickness wounds when the injured area is so large that the number of potential donor sites is limited.  Full-thickness injuries are, those which completely destroy the regenerative capacity of the epithelium.  Re-establishment of the epithelial layer can only occur by harvesting partial-thickness skin grafts from autologous donor sites. In those cases in which the patient does not have enough skin available to close their wounds, Epicel is the only treatment option.  No other device is available which can provide these patients with a permanent autologous epidermal replacement.

During Epicel cultivation, the proliferative cells within the epidermis are isolated from a full thickness skin biopsy and rapidly expanded using the superior growth conditions of tissue culture such that sufficient grafts can be generated to cover the patient’s wounds.  The Epicel culture process utilizes routine mammalian cell culture techniques including an irradiated murine 3T3 cell feeder layer that supports early growth of epidermal keratinocytes.  As the keratinocytes proliferate on the culture flask, the feeder layer is displaced from the flask and removed during subsequent culture media changes.  A description of the characterization testing conducted on the 3T3 cells used in the Epicel™ manufacturing process can be found on page 9 of this document.

During Epicel graft cultivation the cells that comprise the nascent sheet tightly join as they approach confluence forming a protease-resistant cohesive unit.  At the same time, proliferating cells in direct contact with the tissue culture substrate initiate stratification creating a multilayered sheet.  The cell-cell contacts that bind the sheet together are a fundamental property of keratinocytes.  Epicel grafts are enzymatically released as intact sheets from the tissue culture substrate and attached to a petrolatum impregnated gauze dressing in such a way that the basal layer of cells will come into direct contact with the wound surface. The petrolatum gauze dressing provides structural support and prevents desiccation during transport of the grafts to the surgeon and through the initial phases of engraftment to the wound and epithelial maturation.  Engraftment of Epicel grafts leads to a reconstitution of the epidermal layer.

The primary benefit of the Epicel product is that the total number of grafts required to treat a patient can be produced from a single biopsy of unburned skin.  Patients suffering burns over a significant body surface area can be completely covered regardless of the amount of unburned skin available for split thickness skin grafts.  This minimizes the time to wound closure and minimizes the time in which the patient is most susceptible to serious and potentially life threatening complications.  For example, Munster et al reported on a series of patients treated with cultured epidermal autografts who had a significantly reduced mortality rate (14%) when compared with control patients (48%).

# The Importance of Epicel

Epicel (cultured epidermal autografts) is indicated for use in a subgroup that represents the most severely injured patients in the burn population.  Approximately 2700 patients present to burn centers each year with burns (30% of the total body surface area.  Treatment for these patients requires that the wounds be closed as rapidly as possible.  The Epicel product is one of the medical treatments in the surgeon’s armamentarium that allows for the treatment and closure of these wounds.  In those cases in which the patient does not have enough skin available to close their wounds, Epicel is the only treatment option.  No other device is available which can provide these patients with a permanent autologous epidermal replacement.

Although the incidence of these extensive burns is low, the production of an autologous epidermal replacement requires an extensive investment by the manufacturer.  The episodic nature of these injuries necessitates a well-maintained facility and staff that is ready to respond when required.

Genzyme Tissue Repair maintains a state of the art cell culturing facility that can respond to those fluctuations and consistently produce a quality product.  From the time of biopsy receipt, to the product release, the product is managed by an extensive set of standard operating procedures and well-trained personnel.  These technicians are knowledgeable in the in-process, final product, and environmental controls needed to ensure that the Epicel product to be grafted on the patient is clinically acceptable.

Products such as Epicel provide a clear benefit to a critically ill population of patients who might otherwise not survive with conventional treatment.  Because the patient population is very small, FDA granted Humanitarian Use Device designation for Epicel.  HUD designation also requires identification of the product cost and a demonstration that the amount charged does not exceed the cost of research, development, fabrication and distribution of the product.  There is a very high cost of developing and maintaining a state of the art program such as Epicel.  The additional requirements imposed by regulation of Epicel as a xenotransplant product in accordance with the expanded definition, would increase the burden to the patient, the burn center and the manufacturer and delay the time to treatment without necessarily increasing the benefit to the patient.

# Characterization of Swiss Mouse Embryo 3T3 Cell Bank used in Epicel( Autograft Cultivation

Testing of the 3T3 Master cell bank and Master Working cell bank were carried out by Microbiological Associates Inc. (now Bioreliance) following the requirements in FDA’s Points to Consider in the characterization of a cell substrate for production of a biological material.

## Master Cell Bank Testing

### Isoenzyme Analysis

The identity of a species of origin of a cell line can be determined on the basis of the electrophoretic mobilities and banding patterns of the isoenzymes of a number of intracellular enzymes including lactate dehydrogenase (LD), glucose-6-phosphate dehydrogenase (G6PD), malate dehydrogenase (MD), nucleoside phosphorylase (NP), and peptidase B (PepB).  Cells are burst and the proteins recovered run on an electrophoresis gel to determine the banding pattern.  The migration distances for each isoenzyme band are determined and compared to a reference table with standard migration distances for different species.  The results obtained on the cells from the Master cell bank were consistent with the expected species of origin for the test article (i.e., mouse).

### Bacterial and Fungal Contaminants

The direct inoculation method used for this testing meets or exceeds the USP 23 and/or 21 CFR 610.12 requirements for sterility testing.  Cells are lysed and lysate inoculated into fluid thioglycollate broth and soybean-casein digest broth (as required by the sterility test, 21CFR 610.12), pre-reduced peptone yeast glucose (PYG) medium (to detect anaerobic contaminants), and Sabouraud-dextrose agar (to detect fungal contaminants).  The test article showed no evidence of fungal or bacterial contamination in any of the culture conditions during the observation period.

### Mycoplasma

Mycoplasma testing was performed using a large sample volume and exceeds recommendations in the Points to Consider in the Characterization of Cell Lines Used to Produce Biologicals  (CBER 1993).  This assay is designed to detect the presence of agar-cultivable and non-cultivable mycoplasma species under aerobic and anaerobic growth conditions.  For agar-cultivable mycoplasma, a lysate of the test article was inoculated onto Type A and Type B agar.  For amplification of low level organisms, lysate material was inoculated into broth culture followed by subculture onto agar plates on days three, seven, and fourteen.  Under these conditions, the test article was negative for agar-cultivable mycoplasma.  For detection of non-cultivable mycoplasma, the test article was inoculated directly onto cultures of Vero cells.  After incubation for three to five days, the cultures were stained with Hoechst bisbenzimid stain and examined by ultraviolet microscopy.  All wells of Vero cells inoculated with the test article were negative for the presence of mycoplasma.

### In Vitro Assay for Viral Contaminants

The in vitro  assay employs three indicator cell lines (MRC-5, human diploid lung cells; Vero, African green monkey kidney cells; and NIH/3T3, Swiss mouse embryo cells) to screen for potential viral contaminants that may be present in the test article.  The indicator cells are inoculated with a low speed supernatant collected from a thawed cell lysate, and observed at least three times per week for the appearance of cytopathic effects (CPE).  In addition, indicator cell culture supernatants were tested for hemagglutination (HA) and hemadsorption (HAD) using chicken, guinea pig, and rhesus monkey erythrocytes at 4(C and 36(C.  No cytopathic effects, hemadsorption, or hemagglutination was observed in any of the indicator cell cultures inoculated with the test article.

### Test for the Presence of Inapparent Viruses

Since inapparent (latent) viruses may not always be detected by inoculating a battery of indicator cells and observing cytopathic effects or other indications, an in vivo  screening assay was performed on the test article.  Aliquots of the test article were inoculated into both adult and suckling mice, guinea pigs, and embryonated hens’ eggs.  In all cases, multiple inoculation routes were employed to maximize detection of infectious agents.  Adult mice and guinea pigs were observed for twenty eight days for clinical signs.  Suckling mice were observed for fourteen days, after which all surviving mice were euthanized using cervical dislocation.  A homogenate was prepared and inoculated into a new group of suckling mice by the same routes and in the same volumes as the original group.  These newly inoculated mice were observed for a period of fourteen days.  Embryonated eggs were inoculated via two routes, allantoic and yolk sac.  Following allantoic inoculation, the eggs were checked for viability at 1 and 3 days post-inoculation, and fluid collected from the eggs was tested for hemagglutination at 4(C (( 2(C) and 25(C (( 2(C) using chicken, guinea pig and human type O erythrocytes.  Pooled allantoic fluids were passaged to a new group of embryonated eggs and observed and tested in the same way.  Following yolk sac inoculation, eggs were checked for viability at 1 and 10 days post-inoculation.  At 10 days, pooled yolk sacs were passaged into a new group of embryonated eggs and were checked for viability at 1, 2  and 10 days post-inoculation.  All adult and suckling mice, and guinea pigs inoculated with the test article appeared normal and healthy during the observation period.  All hens’ eggs inoculated via the allantoic or yolk sac routes were viable and fluids were negative for hemagglutination.  No evidence of viral contamination was observed due to the test article.

### Presence of Murine Specific Adventitious Agents

The Mouse Antibody Production test determines whether any of sixteen murine specific adventitious agents were present.  Viral antibody free mice were inoculated with the test article via multiple routes to maximize the potential immune response in the animals.  Plasma or serum from one group of mice was collected within 10 days of inoculation and tested for lactate dehydrogenase (LDH) activity; elevated levels of this enzyme are indicative of the presence of lactic dehydrogenase virus (LDV).  This same group was subsequently inoculated with a known lethal strain of Lymphocytic Choriomeningitis Virus (LCMV), and monitored for morbidity and mortality for up to 3 weeks.  Animals would be protected from this lethal challenge only if antibody to this virus is present.  The other groups of mice were bled no sooner than 28 days post-inoculation and their serum assayed by ELISA, Hemagglutination Inhibition (HI) or Indirect Fluorescent Antibody (IFA) assays for the presence of antibody to known murine viruses.  All sera from animals inoculated with the test article were negative for the presence of antibody to Ectromelia, GDVII, Lymphocytic Choriomeningitis (LCM), Hantaan Virus, Minute Virus of Mice (MVM), Mouse Adenovirus, Mouse Hepatitis Virus (MHV), Pneumonia Virus of Mice (PVM), Polyoma, Reovirus Type 3, Epizootic Diarrhea of Infant Mice (EDIM), Mouse Salivary Gland Virus (Mouse Cytomegalovirus - CMV), K, Mouse Thymic Virus (MTV) and Sendai viruses as determined by ELISA, IFA, or HI.  All plasma from animals tested for lactate dehydrogenase activity showed normal levels except for LDV inoculated control animals, which showed elevated levels of LDH activity.  All animals challenged with LCM virus died within 10 days of challenge, indicating that they were not protected by antibody to LCMV produced in response to the original test inoculum.  Based on the data obtained in the assays performed, the test was shown to be free of all of the 16 murine viruses for which it was examined.

### Presence of Bovine Viruses

An in vitro assay was performed to determine whether the cells were contaminated with any of five bovine viruses.  Indicator cells specific for these viruses, bovine turbinate cells, were inoculated with cell lysate prepared from the test article and cultured for a period of 14 days.  During this time, all cultures were monitored for cytopathic effect (CPE).  Any cultures exhibiting CPE were harvested and fixed when approximately 50% of the monolayer was involved.  All remaining cultures were harvested and fixed after the 14 day incubation.  The cells were then examined by fluorescent antibody staining techniques for the following bovine specific viruses: Bovine Viral Diarrhea virus (BVD),  Bovine Adenovirus Type 3 (BAV3), Bovine Parvovirus (BPV), Infectious Bovine Rhinotracheitis virus (IBR), and Bovine Parainfluenza Virus Type 3 (PI3).  None of the cultures inoculated with the test article exhibited CPE.  Fixed cells did not exhibit fluorescence when tested with antisera specific for the 5 viruses used in this analysis. Based on these data, the test article was determined to be free of the 5 specific bovine viruses for which it was examined.

### Retrovirology:  Presence of Xenotropic Murine Retrovirus

The extended S+L- focus assay was used to test the cells for the presence of infectious xenotropic retrovirus.  The development of foci in the mink S+L- cells is indicative of the presence of xenotropic retrovirus contamination.  The test article was inoculated directly onto mink S+L-  cells as well as onto mink lung cells (extended portion) for amplification of potential low level contaminants.  After 2 passages on the mink cells, culture fluids were harvested and assayed in the S+L- focus assay.  No foci were observed on either the direct or extended portions of this assay for cultures inoculated with the test article. Based on these data, the test article was found to be negative for xenotropic murine retrovirus.

### Presence of Ecotropic Murine Retrovirus

The extended XC plaque assay was used to test the cells for the presence of infectious murine ecotropic retrovirus (both N-tropic and B-tropic).  The test article was inoculated onto SC-1 cells for both the direct and extended portions of the assay.  The extended portion is used to amplify potential low level contaminants.  In the direct portion, the cultures were incubated for 5-6 days, irradiated, and overlaid with XC cells.  In the extended portion, the cells were passaged twice, after which the culture fluids were inoculated onto fresh SC-1 cells and processed as in the direct assay.  No plaques were observed in either the direct or extended portions of the assay for cultures inoculated with the test article.  Based on these data, the test article was found to be negative for murine retrovirus.

### Presence of Viral-like Particles

Cells are examined by transmission electron microscopy (TEM) for the presence of virus-like particles.  Cells are pelleted and fixed with Trump’s fixative.  Thin sections of the cell pellet are cut, mounted, and stained with 5% methanolic uranyl acetate and Reynold’s lead citrate.  The cell types are characterized via a morphological examination.   Two hundred test article cells are examined for any particles with a virus-like morphology.  The number of viral particles are tabulated and the morphology of each particle examined to determine whether it is A-, B-, C-, D-, and R-type retrovirus.  Examination of the fixed test article cells revealed no identifiable virus-like or retrovirus-like particles.

### Presence of Retrovirus

A lysate prepared from the test article was assayed for retroviral-derived reverse transcriptase activity as an indicator for potential retrovirus contamination.  The test article was incubated in reaction mixtures containing the appropriate concentrations of Mg++ or Mn++ to differentiate between C type and D type retrovirus reverse transcriptase.  All reaction mixtures containing the test article (diluted or undiluted) incorporated less than 500 cpm of 3H-thymidine above background.  The diluted test article (1:2) did not inhibit 3H-thymidine incorporation in a control mixture containing Rauscher Murine Leukemia Virus (R-MuLV).   Based on these data, the test article showed no evidence of the presence of type C or type D retrovirus reverse transcriptase activity.

### Tumorigenicity

The ability of the test article cells to form colonies in soft agar was assessed as an indicator of possible transformation and the potential ability for the cell line to form tumors in animals.  A single-cell suspension of test article cells were inoculated in a soft agar overlay of a hard agar base.  Cultures were monitored for 28 days for the formation of colonies of greater than 10 cells.  No colonies formed in cultures inoculated with the test article cells.

## Master Working Cell Bank Testing

### Bacterial and Fungal Contaminants

The direct inoculation method used for this testing meets or exceeds the USP 23 and/or 21 CFR 610.12 requirements for sterility testing.  Cells are lysed and lysate inoculated into fluid thioglycollate broth and soybean-casein digest broth (as required by the sterility test, 21CFR 610.12), pre-reduced peptone yeast glucose (PYG) medium (to detect anaerobic contaminants), and Sabouraud-dextrose agar (to detect fungal contaminants).  The test article showed no evidence of fungal or bacterial contamination in any of the culture conditions during the observation period.

### Mycoplasma

Mycoplasma testing was performed using a large sample volume and exceeds recommendations in the Points to Consider in the Characterization of Cell Lines Used to Produce Biologicals  (CBER 1993).  This assay is designed to detect the presence of agar-cultivable and non-cultivable mycoplasma species under aerobic and anaerobic growth conditions.  For agar-cultivable mycoplasma, a lysate of the test article was inoculated onto Type A and Type B agar.  For amplification of low level organisms, lysate material was inoculated into broth culture followed by subculture onto agar plates on days three, seven, and fourteen.  Under these conditions, the test article was negative for agar-cultivable mycoplasma.  For detection of non-cultivable mycoplasma, the test article was inoculated directly onto cultures of Vero cells.  After incubation for three to five days, the cultures were stained with Hoechst bisbenzimid stain and examined by ultraviolet microscopy.  All wells of Vero cells inoculated with the test article were negative for the presence of mycoplasma.

### In Vitro Assay for Viral Contaminants

The in vitro  assay employs three indicator cell lines (MRC-5, human diploid lung cells; Vero, African green monkey kidney cells; and NIH/3T3, Swiss mouse embryo cells) to screen for potential viral contaminants that may be present in the test article.  The indicator cells are inoculated with a low speed supernatant collected from a thawed cell lysate, and observed at least three times per week for the appearance of cytopathic effects (CPE).  In addition, indicator cell culture supernatants were tested for hemagglutination (HA) and hemadsorption (HAD) using chicken, guinea pig, and rhesus monkey erythrocytes at 4(C and 36(C.  No cytopathic effects, hemadsorption, or hemagglutination was observed in any of the indicator cell cultures inoculated with the test article.

### Test for the Presence of Inapparent Viruses

Since inapparent (latent) viruses may not always be detected by inoculating a battery of indicator cells and observing cytopathic effects or other indications, an in vivo  screening assay was performed on the test article.  Aliquots of the test article were inoculated into both adult and suckling mice, guinea pigs, and embryonated hens’ eggs.  In all cases, multiple inoculation routes were employed to maximize detection of infectious agents.  Adult mice and guinea pigs were observed for twenty eight days for clinical signs.  Suckling mice were observed for fourteen days, after which all surviving mice were euthanized using cervical dislocation.  A homogenate was prepared and inoculated into a new group of suckling mice by the same routes and in the same volumes as the original group.  These newly inoculated mice were observed for a period of fourteen days.  Embryonated eggs were inoculated via two routes, allantoic and yolk sac.  Following allantoic inoculation, the eggs were checked for viability at 1 and 3 days post-inoculation, and fluid collected from the eggs was tested for hemagglutination at 4(C (( 2(C) and 25(C (( 2(C) using chicken, guinea pig and human type O erythrocytes.  Pooled allantoic fluids were passaged to a new group of embryonated eggs and observed and tested in the same way.  Following yolk sac inoculation, eggs were checked for viability at 1 and 10 days post-inoculation.  At 10 days, pooled yolk sacs were passaged into a new group of embryonated eggs and were checked for viability at 1, 2  and 10 days post-inoculation.  All adult and suckling mice, and guinea pigs inoculated with the test article appeared normal and healthy during the observation period.  All hens’ eggs inoculated via the allantoic or yolk sac routes were viable and fluids were negative for hemagglutination.  No evidence of viral contamination was observed due to the test article.

### Tumorigenicity

The ability of the test article cells to form colonies in soft agar was assessed as an indicator of possible transformation and the potential ability for the cell line to form tumors in animals.  A single-cell suspension of test article cells were inoculated in a soft agar overlay of a hard agar base.  Cultures were monitored for 28 days for the formation of colonies of greater than 10 cells.  No colonies formed in cultures inoculated with the test article cells.

## Production Cell Bank Testing

Production banks of 3T3 cells undergo raw material Quality Control testing prior to acceptance for use in the Epicel( manufacturing process.

### Source

Documentation accompanying each lot of Production 3T3 cells must specify that these cells were grown from the 3T3 Master Working Cell bank.

### Cell Number

Documentation accompanying each lot of Production 3T3 cells must specify that there are > 1.4 x 108 cells/vial.

### Bacterial and Fungal Contaminants

Sterility testing is performed in-house by GTR’s Quality Control Department on 5% of each lot of Production 3T3s.  The direct inoculation method used for this testing is done according to USP “Sterility Tests” <71>.  Cells from each test vial are inoculated into and tryptic soy broth to test for the presence of aerobic and fungal contaminants and fluid thioglycollate medium to test for the presence of anaerobic contaminants.  The tryptic soy broth is incubated for fourteen days and monitored for turbidity, which is an indicator of microbial growth.  Because the fluid thioglycollate medium has a gel base and is thicker than tryptic soy broth, the inoculated 3T3s do not settle and the medium remains cloudy.  Therefore, in accordance with USP, between days 3 and 5, 1 mL of the medium is withdrawn and directly inoculated into another bottle of fluid thioglycollate medium.  This medium is then incubated for fourteen days and monitored for turbidity.  The test article must show no evidence of fungal or bacterial contamination in any of the culture conditions during the observation period to be acceptable for use.

### Mycoplasma

Mycoplasma testing is performed using three vials of cells from each Production cell bank.  For each vial, five flasks of cells are plated and grown in antibiotic-free medium.  Cells are scraped from the flasks and the cells and supernatants pooled.  Two aliquots from the pooled sample are sent to Bionique Testing Laboratory for mycoplasma testing in accordance with the procedure outlined in the Points to Consider in the Characterization of Cell Lines Used to Produce Biologicals  (CBER 1993).  This assay is designed to detect the presence of agar-cultivable and non-cultivable mycoplasma species under aerobic and anaerobic growth conditions.  All samples must be negative for mycoplasma in order to be acceptable for use.

## Irradiated Production 3T3 Testing

Production 3T3 cells were tested for retroviral activation, as a result of the cells receiving a radiation dose of 6000 rads during the manufacturing process.  An amp of Production 3T3 cells was thawed and irradiated in accordance with current manufacturing practices.  The irradiated cells (either frozen or in monolayer culture) were sent to Bioreliance to undergo various retroviral tests.

### Retrovirology:  Presence of Xenotropic Murine Retrovirus

The extended S+L- focus assay was used to test the cells for the presence of infectious xenotropic retrovirus.  The development of foci in the mink S+L- cells is indicative of the presence of xenotropic retrovirus contamination.  The test article was inoculated directly onto mink S+L-  cells as well as onto mink lung cells (extended portion) for amplification of potential low level contaminants.  After 2 passages on the mink cells, culture fluids were harvested and assayed in the S+L- focus assay.  No foci were observed on either the direct or extended portions of this assay for cultures inoculated with the test article. Based on these data, the test article was found to be negative for xenotropic murine retrovirus.

### Presence of Ecotropic Murine Retrovirus

The extended XC plaque assay was used to test the cells for the presence of infectious murine ecotropic retrovirus (both N-tropic and B-tropic).  The test article was inoculated onto SC-1 cells for both the direct and extended portions of the assay.  The extended portion is used to amplify potential low level contaminants.  In the direct portion, the cultures were incubated for 5-6 days, irradiated, and overlaid with XC cells.  In the extended portion, the cells were passaged twice, after which the culture fluids were inoculated onto fresh SC-1 cells and processed as in the direct assay.  No plaques were observed in either the direct or extended portions of the assay for cultures inoculated with the test article.  Based on these data, the test article was found to be negative for murine retrovirus.

### Presence of Viral-like Particles

Cells are examined by transmission electron microscopy (TEM) for the presence of virus-like particles.  Cells are pelleted and fixed with Trump’s fixative.  Thin sections of the cell pellet are cut, mounted, and stained with 5% methanolic uranyl acetate and Reynold’s lead citrate.  The cell types are characterized via a morphological examination.   Two hundred test article cells are examined for any particles with a virus-like morphology.  The number of viral particles are tabulated and the morphology of each particle examined to determine whether it is A-, B-, C-, D-, and R-type retrovirus.  Examination of the fixed test article cells revealed no identifiable virus-like or retrovirus-like particles.

### Presence of Retrovirus

A lysate prepared from the test article was assayed for retroviral-derived reverse transcriptase activity as an indicator for potential retrovirus contamination.  The test article was incubated in reaction mixtures containing the appropriate concentrations of Mg++ or Mn++ in order to differentiate between C type and D type retrovirus reverse transcriptase.  All reaction mixtures containing the test article (diluted or undiluted) incorporated less than 500 cpm of 3H-thymidine above background.  The diluted test article (1:2) did not inhibit 3H-thymidine incorporation in a control mixture containing Rauscher Murine Leukemia Virus (R-MuLV).   Based on these data, the test article showed no evidence of the presence of type C or type D retrovirus reverse transcriptase activity.

## End-Of-Production 3T3 Testing

Testing was conducted on production 3T3s which had been grown several population doublings beyond the point at which they would be used in the Epicel( process.  An amp of Production 3T3 cells was thawed in accordance with the current manufacturing process and plated into culture using the medium which is used by Quality Control to grow 3T3 cells for raw material testing.  The cells were sent to Bioreliance (tumorigenicity) and Applied Genetics Laboratory (karyology) for testing.

### Tumorigenicity

The ability of the test article cells to form colonies in soft agar was assessed as an indicator of possible transformation and the potential ability for the cell line to form tumors in animals.  A single-cell suspension of test article cells were inoculated in a soft agar overlay of a hard agar base.  Cultures were monitored for 28 days for the formation of colonies of greater than 10 cells.  The cultures were equivalent to that acceptable for negative control.

### Karyology

The cells underwent karyological analysis to determine the number of chromosomes and to check for the presence of abnormalities.  Cells are swollen, fixed onto slides, and stained with Giemsa stain.  One hundred cells in metaphase were examined for chromosome count.  Fifty of these metaphases were examined for cytogenetic aberrations.  The types of chromosomal aberrations present were classified.  Two chromosome aberrations were found with 2% cells aberrant.  Cytogenetic analysis showed that the cells were of mouse origin.  There was a wide variety of normal chromosomes and chromosomal rearrangements, as is consistent with a cell line.  The modal chromosome number was 60 (range 47 to 128) and chromosomes 18 and Y were absent in all karyotypes.  There results are consistent with the results of karyology on a sample of this cell line deposited with ATCC, which had a modal number of 68, no Y chromosomes, and very few cells with chromosome 18.

## Characterization of Epicel( Final Product

Characterization testing was conducted on three lots of Epicel( final product.  Because the final product was an intact sheet of cells which could not be disaggregated without damaging the cells, testing was performed on keratinocytes in culture prior to achieving confluence (formation of an intact sheet).  Lots of Epicel( were grown by Manufacturing in accordance with current practices.  The 3T3s were removed from the culture flasks by washing and shaking with EDTA.  The culture flasks with only keratinocyte colonies remaining were sent to Bioreliance (tumorigenicity) and Applied Genetics Laboratory of Genzyme Genetics (karyology) for testing.

### Tumorigenicity

The ability of the test article cells to form colonies in soft agar was assessed as an indicator of possible transformation and the potential ability for the cell line to form tumors in animals.  Typically, a single-cell suspension of test article cells is inoculated in a soft agar overlay of a hard agar base.  Because the sample consisted of keratinocyte colonies which did not easily disaggregate into single cells, the protocol was modified to include a baseline colony count of the agar plates at 24 hours.  All cells in the entire plate were examined.  If any aggregates greater than 10 cells were present at 24 hours, these were attributed to incomplete disaggregation, and the test was discontinued.  Cultures were monitored for 28 days for the formation of colonies of greater than the original baseline.  No colonies formed in cultures inoculated with the three lots of test article cells.

### Karyology

Karyological analysis was conducted on cells to determine the number of chromosomes and check for the presence of abnormalities.  Cells are swollen, fixed onto slides, and stained with Giemsa stain.  Twenty metaphases were examined for chromosome count and three metaphases were karyotyped.  In the one lot analyzed by Genzyme Genetics, the results showed a normal 46,XX karyotype consistent with the patient’s known sex.  There was no evidence of clinically significant numerical or structural chromosome abnormalities, although this method would not detect small rearrangements, microdeletions, and low level mosaicism.  Karyotyping studies on the last two lots are on-going at this time at Applied Genetics Laboratory.

# Endnotes

Herndon DN.  Total Burn Care, 1996 Saunders Publ. Pg 35-36

Ryan C, Schoenfeld D, Thorpe W.  Objective Estimates of the Probability of Death From Burn Injuries.  NEJM  February 1998; 362-366

Parks S.  Inhalation Injury in Burn Patients.  Trauma Rounds.  West J. Med.  1976; 124:244-248

Herndon DN.  Total Burn Care.  1996 Saunders Publ.  Pg 95-96

Collini FJ, Kealy GP.  Burns:  A Review and Update.  Contemp Surg. 1989; 34:75-77

McAree KG, Klein RL, Boeckman CR.  The Use of Cultured Epithelial Autogratfs in the Wound Care of Severely Burned Patients.  J. Pediatr Surg. 1993; 28:166-168

Munster A.  Cultured Skin for Massive Burns.  Annals of Surg. 1997; Vol 224 No. 3 372-377

Green H, Kehinde O, Thomas J (1979) Growth of cultured human epidermal cells into multiple epithelia suitable for grafting. PNAS USA 76(11): 5665-5668.

Leigh IM and Watt FM (1994)  The culture of human epidermal keratinocytes. In: The Keratinocyte Handbook, Chapter 2, Leigh, I, Lane B, and Watt F, eds, Cambridge University Press,  pp 43-51.

O’Conner NE, Muliken JB, Banks-Schlegel S, Kehinde O, and Green H (1981) Grafting of burns with cultured epithelium prepared from autologous epidermal cells. Lancet I: 191-199.

Munster AM.  Cultured Skin for Massive Burns A Prospective, Controlled Trial.  Annals of Surgery 1996; 224:372-377