Data Acquisition and Processing on BL11.3.1

Last Updated April 2005, S. Fakra

How to acquire XRD data

- Start Database by double clicking on the Database shortcut icon on the desktop

- Start APEX2 software by double clicking on the APEX2 shortcut icon on the desktop.

- Logon as user: guest, Pwd: guest. The APEX main window appears.

- Select from the menu: Instrument -> Connect, a Host name IP# will appear, just click on connect.

- Select file-> new and type the title of your experiment and/or your name and date

- Click on the ‘Collect’ tab located on the bottom left side of the window and then click on the Experiment button on the top left side of the window

- A window ‘Set up experiment’ appears. Enter a filename for your sample

- Check the box ‘attenuate if topped’ (should be on by default)

- Check the box ‘Correlate Frames’ for long exposures (>10s)

- Uncheck the box ‘Generate new Dark Images’

- Enter a default time: that will be the exposure time you want

- Go to ‘Operation’ column on the 1st row and select from the pull down menu ‘wavelength’ and the wavelength is then loaded, typically 0.73Å (17keV).

- Go to the 2nd row and select from the pull down menu ’Still’

- Go under the omega column and set it to -90 and under the phi column set it to 0, 2theta should be set to 0.

- Click on validate (only once at the beginning of your run)

- Click on ‘Execute’

- After few seconds the ‘Monitor Experiment’ window will appear and the frame will be displayed when the acquisition/processing is finished.

- DO NOT OVERWRITE THE FILE NAME, as you’ll get no answer if you do so.

- USE ONE OF THE FOLLOWING EXPOSUES TIMES. The pre-recorded dark frames are for exposure times for the 1024x 1024 bining mode of: 0, 1, 2, 3, 5, 8, 9, 10, 15, 16, 20, 24, 27, 29, 30, 40, 45,47,48, 49, 50, 52, 55, 60, 70, 73, 80, 84,90, 100, 110, 120, 240, 300, 600, 800 seconds. If you use a different exposure time than those, expect a large delay (several minutes to more) as the software will start to record automatically a Dark frame by doing 9 exposures!

- When you’re done with 11.3.1 XRD data acquisition, exit the APEX2 program, click on the menu: file-> exit and save the changes you made since you saved your session the last time.

- Then double-click on the “Stop Database” shortcut icon on the desktop, to close the database.

Analysis of data with Fit2D:

- To access your data, go the analysis computer (BL1131-diff), and double click on the shortcut “1131data”.

- You need to import your data on the C:\Users\your name directory on the BL1131-diff computer.

- You then need to convert your .sfrm data into a bruker format. To do so:

- Double click on the “FrmUtility” shortcut icon on the desktop.

- Browse for the file you want to convert

- Enter a filename for the converted frame in the output box (usually copy/paste the core name of your original .sfrm file) and end it with a .000 extension, so that  Fit2D can recognize it directly as a Bruker file.

- Select “Bruker format”

- Click on ‘convert’. A pop-up window will appear showing that the file has been converted successfully.

- Double click on the Fit 2d icon on the desktop and then click on “I accept”

- Choose X dimension=Y dimension= 1024 typically, then click on Ok.

- Click on Powder Diffraction (2-D) and then click on imput. Click on UPDIR until you get to your data folder.

- Open your converted frames. Don’t forget to apply a spatial distorsion correction and to load the “SpatialOct04.spline” file to do so. Then you can mask the beamstop and corners (using the mask and comparison options typically) and integrate directly your data. The integration results are generally saved as a .chi file (2D txt file).

Analysis of data with XMAS:

- Copy the spatial distorsion correction file NewAppLabXXX._ix into your data directory.

- Double click on FrmUtility icon on the desktop of BL1131-diff computer.

- Browse for the file you want to unwarp.

- Click on the ‘Unwarp frame’ tab

- From the pulldown menu select “36.231” for PixelsPerCM  and “0.726 for CMTo Grid

- Then select in the pulldown menu the Spatial distorsion correction file and press load

- Click on Unwarp: the Unwarped frame is saved over the original file. The original file XXX.sfrm is saved a XXX.sfrm.bak.

- Double click on the XMAS icon on the desktop of the BL1131-diff computer. XMAS reads directly .sfrm files, so no need to convert.

- Click on file-> set directories. 1st directory line should be C:\analysis\parameters. The 2nd directory line could be: C:\Users\your name\datadate\Xmasanalysis

- Then click on file->open image.

- Then click on image and rotate it twice, so that to get the beamstop location at the bottom of the image.

- Proceed with analysis of the frame.