DYEnamic ET-Terminator Cycle Sequencing

on Rolling Circle Amplified DNA

10μl Automated Reaction

Version Number:		2.1

Production Start Date:	06/01/03

Version 2.1 Date:		08/23/04

Authors:			Arshi Khan, Nancy Hammon, Catherine Gordon,

Reviewed/Revised by:	Amber Nivens

# Summary

The DYEnamic ET-Terminator sequencing kit is used to cycle sequence plasmid DNA, which has been amplified by rolling circle amplification, in preparation for sequencing on the MegaBACE 4000 instruments.

# Materials & Reagents

| Materials/Reagents/Equipment Disposables 384-well PCR Microplate, Clear 50-ml Centrifuge Tube PlateLoc Peelable Heat Seal LTS Pipette Tips (2, 10, 100, 200, 1000μl) Reagents DYEnamic ET-Terminator Sequencing Premix FW Sequencing Primer RV Sequencing Primer Coulter Clenz 100% Ethanol (200 Proof) Milli-Q Water Equipment 384-well Hydra Twister Hydra Wash Reservoir Cavro Dispensing Robot Pipettes (2, 10, 100, 200, 1000μl) GeneAmp PCR System 9700 PTFE Release Agent Dry Lubricant PlateLoc Thermal Plate Sealer 250 ml, 500 ml bottles 1000 ml beakers Multi-channel 250µl Pipette 250-ml Double pharmaceutical graduates | Vendor Axygen/ISC BioExpress Corning/VWR Velocity 11 Rainin GE Healthcare* IDT IDT Beckman Coulter AAPER Alcohol (LBNL) Millipore Milli-Q System Robbins Scientific Robbins Scientific JGI Rainin Perkin Elmer (Applied Biosystems) Miller-Stephenson Velocity 11 Pyrex Pyrex Matrix Falcon/VWR | Stock Number T-3059-1CS - 430290 06643-001 - 93-81096 - - PN 8546930 6810-13289 - - - - - - - MS-122V - - - |
| --- | --- | --- |

* GE Healthcare, formerly Amersham Biosciences

# Calibration Check

Hydra Twisters

NOTE: This should be performed on one FW and one RV plate within the first 10 plates per batch per Hydra Twister at step 2.12.

- Without disturbing the Hydra Twister, retrieve an empty FW destination plate from Stacker 4.

- Place the empty FW destination plate on the weighing balance and press the TARE button.

- Return the tared empty FW destination plate back to Stacker 4.

- After the tared FW destination plate has been aliquoted, remove it from Stacker 5 and replace it with an empty Axygen plate.

- Reweigh the aliquoted FW destination plate on the weighing balance.

NOTE 1: If the weight ranges fall outside of 0.3366g – 0.3743g, WARNING.

NOTE 2: If the weight ranges fall outside of 0.3272g – 0.3837g, DO NOT USE.

- Enter the weight into the Plate Weight Excel Spreadsheet.

NOTE 1: To open excel spreadsheet click on Shortcut to Plate Weight on desktop. If shortcut is unavailable go to Network Places on desktop and click on Entire Network( click on Entire Contents( click on Microsoft Windows Network( click on jgi_psf( click on Octopus( click on ProdSeq( click on RCA( click on Calibrations( click on Hydra Twister Calibration ( click on Plate WeightData.

NOTE 2: If the weight entered in the excel spreadsheet is too close to the high or low range, the number entered in the weight column will be highlighted in a red box and the text WARNING will appear in orange under the Spec. column.   If this occurs, continue to use the Hydra but monitor it carefully so that it does not fall out of range.  If it is continuously high or low and in the WARNING zone, contact your supervisor to determine how to proceed.

NOTE 3:  If the weight entered in the excel spreadsheet falls out of the range, the number entered in the weight column will be highlighted in a red box and the text DO NOT USE will appear in a red box under the Spec. column.  If this occurs, discontinue use of the hydra and submit a work request for calibration.

- Remove the empty Axygen plate from Stacker 5 on the Hydra Twister and replace it with the aliquoted weighed FW destination plate.

- Repeat steps 1 – 7 with a RV destination plate.

- Click Save on the Excel Spreadsheet.

Cavro 1

NOTE: This should be performed on the first plate of each batch dispensed during step 3.6c.

- Place the first FW sample plate on the weighing balance and press the TARE button.

- Place the first RV sample plate on the weighing balance and note the difference with the FW plate.

- Place both plates onto the appropriate Cavro stage.

- Click on START.

  - Watch for droplets on the tips and missed wells on the plate.

  - Visually check the bottom of the plate to make sure that the volumes in all of the wells are uniform.

NOTE:  If tips are clogged, wells are missing, or volumes are not dispensed correctly, submit a work request & call Instrumentation to get the problem fixed.

- Upon completion of dispensing, reweigh both plates to check the volume of sequencing chemistry cocktail that is being dispensed and record the data in the Cavro Excel spreadsheet.

NOTE 1: To open excel spreadsheet click on Shortcut to Cavro Calibrations on desktop. If shortcut is unavailable go to Network Places on desktop and click on Entire Network( click on Entire Contents( click on Microsoft Windows Network( click on jgi_psf( click on Octopus( click on ProdSeq( click on RCA( click on Calibration( click on Cavro Calibrations( select Cavro 3 worksheet.

- If weights are all within the range, continue to dispense all the plates in the batch. Acceptable dispense range for 9ul of ET-Terminator chemistry is 3.415-3.497g.

NOTE 1: If the weight entered in the excel spreadsheet is too close to the high or low range, the number entered in the weight column will be highlighted in a red box and the text WARNING will appear in orange under the Spec. column.   If this occurs, continue to use the cavro but monitor it carefully so that it does not fall out of range.  If it is continuously high or low and in the WARNING zone, contact your supervisor to determine how to proceed.

NOTE 2:  If the weight entered in the excel spreadsheet falls out of the range, the number entered in the weight column will be highlighted in a red box and the text DO NOT USE will appear in a red box under the Spec. column.  If this occurs,  submit a work request to instrumentation for calibration.

# Procedure

NOTE 1:  Make sure to wear gloves throughout the entire Sequencing Chemistry process.

NOTE 2: Heat Inactivation (Step 8 of the Rolling Circle Amplification on Plasmid Glycerol Stock protocol) must be completed prior to aliquoting the sequencing template into the FW and RV plates.

NOTE 3: If you are resequencing plates, go to the Venonat database ( click on RCA DB ( click on Scan in RCA plates for Re-Sequencing only! ( enter in: operator and direction ( click on SUBMIT ( print the barcodes and affix to 384-well PCR plates.

- Sample Preparation

1.1   	Remove the DYEnamic ET-Terminator Sequencing Premix Kit and respective FW/RV

primers from the -20(C freezer #202 located in Rm 140.

1.2    	Let the reagents thaw on ice.

NOTE 1: Make sure to keep different primers in different ice buckets.

NOTE 2: The ET-Terminator Kit usually takes 30 minutes to thaw.

- Aliquoting (Hydra Twister Robot)

NOTE:  Make sure to randomize aliquoting batches with at least 2 or more RCA batches that have the same primer association.

  - At the completion of Heat Inactivation (Step 8 of the Rolling Circle Amplification on Plasmid Glycerol Stock protocol) FW and RV destination barcodes were printed:

- Mark the side of these barcodes with a purple stripe, making sure not to cover the barcode itself.

- Affix these FW and RV barcodes to new Axygen plates (on the numbered and indented side of the plate).

NOTE: Make sure that the barcodes on the destination plates are marked with the correct primer association color against the database and against the tubes you pulled (key located on the countertop).

  - Obtain the RCA source plates from final heat off of the thermocyclers or from the 4oC deli in the Sequencing Preparation room (Rm140) and remove the foil seals.

Checkpoint:  Visually check the RCA source plate after removing the seal to ensure that the proper volume is in each plate.

  - Match the source plates with their corresponding destination plates (the stacks of plates should be in the same order).

  - In the Venonat database, click on RCA DB ( STEP 4A - Scan RCA Plates to Update Aliquot   Information ( enter in:

- Select the correct information from the drop-down menus: Hydra, Aliquot operator.

- Use the barcode scanner to scan in the RCA source plates.

- Click on SUBMIT.

- Scan in the corresponding FW and RV plates and click on SUBMIT.

Checkpoint: If there are any mismatched plates correct the mismatches before loading these plates onto the Hydra Twister stackers.

  - Place the RCA templates (source plates) in Stacker 1 and the destination plates (FW and RV pairs) in Stackers 3 and 4 (load stacker 4 up with a maximum of 32 plates before loading stacker 3).  The top source plate should match the top destination plate in stacker 4. To distinguish Twister 2 plates from Twister 1 plates, add black dots to the FW and RV barcodes of Twister 2 plates.

NOTE: Press down the emergency stop button located next to the Hydra Twister to manually move the Twister arm.

  - At the beginning of each week, move the supply and drain Masterflex tubing to a new slot and write the date on the tubing.

NOTE: There are 4 slots on the tubing. When all slots have been filled once, submit a work request to have the tubing changed (approx. every month).

  - At the beginning of each day, turn on the Hydra Twister computers, log on, then push down and twist up each emergency stop button located next to the Hydra Twister.

  - Open the Aliquot AUTOSTART1 program on the computer monitor.

  - Run the WASH_100μlx3 program.

  - After completion of the WASH_100μlx3 program, exit the program.

NOTE: Each time a task is completed in the Aliquot Autostart1 program, you must close the program and re-open it.

  - Re-open Aliquot AUTOSTART1 and click on Aliquot Axygen 384’s.

Checkpoint: The program will ask if there is enough water. Check the supply container and if it is less than half full, press the START button located above the supply container. If the water container is greater than half full, click OK on the program window. This must be done for every batch.

NOTE 1: This program will aliquot 1μl per well from the RCA plates into each well of the corresponding FW and RV plates in preparation for sequencing.

NOTE 2: Periodically monitor the water trough to ensure water is filling and draining correctly. If the water trough is not filling or draining properly or if the program was stopped abruptly then manually fill or drain the trough by pushing the corresponding button towards manual on the Twister Hydra Pump Controller located next to the Hydra Twister. Make sure to switch the button back to automatic when finished.  If the water trough is not filling or draining properly, also submit a work request.

  - Weigh one FW and one RV plate from each Hydra Twister (Refer to Hydra Twisters under the Calibration Check section).

  - After the program is complete and the pump has stopped running, close the window and push down and twist up the emergency stop button.

NOTE: This will move the robotic arm so that the RCA source plates may be removed.

  - Remove source and destination plates from the Hydra-Twister stackers.

  - Re-open Aliquot AUTOSTART1 and run the WASH_100μlx3 program.

NOTE:  This program must be run in between each batch/run.

  - For the source RCA plates:

- Place clear adhesive seals on all source RCA plates.

- From each RCA batch:

- Isolate 6 random plates.

  - Wrap the plates with labeling tape.

  - Label: RCA batch # and RCA date.

  - Put the plates in the 4oC deli located in Rm 140 in the QC Gels bin.

- Wrap the remaining plates in each batch with labeling tape.

- Label: MB sequencing batch # and date.

- Put the plates in the 4oC deli located in Rm 140 on the correct shelf for that day.

NOTE: Steps 2.16 a-e can be done while running the Cavro during step 3.6. After the plates have been stored in the 4oC deli for 2 weeks, throw them away in the autoclave waste container.

  - For the FW and RV plates:

- Stack FW and RV plates into FW and RV sets.

- Visually check destination plates to make sure that all 384 wells contain RCA template.

Checkpoint: If 4 or more wells per plate are empty or overfilled, re-aliquot the plate manually with a Hydra. If there is something wrong with a single tip (e.g. clogged tip), submit a work request and have it repaired immediately. If the source plate is the cause of the missing wells, see the supervisor to make a determination on how to proceed.

- Place a clear adhesive seal on the top plate of each stack.

- Store the stacks in the 4oC deli located in Rm 140 until it is time to add the cocktail. If the plates are aliquoted one hour or more before the cocktail will be added, add a clear adhesive seal to each plate in the stack before storing the plates in the deli.

  - At the end of the day, place a trough with Coulter Clenz cleaning reagent (located under the sink in room 140) in the destination area and run the Trough Wash program.

- Enter needle fill amount: 80 %. Click OK.

- Enter the number of wash cycles: 5. Click OK.

- Enter the soak time: 60 sec. Click OK.

- Enter the number of bath wash cycles: 5. Click OK.

  - To shut down the Hydra Twister Robot for overnight storage:

- Click on End of the Day Fill to fill the needles with water.

- Turn off the robot by pushing down on the red emergency stop button located next to the Hydra Twister.

- Log off of the computer.

- Sequencing Chemistry

  - Wrap two 250-ml bottles in foil and label FW or RV, and place them on ice.

  - Prepare the ET-Terminator cocktail (Refer to the Reagent/Stock Preparation section) using the two 250-ml bottles (one bottle for FW and one bottle for RV).

NOTE 1: When dispensing on Cavro 1 – make enough cocktail for 3 extra plates of FW and RV so the syringes don’t run dry. When dispensing by hand – make 20% extra cocktail for each.

NOTE 2:  Always make fresh ET-Terminator cocktail, never use recycled cocktail.

  - Cap each FW and RV ET-Terminator cocktail bottle and mix by swirling gently and inverting each bottle 3 to 4 times.

  - Place two 250-ml double pharmaceutical graduates wrapped in foil into two 1000-ml beakers of ice labeled FW and RV.

  - Remove foil from the opening of the double pharmaceutical graduates.  Pour the FW cocktail into the 250-ml double pharmaceutical graduate located in the FW beaker and pour the RV cocktail into the 250-ml pharmaceutical graduate located in the RV beaker.

  - Using Cavro 1, dispense 9μl of the FW and RV ET-Terminator cocktail into each well of corresponding FW and RV 384-well plates:

  - Log onto the computer for Cavro 1.

  - Click on the Cavro 1 Barcodes icon located on the desktop.

  - Check the Cavro 1 calibration for each batch (Refer to the Calibration Check section).

  - Remove the FW Cavro tubing from the FW water bottle.

  - Replace the FW water bottle located on the FW Cavro stand with the FW beaker containing FW cocktail.

  - Place the FW tubing into the FW double pharmaceutical graduate.

  - Make sure the FW stopper is at the bottom of the double pharmaceutical graduate.

  - Repeat steps 3.6.d – 3.6.g for the RV Cavro and RV cocktail.

  - Under Select Stations, select BOTH.

  - Under Select Process, select PRIME.

NOTE 1: Before priming, make sure that the troughs do not have water in them.

NOTE 2: Ensure that the feed tubes are submersed in ET-Terminator cocktail and contain no air bubbles

  - Dump the FW and RV trough contents back into corresponding FW and RV double pharmaceutical graduates of ET-Terminator cocktail.

  - Remove the FW and RV plate stacks from the 4oC deli located in Rm140.

  - Place a FW plate onto its corresponding FW plate stand (bottom Cavro stage) and place a RV plate onto its corresponding RV plate stand (top Cavro stage).

  - Click on START.

  - Upon completion of the dispensing run, take the plates off of the stands and visually check to make sure that cocktail has been dispensed into each well.

Checkpoint: If 4 wells or less per plate are empty, manually add 9μl of cocktail to the empty wells with a pipette. If more than 4 wells are empty or overfilled, re-aliquot the plate with the hydra and run it through the Cavro again. If there is something wrong with a tip (e.g. clogged tip), submit a work request and contact Instrumentation to get the problem resolved.

NOTE: While Cavro is dispensing, the source plates can be sealed.

  - Place the plates on ice.

  - Repeat steps 3.6m – 3.6p until cocktail has been dispensed into all of the plates.

NOTE: Instead of clicking on START in step 3.6n, you may now press ENTER on the keyboard to start a dispensing run.

  - Purge the leftover ET-Terminator cocktail from Cavro 1 after each batch:

  - Under Select Process, select PRIME.

  - Pour the leftover ET-Terminator cocktail remaining in the troughs down the sink.

  - Place plates in a centrifuge for a quick spin (1 minute at 1,000rpm).

  - Use the PlateLoc Thermal Plate Sealer to seal each plate with a white foil seal.

  - In the Venonat database, click on RCA DB ( STEP 4B - Scan Plates for Cycle Sequencing:

- Fill in the boxes: Dispense Operator, Dispense Instrument, FW Primer Lot, FW Kit Lot, RV Primer Lot, RV Kit Lot, Mix Operator, Days.

- Scan in the plates in FW and RV pairs to thermocyclers according to which side of the thermocycler a plate will be loaded.

- Click on SUBMIT.

- Print this page and add it to the sequencing chemistry binder.

  - Place the plates on the correctly matched thermocyclers.

  - Run the ET-Terminator program on the thermocyclers (refer to Instrument Maintenance & Settings for cycle program).

NOTE: Use the computer to start thermocyclers if all thermocyclers are filled on the corresponding rack. Otherwise, start thermocyclers individually.  After starting thermocyclers, verify that each thermocycler has been started on the correct program.

  - Wash Cavro 1 after each batch.

a.   Place the double pharmaceutical graduates in the autoclave bin.

b.   Place the 500-ml water bottles back on their respective FW and RV stands.

c.   Place the Cavro tubing back into the bottles.

d.   Under Select Process, select PRIME and repeat 3 times.

NOTE: Empty the FW and RV troughs into waste containers after each wash

- Rinse out the FW and RV troughs by running them under Milli-Q water.

- Add fresh Milli-Q water to each 500-ml bottle when the water levels get low.

- Place the FW and RV troughs back into their slots

  - After the thermocycler program ends (approx. 2 hours later), remove the paired plates from the thermocyclers and place them in a bin labeled: MB batch # and date.

NOTE: Stop the thermocyclers either manually or by computer. Use another plate or a plate release tool to pry the plates up off the thermocyclers. When the plates start getting difficult to remove, spray the thermocyclers with PTFE Release Agent Dry Lubricant.

  - Place the plates in the 4oC fridge located in Rm141 for the BET cleanup protocol.

  - At the end of the day:

  - On Mondays, clean the Cavro by priming 70% ethanol through the FW & RV tubes 3 times.  Follow the ethanol wash with a water wash by priming Milli-Q water through the tubes 5 times.

  - Replace the two 500-ml water bottles on the FW and RV Cavro stands with clean bottles filled with fresh Milli-Q water.

  - Under Select Process, select PRIME.

  - Leave the Cavro tubing filled with water.

  - Empty & rinse out the FW and RV troughs by running them under Milli-Q water.

  - Close the Cavro 1 Barcodes program.

  - Log off of the computer.

# Reagent/Stock Preparation

DYEnamic ET-Terminator cocktail, 10μl reaction for Plasmid Template

Per one 384-well Plate:

1.92   ml Milli-Q Water

1.536 ml DYEnamic ET-Terminator mix

  - μl  Primer (250μM)

- When preparing a fresh cocktail mix, prepare enough cocktail for the number of source plates in your sample plus 3 additional plates according to the cocktail volume per number of plates chart located in the Sequencing Chemistry bay and also at the end of this protocol in Appendix B.

- Wrap two 250 ml bottles with Aluminum foil and place bottles in an ice bucket.  Mark each bottle with FW or RV for the two different primer directions.

NOTE:  The ET-Terminator is light sensitive.

- Add water, ET-Terminator, and primer to each 250-ml bottle and mix by swirling and inverting the bottle.

NOTE: Keep all reagents on ice while preparing cocktail and dispensing into the 384-well plates.

70% Ethanol

- Pour 350 ml of 100% ethanol into a 500-ml bottle.

- Add 150 ml of Milli-Q water.

- Mix well by inverting bottle.

- Label bottle: 70% ethanol, date prepared, and your initials.

# Instrument Maintenance & Settings

Perkin Elmer 9700/Applied Biosystems 9700

Program ETMB30

95oC for 30 seconds

95oC for 25 seconds

50oC for 10 seconds	   30 cycles

60oC for 2 minutes

Hold at 4oC

Volume = 10μl

Cavro 1

Setup:

  - Log onto the computers located next to the Cavros.

  - Double-click on the Cavro 1 Barcodes e55A icon.

Calibration:

Refer to the Calibration Check section on page 2.

Shut-down:

  - Close the Cavro 1 Barcodes program.

NOTE:  Be sure that the tubes are full of water and that the water bottles are replaced daily with clean glassware and Milli-Q water.

  - Log off the computer.

# SOP Approval

| DEPARTMENT | APPROVED BY | DATE |
| --- | --- | --- |
| Lab Supervisor |  |  |
| Research & Development |  |  |
| Instrumentation |  |  |
| QC |  |  |
| Purchasing |  |  |
| EH & S |  |  |
| Informatics |  |  |
| Seq Assessment & Analysis |  |  |
| Dept Head of Prod Seq |  |  |

# Appendix A

|  |  |  |  | ET cocktail Recipe (MB) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Plate | H2O (mL) | KIT (mL) | Primer (uL) |  | Plate | H2O (ML) | ET Kit (mL) | Primer (UL) |
| 1 | 1.92 | 1.536 | 6.144 |  | 37 | 71.04 | 56.832 | 227.328 |
| 2 | 3.84 | 3.072 | 12.288 |  | 38 | 72.96 | 58.368 | 233.472 |
| 3 | 5.76 | 4.608 | 18.432 |  | 39 | 74.88 | 59.904 | 239.616 |
| 4 | 7.68 | 6.144 | 24.576 |  | 40 | 76.8 | 61.44 | 245.76 |
| 5 | 9.6 | 7.68 | 30.72 |  | 41 | 78.72 | 62.976 | 251.904 |
| 6 | 11.52 | 9.216 | 36.864 |  | 42 | 80.64 | 64.512 | 258.048 |
| 7 | 13.44 | 10.752 | 43.008 |  | 43 | 82.56 | 66.048 | 264.192 |
| 8 | 15.36 | 12.288 | 49.152 |  | 44 | 84.48 | 67.584 | 270.336 |
| 9 | 17.28 | 13.824 | 55.296 |  | 45 | 86.4 | 69.12 | 276.48 |
| 10 | 19.2 | 15.36 | 61.44 |  | 46 | 88.32 | 70.656 | 282.624 |
| 11 | 21.12 | 16.896 | 67.584 |  | 47 | 90.24 | 72.192 | 288.768 |
| 12 | 23.04 | 18.432 | 73.728 |  | 48 | 92.16 | 73.728 | 294.912 |
| 13 | 24.96 | 19.968 | 79.872 |  | 49 | 94.08 | 75.264 | 301.056 |
| 14 | 26.88 | 21.504 | 86.016 |  | 50 | 96 | 76.8 | 307.2 |
| 15 | 28.8 | 23.04 | 92.16 |  | 51 | 97.92 | 78.336 | 313.344 |
| 16 | 30.72 | 24.576 | 98.304 |  | 52 | 99.84 | 79.872 | 319.488 |
| 17 | 32.64 | 26.112 | 104.448 |  | 53 | 101.76 | 81.408 | 325.632 |
| 18 | 34.56 | 27.648 | 110.592 |  | 54 | 103.68 | 82.944 | 331.776 |
| 19 | 36.48 | 29.184 | 116.736 |  | 55 | 105.6 | 84.48 | 337.92 |
| 20 | 38.4 | 30.72 | 122.88 |  | 56 | 107.52 | 86.016 | 344.064 |
| 21 | 40.32 | 32.256 | 129.024 |  | 57 | 109.44 | 87.552 | 350.208 |
| 22 | 42.24 | 33.792 | 135.168 |  | 58 | 111.36 | 89.088 | 356.352 |
| 23 | 44.16 | 35.328 | 141.312 |  | 59 | 113.28 | 90.624 | 362.496 |
| 24 | 46.08 | 36.864 | 147.456 |  | 60 | 115.2 | 92.16 | 368.64 |
| 25 | 48 | 38.4 | 153.6 |  | 61 | 117.12 | 93.696 | 374.784 |
| 26 | 49.92 | 39.936 | 159.744 |  | 62 | 119.04 | 95.232 | 380.928 |
| 27 | 51.84 | 41.472 | 165.888 |  | 63 | 120.96 | 96.768 | 387.072 |
| 28 | 53.76 | 43.008 | 172.032 |  | 64 | 122.88 | 98.304 | 393.216 |
| 29 | 55.68 | 44.544 | 178.176 |  | 65 | 124.8 | 99.84 | 399.36 |
| 30 | 57.6 | 46.08 | 184.32 |  | 66 | 126.72 | 101.376 | 405.504 |
| 31 | 59.52 | 47.616 | 190.464 |  | 67 | 128.64 | 102.912 | 411.648 |
| 32 | 61.44 | 49.152 | 196.608 |  | 68 | 130.56 | 104.448 | 417.792 |
| 33 | 63.36 | 50.688 | 202.752 |  | 69 | 132.48 | 105.984 | 423.936 |
| 34 | 65.28 | 52.224 | 208.896 |  | 70 | 134.4 | 107.52 | 430.08 |
| 35 | 67.2 | 53.76 | 215.04 |  | 71 | 136.32 | 109.056 | 436.224 |
| 36 | 69.12 | 55.296 | 221.184 |  | 72 | 138.24 | 110.592 | 442.368 |

# Appendix B

AUDIT TRACKING

6/07/04 – Steve Wilson and Jamie Jett performed an audit between 2:30PM – 6:00PM. The operator was Megha Ambati.

PROCESS CHANGES

8/10/04 - No longer taking individual well volumes per FW & RV plate per batch per Hydra Twister.

9/02/04 - Cavro 1 calibration range changed from 3.26g-3.65g to 3.30g-3.50g.

10/20/04 – HTS weight ranges have changed.  If weights fall outside 0.3366g – 0.3743g it gets a Warning level, if it falls outside 0.3272g – 0.3837g it gets a do not use level.

10/28/04 – Taking Cavro calibrations on the first FW & RV plate using cocktail instead of water.  Acceptable weight range is now 3.415-3.497.