MD-80045

Method: A-006

Cyanide/CLP SOW ILM03.0

Issue 1

Miamisburg Environmental Management Project

Miamisburg, OH

Source Document:  QAPP (April 1995)

# INTRODUCTION

## Cyanide Analysis Description

Cyanide will be analyzed according to the CLP SOW ILM03.0 for water and soil samples. Soil samples must be prepared per Attachment A. This method uses spectrophotometry. The required detection limits for cyanide are 10g/L for water and 2 mg/kg for soil.

## References

DOE 1995. “Remedial Investigation/Feasibility Study Operable Unit 9, Site-Wide Quality Assurance Project Plan,” Final Revision 4, U.S. Department of Energy, April 1995.

EPA. 1996. “U.S. EPA Contract Laboratory Program, Statement of Work for Inorganic Analysis, Multi-media, Multi-Concentration.” Document No. ILM03.0. U.S. Environmental Protection Agency, March 1990.

# PRESERVATION

Cyanide Analysis
Sample Containers, Volumes, Preservation, and Holding Times

| Matrix | Parameters1 | Analytical Method | Container | Minimum Volume | Preservation | Holding Time |
| --- | --- | --- | --- | --- | --- | --- |
| Water | Cyanide | CLP SOW | Polyethylene bottle | 1500 mL | NaOH to pH  12 Cool 4°C | 14 days |
| Soil | Cyanide | CLP SOW | Wide-mouth polyethylene bottle | 100 grams | Cool 4°C | 14 days |

1	the analyses have been listed to ensure field personnel know which analyses can be taken from the same container.

# CALIBRATION

Cyanide will be analyzed according to the procedures presented in the CLP SOW for inorganic analyses. The spectrophotometer is calibrated daily, or once every 24 hours, and each time the instrument is set up. The minimum correlation coefficient for cyanide calibration is 0.996.

A initial calibration verification (ICV) is performed to assess the accuracy of the initial calibration using a standard of a certified concentration from an external source. When the measurement exceeds the CLP-established control limits, the problem is corrected, the instrument is re-calibrated and the ICV is run again. The initial calibration is verified by analysis of a continuing calibration verification (CCV) standard every two hours during an analysis run or at a frequency of 10%, whichever is more frequent. This standard is also analyzed for at the beginning and end of each sample run. The concentration and source of the CCV and acceptance criteria are specified in the CLP SOW.

# QC CRITERIA

Table 4.1 - Cyanide Analysis
Field QC Sample Frequency

| Parameter | Quality Control Check | Frequency | Acceptance Criteria | Corrective Action |
| --- | --- | --- | --- | --- |
| Cyanide, CLP SOW ILM03.0 | Duplicate | 1 every 10 or fewer field samples (water) 1 every 10 or fewer field samples (soil) | ( 25% RPD Not applicable | Evaluate data for usability. Evaluate variability. |
|  | Equipment (rinsate) blank | 1 every 10 or fewer field samples (water) | ( 10 ( level in associated samples | Evaluate potential sources; Evaluate associated data for usability. |

Table 4.2 - Cyanide Analysis CLP SOW ILM03.0
Laboratory QC Sample Frequency

| Parameter | Quality Control Check | Frequency | Acceptance Criteria | Corrective Action |
| --- | --- | --- | --- | --- |
| Cyanide, CLP SOW ILM03.0 | Initial and continuing calibration blanks (ICB,CCB) | After every ICV and CCV or 10% or every 2 hours, whichever is more frequent. | ( CRDL | Correct problem; recalibrate; reanalyze preceding 10 samples or all since last good blank. |
|  | Spike sample (S) | 1 per group of samples of a given matrix, concentration, or sample delivery group, whichever is more frequent. | 75-125% Recovery | Evaluate data for usability. |
|  | Sample duplicate (D) | 1 per group of samples of a given matrix, concentration, or sample delivery group, whichever is more frequent. | If result ( 5 ( CRDL: ( 20% RPD If result ( 5 ( CRDL: ( CRDL | Evaluate data for usability. |
|  | Laboratory Control Sample (LCS) | One per laboratory batch. | 85-115 | Evaluate data for usability. |

# ANALYTE LIST AND REPORTING LIMITS

These are expected quantitation limits based on reagent grade water or a purified solid matrix. Actual quantitation limits may be higher depending upon the nature of the sample matrix. The limit reported on final laboratory reports will take into account the actual sample volume or weight, percent moisture (where applicable), and the dilution factor, if any.

| Table 5.1 Cyanide - CLP SOW ILM03.0 Target Analyte List |
| --- |
| Analyte | Water (g/L) | Soil (mg/kg) |
| Cyanide | 10 | 2 |

attachment a for
method a-006

Attachment A for Method A-006

Statement of Work for Soil Preparation for Common Organic, Inorganic, and 
Selected Radiological Analyses

1.0	Scope and Application

This procedure describes how to aliquot Mound soil and soil-like samples for laboratory preparation and analysis. This procedure applies to soil analysis for metals, semi-volatiles, pesticide/PCBs, cyanide, anions, explosives, and radiological analyses which do not have a prescribed soil preparation procedure. This procedure should not be used for volatile organic analysis. Soils for volatile organic analysis will be prepared and homogenized as described in the method of analysis.

2.0	Summary

A representative aliquot of a sample is taken in the laboratory by either visually examining and taking a representative portion from each layer in a sample or taking a core of the sample.

3.0	Interferences

Soil samples are heterogeneous by nature. Because of this nature, target analytes are often channeled and concentrated in the soil in specific layers or locations. This heterogeneity may affect both how representative the sample is of the field location and how representative the laboratory aliquot is of the sample.

Heterogeneous nature of soils can sometimes be eliminated in laboratory aliquoting by visually inspecting the sample for layering and selecting a representative aliquot or by taking a core of the sample.

4.0	Equipment

4.1	Spatula or Scoop

4.2	Glass tray, plastic tray, or other material for containing spilled soil

4.3	Large container, i.e. 1000 mL Pyrex beaker

5.0	Reagents/Supplies

5.1	Disposable gloves

6.0	Sample Collection/Holding Time/Preservation

6.1	See Section 2.0 of Method A-006.

7.0	Procedure

7.1	Place a glass tray, plastic tray, or disposable paper beneath the sample container. The tray or paper will be used to contain any soil which accidentally falls off the bottle lip when the cap is opened or falls out while the sample is taken.

7.2	Visually examine the contents of the sample container. If obvious layering is present, then representative portions of each layer must be taken for the aliquot.

- If the sample is obviously a core sample (cylindrical soil mass), then use the spatula to core from the top of the sample to the bottom of the sample. This procedure should be representative of the entire core.

- If the sample cannot be easily cored, it may be necessary to transfer the sample to a large container and thoroughly and carefully mix the sample with a spatula or scoop. Mixing will not be performed on soil samples for volatile and semi-volatile analyses.

- If the sample is neither layered nor a core sample, then use a spatula to core through the middle of the sample. The core should be representative of the entire sample.

7.3	Process the sample as specified in the applicable method.

8.0	Quality Control

8.1	Each analytical method has specific types of quality control samples introduced to evaluate laboratory precision and reproducibility of sample results. Typically, these quality control samples are laboratory duplicates or matrix spike duplicates. These quality control samples permit the laboratory to calculate the relative percent difference and evaluate the soil aliquoting procedure and the precision of the method.

9.0	References and Associated Standard Operating Procedures

None