HELIUM BACKGROUND IN THE D-ZERO DETECTOR

RELATED TO THE

PHOTOMULTIPLIER TUBES

D-ZERO ENGINEERING NOTE # 3823.130-EN-490

Approved :

Kurt Krempetz, PPD/ETT/D0 Upgrade/Mechanical group leader

April 9, 1998

Author:  Russ Rucinski

PPD/ETT/D-Zero Upgrade project

APPENDIX

April 8, 1999

Author:  Jason Lambin

PPD/ETT/D-Zero Upgrade project

Summary

| Source | #### Max. Continuous | Max. Discrete | Effect on Tubes During 10 year run |
| --- | --- | --- | --- |
| ## Solenoid system | 180 ppm Helium [Away from tubes] | 100 % Helium < 1 min. [Away from tubes] | None |
| VLPC system | 270 ppm Helium [Away from tubes] |  | None |
|  | 10 ppm Helium [4 to 6 “B,C” layer tubes] |  | None |
|  |  | 1.3 % He decaying to <100 ppm in 5 min. [By 10 to 20 “B” layer tubes] | None Lifetime at 1.3% is 45 days. Expect 1.3% for total 80 min. over 10 year run |
| Calorimeter cold valve actuation gas | Negligible probability | 7 % He decaying to <100 ppm in 30 minutes [By 2 “A” layer tubes] | None Lifetime at 7% is 8 days. Expect to actuate cold valve 2 to 4 times over 10 yr run. |
| Tev Magnet | 100 to 200 ppm Helium [Away from tubes] | 5 % He decaying to <100 ppm in 25 minutes. [Magnet rupture, Envelopes D0 detector] | None- Lifetime at 5%: Borosilicate tubes, 11 days Quartz tubes, 167 minutes Expect maybe 1 discrete event during entire run. |

Discussion and Background:

Helium is present in the earth’s atmosphere at about 5 parts per million.  (ref. Technology of liquid helium, NBS monograph 111).  The D-Zero detector uses helium for the cryogenic cooling of its superconducting magnet and visible light photon counter (VLPC) electronics chips.  In addition, the tevatron accelerator has superconducting magnets that use helium.  Due to the possibility of leaks or releases of helium from these helium lines and components, the background helium level in the collision hall may exceed the natural level of 5 ppm.  This engineering note will quantify the probability and level of helium background in the D-Zero detector.

The photomultiplier tubes used in the D-Zero detector are sensitive to an elevated helium atmosphere.  This is due to the permeation rate of helium gas through the glass tube, into the vacuum space inside.  It is very important for the helium atmosphere surrounding the photomultiplier tubes is known and controlled.  If the level of helium in the vacuum tube reaches a level above 5 ppm, then the photomuliplier tube may no longer work as designed.  The process is an irreversible one.

### Photomultiplier tube susceptibility

The time for the helium level in the photo tube to reach a certain level by diffusion is given by:

Time =-tau* (1-Ptube/Psurroundings)	where: tau = 1010 seconds for Borosilicate glass

= 109  seconds for UV glass

= 108   seconds for Quartz glass

This information is shown in the Table 1

Table 1.  Time to reach 5 ppm for a sustained exposure

|  |  | ##### Exposure Time |  |
| --- | --- | --- | --- |
| Sustained Helium Exposure | Quartz | UV glass | Borosilicate |
| 5 ppm | infinite | infinite | infinite |
| 10 ppm | 2.2 years | 22 years | 200 years |
| 100 ppm | 59 days | 1.6 years | 16 years |
| 0.1% | 6 days | 58 days | 1.6 years |
| 1 % | 14 hours | 6 days | 58 days |
| 10 % | 83 minutes | 14 hours | 6 days |
| 100 % | 8.3 minutes | 83 minutes | 14 hours |

### Quartz photomultiplier tubes are going to be used for the luminosity monitors.  They will be located between the central calorimeter and end calorimeters surrounding the beam pipe.  Note that their location is at least 5 meters away from any helium source.  From Table 1 it is seen that the quartz photomultiplier tubes can tolerate a high level of concentration for a short time.  A sustained low level concentration of only 10 ppm however, could cause a premature death of the tubes before the 5 to 7 year physics run is complete.

### The muon trigger counters use Borosilicate photomultiplier tubes.  These tubes are more widely distributed through out the D0 detector.  They are mounted on the “A”, “B”, and “C” layers of the proportional drift tube chambers.  These tubes are more hardy and can tolerate a reasonable elevated helium background.

### Helium dispersion behavior

A helium leak disperses into the surrounding air by diffusion and rises upward due to its lesser density.  To better understand qualitatively what it happens, I can calculate some dimensionless parameters.  First, I assume the case of a discrete release of helium that mixes evenly with the surrounding air into a 1 meter diameter sphere.  Assume the helium/air mix is 5 % helium, 95 % air.

The Grashoff number (Gr)2 is the ratio of the buoyancy forces to viscous forces for the gas.  It typically is calculated to understand free convection heat transfer.  A warm or cold surface induces flow by changing the fluid density.

Gr=g*Beta*(Ts-Tinf)*L^3/mu^2 	where:  g= acc. Due to gravity, Beta= fluid expansion coeff.,

Ts=surface temp., Tinf=temp. far away,

L=characteristic length, mu= viscosity

Beta = (1/RHO)*(RHOinf-RHO)/(Tinf-T)  where: RHOinf=fluid density far away,

RHO=fluid density of air/helium mix,

T=Avg.,Fluid temp = (Ts+Tinf)/2

Substituting the second equation into the first, the temperatures cancel out.  By plugging in all the values, and using 1 meter for the characteristic length, Gr= 4000.  The buoyancy force is huge compared to the viscous drag type forces on the fluid mix.

The Schmidt number (Sc) provides a measure of the relative effectiveness of momentum and mass transport by diffusion in the velocity and concentration boundary layers.  It is the ratio of viscous forces and mass diffusivities.

Sc= mu/D	where: mu= viscosity, D = binary diffusion coefficient

The Binary diffusion coefficient for helium in air is not given in reference 2.  They are given for other gasses in air, including hydrogen, which is probably similar to helium in diffusion.  All the binary diffusion coefficients are about the same order of magnitude, and guessing the value to be the same as hydrogen, Sc=0.5.  The viscous transport and diffusion transport are the same order of magnitude.

The ratio of the buoyancy force to mass diffusivities can obtained by multiplying the Grasshoff number by the Schmidt number.  The result is Gr*Sc= 2000.   This shows that the buoyancy force dominates the way the helium moves away from the release sight.  The buoyant transport is about a 2000 times larger influence than the diffusion transport of the helium.  The helium will go straight up and make its way out of the collision hall.

One can also calculate that the 1 meter sphere of 5 % helium and 95 % air will have a net buoyancy force of 0.26 Newtons.  Applying a=F/M, we find that the acceleration on the helium sphere is 0.48 m/s^2.  It takes about 2 seconds for the sphere to move upward 1 meter.

These calculations give one the intuitive sense that the helium will rapidly rise vertically before it disperses outward horizontally.  Trapped helium in inverted pockets could possibly occur, but shouldn’t be a problem for the phototube installations.

# Sources of helium in the D-Zero detector:

- Liquid helium supply/returns for solenoid

- Lead flow returns from solenoid

- Relief valves from solenoid control dewar

- Liquid helium supply/returns for VLPC detector

- Cassette space gas helium volume

- Calorimeter cold valve actuation gas

- Tevatron magnets and tunnel

Each of the above sources will be elaborated on below.  Specifically, I will:  a.) speculate method of the release, b.) describe the location, c.) quantify the potential and amount of a helium release, d.) calculate the helium concentration and duration of  such a release.  Quite a few assumptions are made to arrive at hard numbers.  The volume around the leak is assumed to be ventilated at the rate of ¼ volume change per minute.  This is the approximately the same ratio as the collision hall ventilation rate.  Fans of 23,700 cfm (AHU-1 gives 19,200 cfm, EF-7 gives 4500 cfm) move air through the 100,000 cubic foot D0 detector.  The volume change is assumed to be with fresh air with no elevated concentration of helium.  This assumption is valid because the huge volume of the collision hall dilutes virtually all the helium leaks to a negligible level.

- Liquid helium supply/returns for solenoid:  A reasonable method of a release from theses lines would be through a failed o-ring leak at the u-tube bayonet fitting, a leaking  trapped volume relief, a leaking tube fitting from a pressure tap to transmitter line, valve packing leak etc.  All these potential types of leaks occur at the control dewar location positioned outside the detector at an elevation about 12.5 feet above the beam line.  There are no photo tubes mounted above the location.  The nearest photo tubes are 2 to 3 feet to the north and west of the location.   The likelihood of such a leak is say 25 items * (3E-6 failures/hr) (8766 hrs/yr) = 0.7 leaks per year.  The total leak rate would be a maximum of 1 cc/sec air = 0.005 scfm helium.  (Note that a leak detectable by snoop is 1/1000th of this = 0.001 atm-cc/sec air).  For the sake of a number, assume the helium is distributed through a spherical volume with 3 foot radius.  The resultant helium background would be around 180 ppm.  The concentration level should reduce as cubic function of the distance from the source.  The duration of such a release could be for the whole run since the leak may not be noticed or discovered.  Even though the probability is that one such leak would occur every year  and a half, the total leak rate assumed is rather large and should account for any combination of multiple leaks that may occur during the run.

180 ppm Helium [Away from tubes]

- Lead flow returns from solenoid:   The method and location of a leak from these items is similar to the solenoid supply lines above.  The elevation of such a leak is a few feet higher, within 5 feet of  the top of the detector.  There are going to be 6 potential leak sites all enclosed in the lead flow purge box.  The GN2 purge box would dilute any such leak by about 0.5 scfm purge/0.005 scfm helium = 100 before it’s release out of the box.  The elevated helium concentrations are 100 times less than those calculated for source number 1.  That is (5 ppm (natural background) + 6 ppm=) 11 ppm at 2 feet away and (5 ppm (natural background) + 2 ppm=) 7 ppm at 3 feet away.

Negligible ppm [Away from tubes]

- Relief valves from solenoid control dewar: The solenoid has relief valves that protect the process lines and helium subcooler vessel.  The cause of a relief discharge could be a quench or an operational upset.  The location of the discharge and elevated helium atmosphere is the control dewar location described in the previous cases.  It’s hard to estimate how often a release could happen, in the operational test of the magnet at Toshiba the relief valves did not come close to having to open from a forced quench of the magnet.  For a rough guess, say it might occur 4 times a year.  The rate of release is large 616 cfm helium, for a short duration, say 5 seconds at most, for a release of about 50 scf helium (equivalent to 2 liquid liters of helium).  The local concentration would initially be high, approximately 100 % up to 3 feet from the relief valve.  Modeling the helium  as a sphere, the sphere would rise above and away from the detector in a matter of seconds.  Any such releases from the relief valves would be discrete short term exposures of helium to the tubes.

100 % Helium < 1 min. [Away from tubes]

- Liquid helium supply/returns for VLPC detector:  These lines connect via u-tubes in the same area as the solenoid control dewar.  There are 6 bayonets and 6 trapped volume relief valves at that location. The conservative failure rate of the bayonets and relief valves is about 0.3 failures per year.  A leak rate of 0.005 scfm helium is reasonable.  The duration of such a leak would probably be for the duration of the run.  The maximum local concentration at the end of 5 years above the bayonet feed can could be (5 yrs)(0.3/yr)(180 ppm)= 270 ppm.  This location is not in the area of any photo tubes.

270 ppm Helium [Away from tubes]

The other leak sources are the stems of two control valves located in a valve box in the center-east section of the detector platform.  These stems are bellow sealed and each have a failure rate of about (1E-8/hr)*(8677 hrs/yr)=1.7E-4 per year.  The valve box location is probably within 10 feet of a “B” or “C” layer photo tube.  The failure rate is low enough that it can be neglected.  If it did fail, the mixing and distances involved would probably cause a maximum helium concentration of about 80 ppm at the closest photo tube location.  The helium concentration would probably dilute by the distance from the source cubed.  So photo tubes 20 feet away would see at most 10 ppm above the normal atmospheric concentration.

10 ppm Helium [4 to 6 “B,C” layer tubes]

- Cassette space gas helium volume:  The cassette space volume is located in the center of the detector platform.  The cassette space is a stagnant volume that surrounds the cold end of each of the 102 VLPC cassettes.  Each VLPC cassette has a gasket and silicone potting at the warm optical fiber/flex circuit bulkhead.  I’ll estimate the leak rate for a cassette bulkhead to be 1E-2 atm-cc/sec helium= 2E-5 cfm helium. Multiplied by the 100 cassettes, then the total steady state leak rate is going to be 0.002 cfm distributed over the 1.5 ft x 12 ft long surface area of the cryostats.  The center section of the platform is ventilated with fans at a rate of 2080 cfm and has a volume of about 1600 cf.  The raised helium concentration from the cassette leaks is on the order of 1 or 2 ppm.

Negligible ppm [By 10 to 20 “B” layer tubes]

There are going to be low pressure relief valves protecting the cassette space.  These will relieve pressure generated from a warm up of the cassette space.  These relief valves are going to have their outlets piped into the argon spill duct which is located not more than 10 feet away.  The spill duct has continuous ventilation out of the detector area.

A cassette change operation would also release helium gas.  A cassette change would occur 2 times per year at most.  Operational spares are included as a subset of the total 102 cassette installation.  During a change out one probably would keep a positive helium supply to the cassette space to prevent back streaming of air and moisture into the cassette space which could be cold.  In between the actual removal of the cassette and blanking off the hole again, helium would escape out of the cassette slot for 1 to 2 minutes.  I estimate 30 scf of helium is released during this operation.  The elevated helium background in the center section of the platform could be as high as 1.25% =12,500 ppm helium for a couple of minutes.  The concentration would then decrease exponentially thereafter.

1.3% He decaying to <100 ppm in 5 min. [By 10 to 20 “B” layer tubes]

- Calorimeter cold valve actuation gas:  The cold valves are fed by ¼” tubing lines that run from the cryo corner area (SE corner of platform) up the cryobridge and to the tops of the three calorimeters.  The tubing for the most part is continuous with some joints.  There are muon chambers and counters within a few feet of the tubing run.  The probability of a leak is around (10 joints)(1E-8/hr)(8766 hr/yr)=0.00087 per year or about 1 in a 1000 odds.  The maximum leak that could reasonably occur is around 0.005 scfm.  This leak could cause an elevated helium atmosphere of 500 ppm at the closest photo tube should a leak occur.  Another release is due to the actuation of the cold valve.  The solenoid that applies helium to the cold valve releases a slight burp when energized.  The manufacturer says there is a slight leakage, so I’ll interpret that to be 1% of the full flow Cv=0.09 (orifice of 1/16”).  The equivalent leak rate is 0.9 cfm.  This is a momentary 5 second puff that could cause a local atmosphere of 7% helium within a foot of the solenoid.  This would rapidly dilute away due to the density difference between the helium and air regardless of whether there was a trapped pocket above the location or not.

7% He decaying to <100 ppm in 30 minutes [By 2 “A” layer tubes]

- Tevatron magnets and the tunnel:  This is a hard source to quantify.  The components are located on the north and south  ends of the collision hall and would impact  the outer most layer of counters on the north and south faces of the detector.  The magnets and tunnel probably would have a steady state leak rate and probability not too different from the solenoid control dewar components, sources 1 through 4.  The ODH analysis (conservative) lists a maximum leak rate from a powered low beta quad magnet at 1898 scfm.  The probability of such a failure is 2E-6/hr = 1 in 60 chance per year (also conservative).  The circulating ventilation rate (non fresh make-up) for the collision hall is 24,000 cfm and the volume of the collision hall is about 100,000 cf.  Just to get a number, assume the failure occurs and the release lasts for 3 minutes.  In that time, about 5700 cf of helium would be released into 100,000 cf of air.  The helium concentration could be as high as 5 to 7% initially.  The gas helium would escape out the top of the collision hall, and the large ODH ventilation fans would be kicked on mixing and diluting the concentration from there.  The ventilation rate during such an event would be at least ¼ volume change per minute.

5% He decaying to <100 ppm in 25 minutes. [Magnet rupture, Envelopes D0 detector]

E-mail communication from Rich Partridge, Brown University, referenced from Hamamatsu photomultiplier data book.  Tau is representative, not exact as D0 is not using Hamamatsu photo tubes.

2 Incropera and Dewitt, Fundamentals of heat and mass transfer, 2nd ed., 1985.

Incropera and Dewitt, Fundamentals of heat and mass transfer, 2nd ed., 1985.