ATDEPARTMENT OF HEALTH AND HUMAN SERVICES

FOOD AND DRUG ADMINISTRATION

CENTER FOR DEVICES AND RADIOLOGIC HEALTH

GENERAL AND PLASTIC SURGERY DEVICES PANEL

OF THE MEDICAL DEVICES ADVISORY COMMITTEE

Tuesday, October 14, 2003

8:00 a.m.

Grand Ballroom

Gaithersburg Marriott Hotel

9751 Washingtonian Boulevard

Gaithersburg, Maryland

PARTICIPANTS

Thomas V. Whalen, M.D., Acting Chair

David Krause, Ph.D., Executive Secretary

VOTING MEMBERS

Brent A. Blumenstein, Ph.D.

Phyllis Chang, M.D.

Michael A. Choti, M.D.

Ann Marilyn Leitch, M.D.

Michael J. Miller, M.D.

Amy E. Newberger, M.D.

CONSUMER REPRESENTATIVE

Alisa M. Gilbert

INDUSTRY REPRESENTATIVE

Debera M. Brown

TEMPORARY VOTING MEMBERS

Benjamin O. Anderson, M.D.

Dennis W. Boulware, M.D.

Emily F. Conant, M.D.

Nancy A. Dubler, LLB

Ruth A. Lawrence, M.D.

Stephen Li, Ph.D.

Ellice S. Lieberman, M.D., DR PH

Barbara R. Manno, Ph.D.

Mary H. McGrath, M.D., MPH

Michael J. Olding, M.D.

FDA

Celia Witten, Ph.D., M.D.

C O N T E N T S

PAGE

Conflict of Interest and Opening Remarks:

David Krause, Ph.D.	7

Panel Introductions:

Thomas V. Whalen, M.D.	10

Welcome Remarks:

CDR Stephen P. Rhodes	16

Dr. David Feigal	18

Open Public Comment

Introductory Remarks:  Thomas V. Whalen, M.D.	20

Public Speakers

Susan Scherr, National Coaltion

for Cancer Survivorship	23

Shery Henderson, Silicone Solutions Outreach	28

Carolyn Kittle	31

Elizabeth Anderson	34

Arlene Nicole Cummings	37

Michelle Colombo	40

Dr. Thomas Joiner	42

Susan Cunningham	45

Anne Stansell	49

Rosella Rust	51

Kathy Pate	54

Nancy Bruning	57

Jill Fallows	63

Elizabeth Santoro	65

Evon Peterson	69

Elizabeth Weber	72

Ed Brent	73

Dr. Bruce Cunningham	76

C O N T E N T S (Continued)

PAGE

Open Public Comment (Continued)

Margaretha McGrail	78

Gail Hamilton	80

Audrey Moritz-Ciancutti	84

Cynthia Teague	87

Pamela Dowd	89

Lisa Bordelon	92

Jessica Forman	94

Diane Fjelstad	98

Dr. Frank Vasay	101

Marcy Gross	104

Dr. Marga Massey	109

Carolyn Wolf	113

Dr. Vicki Hofnagel, (by conference call)	117

Virginia Gallagher, (by conference call)	121

Dr. Mark Jewel      	124

Mary Ann Richards	127

Lisa Bancarz	129

Eileen Wright	132

Lisa Lowenstein	134

Kathleen V.F. Nye	138

Betty Buikema	140

Dr. Caroline Glicksman	143

Dr. Edward Melmed	147

Dr. Paul H. Wooley	152

Nancy Gertner	157

C O N T E N T S (Continued)

PAGE

Open Public Comment (Continued)

Dr. Eugene Goldberg	160

Mindy Tapscott	164

Kim Gandy, National Association for Women 	168

Mary Ann Ward	175

Marlene Keeling, Chemically Associated

Neurological Disorders	180

Vanessa Rose Ferelli	183

Connie Wasserman	190

Dr. Sidney Wolfe, Public Citizen	193

Judy Norsigian, Our Bodies, Ourselves	203

Dr. Deborah Bash	208

Kathy Keithley Johnson, Toxic Discovery	212

Dr. Paul Weiss, National Endowment

for Plastic Surgery	218

Dr. Laurie Casas, American Society

for Aesthetic Plastic Surgery 	223

Dr. Leroy Young, Aesthetic Surgery

and Research Foundation	229

Dr. Martha Burke, National Council

of Women's Organizations	233

Dr. Laurie Young, Older Women's League	237

Christine Lozier	241

Anne S. Kasper, Finding My Way	244

Dr. Shawna Willey, American College

of Surgeons	249

Margaret Volpe, Y-Me	252

Regulatory Overview of Silicone Gel-Filled

Breast Implants:

Cilia M. Witten, Ph.D., M.D.	257

C O N T E N T S (Continued)

PAGE

Applicant Presentation

Inamed Corporation

McGhan Silicone-Filled Breast Implants

Introduction:

Ronald J. Ehmsen, Sc.D.	264

Preclinical Studies Review:

Thomas E. Powell	267

Clinical Studies:

JoAnn M. Kuhne, MSN, RAC	295

FDA Presentation

ODE and Office of Surveillance and Biometrics (OSB)

Introduction and Mechanical Testing:

CDR Samie Allen	393

Chemical Testing:

Sam Arepalli, Ph.D.	399

Toxicology Testing:

David B. Berkowitz, Ph.D., VMD	405

Clinical Data Overview:

Sahar M. Dawisha, M.D.	413

Statistical Testing:

Telba Z. Irony, Ph.D.	439

Medical Devices Surveillance and

Literature Review:

S. Lori Brown, Ph.D., MPH	448

Postapproval Study and Labeling:

CDR Samie Allen	465

Open Panel Discussion	489

Open Panel Discussion and FDA Questions	551

P R O C E E D I N G S

Conflict of Interest and Opening Remarks

DR. KRAUSE:  Good morning, everybody.  We are ready to begin this, the 63rd meeting of the General and Plastic Surgery Devices Panel.  My name is David Krause.  I am the Executive Secretary of this panel and I am also a biologist and a reviewer in the Plastic and Reconstructive Surgery Devices Branch.

I would like to remind everyone that we request that you sign in on the attendance sheets which are just outside the door.  You can also pick up an agenda, a panel-member roster and information about today's meeting out there on the table.  The information includes how to find out about future meetings, future-meeting dates, through the advisory panel phone line and how to obtain meeting minutes or transcripts.

Before I turn the meeting over to Dr. Whalen, I am required to read two statements into the record.  One is the deputization to temporary voting status and the other is the conflict-of-interest statement.

"Pursuant to the authority granted under the Medical Devices Advisory Committee Charter, dated October 27, 1990, and as amended August 18, 1999 I appoint Benjamin Anderson, Dennis Boulware, Emily Conant, Nancy Dubler, Ruth Lawrence, Stephen Li, Ellice Liberman, Barbara Manno, Mary McGrath and Michael Olding as deputized voting members of the General and Plastic Surgery Devices Panel for this meeting on October 14 and 15, 2003.

"In addition, I appoint Thomas Whalen to act as temporary chair for the duration of this meeting.  For the record, these individuals are special government employees and consultants to this panel or other panels under the Medical Devices Advisory Committee.  They have undergone the customary conflict-of-interest review and have reviewed the material to be considered at this meeting."

This is signed by Dr. David Feigal who is the Director of the Center for Devices and Radiological Health.

The conflict-of-interest statement is as follows.  "The following announcement addresses conflict-of-interest issues associated with this meeting and is made a part of the record to preclude even the appearance of an impropriety.  To determine of any conflict existed, the agency reviewed the submitted agenda for this meeting and all financial interests reported by the committee participants.

"The conflict-of-interest statutes prohibit special government employees from participating in matters that could affect their or their employers' financial interests.  However, the agency has determined that the participation of certain members and consultants, the need for whose services outweighs the potential conflict of interest involved is in the best interest of the government.

"We would like to note for the record that the agency took into consideration certain matters regarding Dr. Miller.  He reported his institution's past and current involvement with firms at issue.  The agency has determined, therefore, that he may participate fully in the panel's deliberations.

"In the event that the discussions involve any other products or firms not already on the agenda for which an FDA participant has a financial interest, the participants should excuse him- or herself from such involvement and the exclusion will be noted for the record.

"With respect to all other participants, we ask, in the interest of fairness, that all persons making statement or presentations disclose any current or previous financial involvement with any firm whose products they may wish to comment upon."

At this time, I would like to turn the meeting over to Dr. Whalen.

Panel Introductions

DR. WHALEN:  Thank you, Dr. Krause.

Good morning.  I am Dr. Thomas Whalen.  I am Professor of Surgery and Pediatrics at Robert Wood Johnson Medical School.  I would venture, based on my own experience this morning, there are others who are going to be at this meeting who are probably visibly aging waiting for the elevators to get down here because it takes so long, including some of our panel members.  So we would hopefully tolerate some of them coming in a little bit later than we had anticipated.

Today, the panel will be making recommendations to the Food and Drug Administration on a premarket approval application.   Our next item of business is to introduce the panel members who are giving of their time to help the FDA in these matters and the FDA staff who are here at this head table.

I would ask each person, starting at my left, to introduce themselves and for the panel members to specifically not only give the typical introductions with your names, institutions and titles, but to please give a very brief comment upon the area of expertise that you bring to this particular panel.

Starting to my left, then, from the FDA with Dr. Witten.

DR. WITTEN:  I am Dr. Celia Witten, the Division Director at FDA of the Reviewing Division for these products.

DR. MILLER:  I am Michael Miller.  I am a professor of plastic surgery at the University of Texas M.D. Anderson Cancer Center and I have worked with breast reconstruction and breast surgery and these types of devices for many years.

DR. ANDERSON:  My name is Ben Anderson.  I am an associate professor of surgery at the University of Washington.  I am a breast-cancer surgeon.  I do not place implants but my patients do receive them and I am listening on their behalf.

DR. LI:  My name is Dr. Steve Li.  I am the President of Medical Device Testing Innovations.  My areas of expertise are biomaterials, biomechanics and failure of medical devices.

PROF. DUBLER:  I am Nancy Dubler.  I am Director of the Division of Bioethics at Montefiore Medical Center and a professor of epidemiology and population health at the Albert Einstein College of Medicine.  My expertise is in medical ethics.

DR. NEWBERGER:  I am Amy Newberger.  I am a dermatologist in private practice in Scarsdale, New York.  I teach at St. Luke's Roosevelt Medical Hospital Consortium in New York City.  My expertise is that I have along experience with various forms of silicone.

DR. BOULWARE:  My name is Dennis Boulware.  I am a professor of medicine at the University of Alabama at Birmingham.  I am also a rheumatologist.

DR. McGRATH:  I am Mary McGrath.  I am a professor or surgery at the University of California, San Francisco, and I am a practicing plastic surgeon, have been working with implants for almost 25 years.  My research interests are in wound healing and biomaterials.

DR. LEITCH:  I am Marilyn Leitch.  I am a professor of surgery at the University of Texas Southwestern Medical Center in Dallas.  I am a surgical oncologist so I deal with patients having breast cancer and also breast-screening issues.  That is the expertise I bring.

DR. CHANG:  I am Phyllis Chang.  I am an associate professor at the University of Iowa College of Medicine, Department of the Surgery in the Division of Plastic Surgery as well as in the Division of Hand and Microsurgery.  My expertise is as a reconstructive plastic surgeon.

DR. CHOTI:  My name is Michael Choti.  I am an associate professor of surgical oncology at Johns Hopkins University and my interest is in clinical surgical oncology.

DR. BLUMENSTEIN:  I am Brent Blumenstein, a biostatistician working independently.

DR. CONANT:  I am Emily Conant.  I am an associate professor of radiology and Chief of Breast Imaging at the University of Pennsylvania.  My expertise is in detecting and diagnosing breast cancer with breast imaging.

DR. LAWRENCE:  I am Ruth Lawrence from the University of Rochester School of Medicine and Dentistry, Professor of Pediatrics and OB-GYN.  My particular area of interest and expertise is breast feeding and human lactation, its impact on both the mother and the baby, and the breast and its function and purpose.

DR. LIEBERMAN:  My name is Ellice Lieberman.  I am an associate professor of obstetrics and gynecology at Harvard Medical School and an associate professor of maternal and child health at the Harvard School of Public Health.  I am also Director at Brigham and Women's Hospital of the Center for Perinatal Research and my expertise is in epidemiology.

DR. MANNO:  I am Barbara Manno.  I am from the Louisiana State University Medical Center in Shrevepoint, Louisiana.  I am professor in the Department of Psychiatry.  My specialty is toxicology.  I have been trained as a general toxicologist.  I have worked in poison-center activities from management down to answering the phones.  So I come to you with quite a few years in medical-toxicology support and my subspecialty is forensic toxicology.

DR. OLDING:  My name is Michael Olding.  I am an associate professor of surgery at George Washington University in the Division of Plastic Surgery.  I have had a long interest in plastic-surgery devices in general and, in particular, breast reconstruction.  I bring that experience to this panel.

MS. BROWN:  My name is Debera Brown.  I am the Vice President of Regulatory Affairs for Broncus Technologies.  I am the industry representative for the panel.  I have worked in medical-device development for over twenty years.

MS. GILBERT:  Good morning.  My name is Alisa Gilbert.  I am the co-founder of the Unbroken Circle and Director of Office of Native Cancer Survivorship and the consumer representative today.  I am also a breast-cancer survivor of seven years and I have intimate experience with implants and this whole process.  So I am delighted to be here.

DR. WHALEN:  I would like to thank all the panel members for their introductions and also thank them for the time they are giving these two days to help out the FDA.

Also, because of the panel configuration going from Delaware to Virginia, you may need to not just raise your hand but speak into the microphone if you need get my attention at any time to ask questions or make comments, please.

I would like to note for the record that the voting members present to constitute a quorum as required by 21 CFR Part 14.

Next, I would like to introduce Commander Stephen Rhodes, the Branch Chief of the Plastic and Reconstructive Surgery Devices Branch who will make some introductory remarks.  Commander Rhodes?

Welcome Remarks

CDR RHODES:  Thank you, Dr. Whalen, and good morning.  I am Stephen Rhodes, the Branch Chief of the Plastic and Reconstructive Surgery Devices Branch here at the FDA.  Welcome members of the panel, members of the public and manufacturers to this important two-day meeting of the General and Plastic Surgery Devices Panel, important  because it is the first FDA panel meeting to make recommendations on the approvability of a silicone gel-filled breast implant since 1992.  Tomorrow, you will make recommendations and vote on Inamed Corporation's premarket approval application.

Because of the high public interest in silicone-gel-filled breast implants, in addition to the regular public-comment periods for any PMA discussion, we have scheduled eight hours for public comment on general issues related to silicone-gel-filled breast implants.

After the public comment period this morning, Dr. Celia Witten, Director of the Division of General Restorative and Neurological Devices, will review the history of FDA's regulation of silicone-gel-filled breast implants.  Following Dr. Witten's presentation, Inamed Corporation will present the data in their premarket approval application followed by FDA's review of the application.

While this afternoon's FDA presenters are in the Office of Device Evaluation and the Office of Surveillance and Biometrics, there are other offices in the Center for Devices and Radiological Health that are part of the FDA team and many are here today.  These include: the Office of Compliance who have evaluated the sponsor's manufacturing facility and audited the clinical-study data; the Office of Health Industry Programs who provide assistance to consumers and review the program labeling; the Office of Science; and Technology and the Office of the Director.

Additionally, the FDA team consists of reviewers from the Division of Pediatric Drug Development and the Pregnancy Labeling team in the Office of New Drugs in the Center for Drug Evaluation and Research.  I also want to acknowledge the valuable assistance provided by the Office of Women's Health in FDA Headquarters.

As a reminder, we will not be discussing saline-filled breast implants at this meeting and I request that panel members and the members of the public limit their comments and discussion to the products being evaluated, silicone-gel-filled breast implants.

Panel members, we appreciate your commitment.  The representatives of 27 professional organizations and 92 members of the public who have requested time to address the panel, we appreciate your comments.  And, PMA sponsor, we appreciate your participation in presenting the safety and effectiveness data you have collected to the panel and answering questions that the panel may have.

Thank you for your attention.  I would now like to introduce Dr. David Feigal, the Director for the Center for Devices and Radiological Health.

David?

DR. FEIGAL:  Good morning.  I would also like to extend my welcome here.  I wanted to take a moment to actually address a problem which we are trying to deal with.  We have heard for a number of years that it is very difficult for people who have medical devices to get notifications and information about these devices.

It has been suggested registries could be used for such a mechanism but registries track individuals by name, must keep track of their addresses.  They are very labor intensive and beyond, really, the resources of the Center aside from any of the confidentiality issues and problems that that creates.

What we have done in breast implants is a new program which we hope to extend to other medical devices where we will actually provide, through e-mail contact, information about products of great concern to the public.  There is a brochure that is out on the table that people can sign up and all that is needed is an e-mail name.  If you would like to also give us your name, that is fine, but we are not in any way tracking individuals.  This will be an opportunity where we can communicate with people who would like information about medical devices, about medical devices that they are using, about such products.

So, for example, the summary of this meeting, the web pages where the transcripts may be found, other kinds of information about these implants will be provided as a part of the updates that would be sent by e-mail.  If there are organizations that would like us to load mailing lists and other lists of people who would like to participate in this and hear information from us, we can do that as well.

One of the things that we particularly appreciate is the publicness of this process.  We realize that it is at some expense that people travel here.  It is certainly a commitment of people's time.  We value very much the perspectives that we hear from people who will put a face on the use of these products and we value, very much, your expertise in considering whether or not these products meet the standards for safe and effective products.

So, with that, let me conclude my welcome and wish you a successful two days.

DR. WHALEN:  Thank you.  We now will proceed with the open public hearing session of this meeting.

Open Public Comment

Introductory Remarks

DR. WHALEN:  All those who will be addressing the panel are asked to speak clearly into the microphone at the podium as the transcriptionist is dependent upon this means of providing an accurate record of the meeting.

I would specifically now like to have the attention of all those who are registered or who might be considering speaking to the panel today.  You have each, if you have registered, been given a number corresponding to your order of appearance.  We would ask that you come to the podium area in advance so that we are not spending a great deal of time in transitions from speaker to speaker due to the volume of individuals who will be speaking.

FDA staff will direct you to the appropriate podium.  Please remain within your time constraints as we will have a timer going to help people remain on time and please be aware that you will receive a flashing light when 30 seconds remain in your allocated time.  If you go beyond that, I will try to, as unobtrusively as possible, remind you that you need to come to conclusion.

Both the FDA and the public believe in a transparent process for information gathering and decision making.  To ensure such transparency at the open public hearing session of the advisory committee meeting, FDA believes that it is important to understand the context of an individual's presentation.

For this reason, FDA encourages you, the open public hearing speaker, at the beginning of your written oral statement, to advise the committee of any financial relationship that you may have with the sponsor, its product and, if known, its direct competitors.

For example, this financial information might include the sponsor's payment of your travel, lodging or other expenses in connection with your attendance at the meeting.  Likewise, FDA encourages but does not require you at the beginning of your statement to advise the committee if you do not have any such financial relationships.

If you choose not to address this issue of financial relationships at the beginning of your statement, it will not preclude you from speaking, which is a departure from the last time I chaired a meeting a couple of years ago.

We will start with those individuals who have notified FDA of their request to present in the open session.  I would tell you in advance, for those of you who will be lining up in your assigned order, that, at a particular point in time in mid-morning, I will be interrupting to take a break and then I also will be interrupting at a later point in early afternoon when we will be breaking for 45 minutes for lunch.

If the first speaker would come to the podium, then, please.

Public Speakers

MS. SCHERR:  Good morning.  My name is Susan Scherr and I am with the National Coalition for Cancer Survivorship.  There is no conflict of interest.  I have not received any funds.  This is a statement letter on behalf of twelve different organizations.

To the members of the panel, the National Alliance of Breast Cancer Organizations, Y-Me National Breast Cancer Organization, the National Coalition for Cancer Survivorship and nine non-profit health and cancer organizations that share our concerns believe it is important that women requiring reconstructive surgery for because have silicone breast implants available as an option for reconstruction.

Please note that we offer no opinion or comment on the availability and use of silicone breast implants for cosmetic or other purposes.  Consistent with the missions and expertise of our organizations, we are not endorsing products from any manufacturer and the following comments are limited to the availability of silicone breast implants for women with breast cancer, breast disease or other medical needs and by qualified clinicians.

All of the undersigned organizations support the right of each consumer and patient to consider options for selection and then receive individualized treatment and care that follows the approach she or he desires and believes is best.  Furthermore, we encourage consumers to make healthcare decisions in consultation with an expert clinician after receiving and understanding in-depth information on the benefits, risks and unknowns and uncertainties of a range of state-of-the-art options for care.

Consistent with these rights, we believe that silicone breast implants should be available as an option for women requiring reconstruction.  About 20 percent of patients with breast implants receive them because of a breast cancer diagnosis or other medical need.  Current options for post-mastectomy reconstruction are limited and not all are suitable or possible for every woman.

For many women, an external prosthesis is impractical and saline-filled breast implants often cannot achieve acceptable symmetry, appearance or comfort.  We emphasize that, because women differ by health status and by emotional, physical and personal needs, they require a similarly wide variety of post-mastectomy reconstruction options.

A further important distinction is that women with breast cancer who choose implants do so with a different risk/benefit analysis than those who choose implants for other reasons.  After breast cancer, most women with an implant consider it is necessary, not optional, or even crucial, for their healing and recovery.

Despite the availability of implants over the past decade, in reality, it has been very difficult for women to get implants in controlled studies as many physicians are unaware of the studies or unable to offer them to their patients.  After an extensive review of the science surrounding breast implants, the U.S. National Academy of Sciences' Institute of Medicine reported there is no causal relationship between systemic illness and breast implants as these illnesses occur with the same frequency in women with breast implants and women without breast implants.

The U.S. Federal Court's National Science Panel, the European Parliament's Scientific and Technology Options Assessment and the British Ministry of Health all reached the same conclusion.  In its recent report to Congress, NIH reported that a large meta-analysis following the IOM review of the literature concluded that there was no evidence of an association of silicone-gel-filled breast implants with connective-tissue diseases or other autoimmune or rheumatic conditions.

Women should receive accurate information and give their informed consent regarding the risks and benefits of silicone breast implants.  No breast-cancer or any medical procedure is without risk.  Neither are medical devices which do not last forever.  These treatments and devices are not completely predictable or completely understood.  Silicone breast implants for reconstruction should not be held to a different, higher standard.

However, we feel strongly that women considering silicone-gel-filled or any type of breast implants must receive current detailed and accurate information about the device's risks, benefits and complications including incidence of capsular contracture and rates of rupture.  In addition, we advocate ongoing data collection and reporting based on long-term follow up of all women who receive silicone implants for reconstruction.

Follow-up mechanisms such as privacy-protected recipient registries and close physician tracking will be needed to facilitate medical professional education training and disclosure data compilation and analysis, regulatory oversight and prompt communication of significant findings to consumers, panels and providers.

In conclusion, the undersigned twelve organizations recommend that the panel make silicone breast implants available for reconstruction use for women with breast cancer, breast disease or other medical needs.

Do I have time to read the organizations?

DR. WHALEN:  No; I am afraid that is the end of your time.

MS. SCHERR:  Okay.  Thank you.

DR. WHALEN:  Thank you for your comments.  Just two technical notes before our next speaker.  One is, apparently, the light does not flash at 30 seconds but simply turns yellow.  So you can be aware of that.  Or it will turn red at the end of your time.

The other technical note for the audience, if you are in possession of cell phones or pagers, I give you three choices; silent mode, off or outside.  Thank you.

The next speaker?

MS. HENDERSON:  Good morning.  My name is Sherry Henderson and I am from Bossier City, Louisiana.  I came here today at my own expense with no conflict of interest.  	I had fibrocystic breast disease and precancerous cells which resulted in a double mastectomy and reconstructive surgery with implants.  During my first year of the implants, I had chest pain and headaches.  Within the next three years, I had hair loss, muscle spasms, irritable bowel, reflux, chronic fatigue and rashes.

By the eleventh year, a sonogram verified a silent rupture I had and I had my third surgery with a tram-flap. That was a six-hour surgery and the medical cost was $41,000.  I was fortunate to see a doctor who studied implanted women for over 20 years.  I was diagnosed with several diseases including lupus likeness, rheumatoid arthritis, fibromyalgia, myositis, organic brain syndrome, high blood pressure and autoimmune diabetes.  I am now on disability.

I would like to ask the FDA to have the implant makers to do at least ten years of study and research.  They should be required to pay for a national implant registry and require all women in it to have yearly screening examinations for local and systemic complications.  Let unbiased researchers do the screening.

The Inamed research shows very high complication rates and an increase in symptom-like fatigue in just the first two years after getting implants.  But to know the real consequences, we must have longer research.  Breast implants are not life-saving devices.  If they are not safe, they should not be approved.

Like thousands of breast-implant women, I live in pain every day and I pray that my grandchildren and my nieces do not go through the pain that I have gone through.  The higher rate of brain and lung cancer, more suicides and the study linking implants to fibromyalgia are warning signs that silicone gel should be made less available, not more available.

We are tired of all the suffering.  There are young women becoming 100 percent disabled and the government is finally coming after the manufacturers.  If the FDA refuses to approve this defective product, the implant makers should be better research and develop safe implants.

In the past, the FDA ignored hundreds of thousands of adverse reports reported to them dating back in the '70s.  Breast-implant women have copies of these reports and would be glad to provide them for anyone to see.  These reports document horrible health problems and unbelievable complications with these devices.

Please make the right decisions for the sake of women who trust your opinion and whose lives depend on your trust.

Thank you.

DR. WHALEN:  Thank you.  The first two speakers representing national organizations were each allotted five minutes, for audience knowledge.  The next set of speakers for some time will be individuals and each will be allotted three minutes.

The next speaker, please.

MS. KITTLE:  Good morning.

DR. WHALEN:  You can take advantage of our setting it up and just go ahead and take a few extra seconds.

MS. KITTLE:  What did you say?

DR. WHALEN:  You can go ahead and start.

MS. KITTLE:  Oh, I'm sorry.  My name is Carolyn Kittle.  I am from Shreveport, Louisiana.  I came here at my own expenses.  I have no conflicts of interest.

I am here today to testify about seeing a loved one experience pain from breast implants.  I have two sisters that have had breast implants.  My younger sister had breast implants for cosmetic reasons.  After having two children, her breasts were sagging and her husband wanted her to have bigger breasts.  She was only 22-years old.

She immediately starting having problems.  They became hard.  Five ruptured and she has had at least five breast surgeries.  Her health has deteriorated.  She has been diagnosed with chronic fatigue, fibromyalgia, muscle pain, MS and many strokes and white lesions on the brain.  She is only 46-years old and cannot work due to her pain and the MS.

My other sister had fibrocystic disease and cancer cells, had to have a mastectomy with reconstructive surgery and implants.  She soon was debilitated by headaches, muscle spasms, muscle pain, irritable-bowel syndrome and reflux.  She had a silent rupture and three surgeries.  I have seen her go from a vibrant wife, other and professional to a woman coping with lupus, fibromyalgia, thyroid disease and autoimmune diabetes.  She has had three breast surgeries.

I grieve to see both my sisters go through what they have gone through as a result of a product that, even today, has never been tested for long-term safety.  My sisters have to travel hours away from their home to be seen by a knowledgeable doctor to treat their medical problems.

Inamed's own research shows very high complication rates and they only studied women for two or three years.  Just imagine what the data would have looked like if they followed the women for another few years.

I am asking the FDA to please make the manufacturers study women for at least a few more years before considering approval.  If you approve the implants first and tell the company to continue to study afterwards, the FDA has no authority to make them complete the study and you see, from the company's study so far, they won't do more than required.

The European Union, on September 1, 2003, has reclassified all breast implants from a moderate risk category to the highest risk category for medical devices.  Now is not the time to lower our standards here in the U.S.  Scientists have found breast-implant women have higher rates of suicide and are more likely to die from brain cancer and lung cancer compared to plastic-surgery patients who smoke just as often as implant patients do.

The FDA's own study shows the risk of leaking implants and those leaking implants are difficult and very expensive to remove.  I urge the FDA to reject approval until further research and long-term data are available.  We need to know what implants do to the human body year after year.  I ask the advisory panel to the FDA to listen to all the ill women, make manufacturers do better research before you consider approval.

Thank you and God bless.

DR. WHALEN:  Thank you.

MS. ANDERSON:  My name is Elizabeth Anderson and I come here from Florida at my own expense.  I have no conflicts of interest.

Until last month, I was director of an on-line health hotline and I have spoken with hundreds of women who have had breast implants or who wanted to get them.  They wanted to know if implants were safe and if implants could make them sick.  Many wanted to have their implants removed but they either had no money for the surgery or needed to find a doctor who would take their problem seriously and who would treat them.

Many of the women I spoke with were sick with problems ranging from mild to devastating.  While some got sick shortly after getting implants, most were fine for ten years after implantation and then they started getting sick at that point.

Women told me they were suffering from multiple illnesses and of doctors who were unresponsive to any suggestion that their problems could have been caused by their silicone implants.  Their doctors sometimes told them that they were depressed and that was why they were so ill, but it seems more logical that, because they were so ill, they were depressed.

When you hear so many stories of ruined lives, you understand this.  Some became so sick they could no longer work which, of course, meant the loss of income and health insurance.  Can you imagine believing the medical device you have chosen is perfectly safe and thinking you are going to feel better about yourself only to become so sick that you lose everything, even your spouse.

I have spoken with numerous women who have gone through this and I know women in their 30s and 40s who are completely disabled from their breast implants.  I believe that women who have had mastectomies should have a choice for reconstruction, but what kind of choice do implants offer?

There is no long-term reliable research proving the safety of breast implants, not the Mayo study, not the Harvard study, and not the IOM report that is based on them.  Epidemiologists will tell you that these often quoted are seriously flawed because so many women in them had implants for only a few months or years.  Diseases like scleroderma and MS take a few years to develop and to be diagnosed.

Breast-cancer patients have suffered enough but, according to Inamed's data, these women will have more complications from their breast implants than those who receive implants for augmentation.  So not only does a woman have to worry about a recurrence of because but she has to worry if her breast implants are going to make her sick.

I am sure everyone here realizes the only thing most people need to hear about a medical product is that it is FDA approved.  That is the seal of safety.  People believe the FDA protects them from untested and unsafe products.  To be approved, silicone-gel breast implants should be proven safe for long-term use by research on a large cohort over a long period of time, at the very least, ten years.

The FDA is mandated to make decisions based on science but using all long-term research that is available is a hollow statement when there is no solid long-term research available.  Everyone knows that three years of research is not sufficient.  Since all Inamed's silicone-gel breast implantees since 1998 should have been enrolled a study, we should have at least a four-year study by now.  But, where is it?

The study that Inamed started in 1990 could have been an excellent study but they neglected to include more than a few cancer patients and then most of the augmentation patients were lost to follow up making the study useless.

DR. WHALEN:  Can you conclude, please.

MS. ANDERSON:  Yes.  The FDA did not want to ban silicone-gel breast implants in 1992, allowed them to remain on the market to still be studied.  I ask, where are the data from those women?

MS. CUMMINGS:  Good morning.  My name is Arlene Nicole Cummings.  I came here at my own expense and will not be reimbursed.

At the age of 12, I had a benign breast tumor removed from my right breast.  This left me with an asymmetry problem that only got worse as I got older and had my children.  Fixing my breasts was something I had thought about for years.  After I was done having my children, I felt the time was right.

I did a lot of research on line and found a great deal of information but nothing personal.  This is what prompted me to start implantinfo.com back in 1998 after I had my own surgery.  implantinfo is now the largest community website for breast-augmentation patients where women can read the pros and the cons about breast implants and go through photos and stories of thousands of other women.

I am here today primarily because I am a satisfied breast-augmentation patient.  However, I never had the option of choosing silicone back in 1998.  After talking to women and doing my own research, I would choose silicone now over saline if it was available to me.

On my website, I have corresponded with countless women who, for many reasons, probably should have had the silicone option but did not.  They were denied silicone implants, had an unsatisfactory result with saline, and then had a second surgery to replace the saline implants with silicone.  So they actually had to have two surgeries to get the result they wanted the first time.

I am also aware of the frustration that surgeons feel because they can't give a patient what they know, in some cases, will produce a better result for them.  The risk of a second surgery, in my opinion, is far greater than the risk of silicone implants.

Ironically, breast-cancer patients already have compromised immune systems but, for the most part, receive silicone implants.  Are they at less of a risk of the supposed risks of silicone implants?  I believe silicone implants provide a far better result for reconstruction.  If these patients get breast implants and get a better result with silicone, shouldn't they be available to all of us?

Women, along with their surgeons, should be allowed to choose silicone or saline.  We are neither ignorant nor shallow and we are not seeking breast implants with informed consent.  We have researched the procedure for years.  The average age of a woman on my website is 34.  Most of us are married with children and just want to get back what we lost after pregnancy and nursing.

I believe in informed consent and women should be told all the risks.  But, in the end, after we receive that information, let us decide, along with our surgeons, silicone or saline.

I have passed out to a handful of testimonies from other women who, for many reasons, could not physically be here.  There are thousands more on implantinfo.com and I encourage you to read about them and give us back the choice of saline and silicone.

I think you for your time.

MS. COLOMBO:  Good morning.  My name is Michele Colombo and I come here on my own behalf and I have no conflict of interests,

I am 34-years old and married.  For many years, I have been dissatisfied with the size and appearance of my breasts.  I am very self confident and I am an educated person but have always seriously considered improving my appearance by undergoing breast augmentation.

I have been researching breast augmentation for approximately five years including reading the FDA website as well as other resources on the Internet.  In doing so, I have satisfied myself that silicone implants are a safe alternative for my breast augmentation.

Unfortunately, in the early 1990s, silicone breast implants were banned for sale to people like me in the United States.  If I lived in Europe, I could get silicone breast implants.  If I traveled to Europe, I could silicone breast implants and come back to the U.S. and live here with them even though they are unavailable here.  If I had breast cancer, I could get silicone breast implants.  What's more, if I were a man, I could get silicone testicular implants and if I required a shunt for medical reasons, it would also be silicone.

But, unfortunately, for me, I am just a woman seeking to undergo a cosmetic procedure.  Moreover, what I seek is a procedure using a product which has been shown in numerous studies to have no more long-term heath risks than saline breast implants or other silicone devices.

My problem, it seems, is that the FDA, up until now, has viewed people like me as unable to decide what it best for me.  I could make that decision if I had cancer but simply because I want to improve my physical appearance, that is not enough to justify the risks, in someone else's opinion.

I think it is time to reevaluate the true motivations for the ban in silicone breast implants.  The fact that breast-cancer patients may obtain them and cosmetic patients cannot exposes the true reason for the ban.  What is at stake is a moral judgement rather than a medical one.  If my breasts are completely deflated from breast feeding, weight loss or age, improving my appearance with breast implants is deemed unacceptable because it would be for vanity.  If I had had a mastectomy, it would be acceptable because it could be justified as medically necessary.  The difference is a moral one and not the a medical one.

I urge you to remove the ban on simply breast implants and let women make informed decisions about their bodies the way millions of people do every day with countless other devices and medicines that are approved by the FDA.

Thank you for your time.

DR. JOINER:  Good morning.  I am Dr. Thomas Joiner, the Bright-Burton Professor of Psychology at Florida State University.  I have authored a peer-reviewed paper on the association between breast augmentation and suicide as well as six books and over 200 scientific articles on topics related to mental disorders including suicide.

I was awarded a Guggenheim fellowship for research on suicide and was given the Schneidman Award for Excellence in suicide research from the American Association of Suicidology.  My travel expenses to this hearing were paid by the American Society for Aesthetic Plastic Surgery.  I have no financial ties to any implant manufacturer and no involvement in a pending law suit related to breast implants.

I thank the panel for the opportunity to address the relation between breast augmentation and suicide.  To date, there are four studies indicating that women who undergo breast augmentation have higher suicide rates than women in the general population.  It is crucial to note that breast-augmentation patients are not representative of the general population and the ways in which they differ from the general population are very relevant to suicide risk.

The crucial point it this.  Breast-augmentation patients should be at relatively high risk for suicide for reasons that have nothing to do with breast augmentation.  Given the demographic and other characteristics of breast-augmentation patients, their suicide rate should be higher than that of the general population.

Consider race, for example; virtually 100 percent of breast-augmentation patients are white whereas the general population of women is, of course, more diverse.  This is relevant because white people are more likely to die by suicide than non-white people.  Breast-augmentation patients' risk for suicide is increased partly because of the simple fact that they come from a racial group with relatively high suicide rates.

This same exercise can be repeated with the same conclusions for each of the following factors; age, marital status, alcohol use, personality and symptoms of mood and other disorders.  In each case, breast-augmentation patients appear to differ from the general population of women and, in each case, the difference is such that suicide risk is increased.

I calculated the expected suicide rate in breast-augmentation women given their demographic and other characteristics.  A conservative calculation produces an expected suicide rate among breast-augmentation patients that exceeded the population rate by fourfold.  For reasons having nothing to do with breast augmentation, the subgroup of women who undergo the procedure should have relatively high suicide rates because their demographic and other characteristics put them at high risk.

Given that expected rates are fourfold the population rate, and actual rates are lower in every study to date, the possibility arises that breast augmentation actually confers protection from suicide a possibility that squares with the finding that vast majority of breast-augmentation patients report high satisfaction with the procedure.

A safe conclusion appears to be that actual suicide rates are no higher than and may possibly be lower than expected suicide rates among women who undergo breast augmentation.

Thank you.

MS. CUNNINGHAM:  Good morning.  My name is Susan Cunningham.  I came here at my own expense.  My biggest concern is that these implants do rupture and there is still no medical information on how to recognize a rupture and what to do about it.  It is a frightening experience to be so sick and not have a clue what is wrong with you.

Many women cannot afford an MRI to check and, even if they find out their implants are ruptured, they don't have the money needed for surgery.  My own experience began after I had breast-fed both of my children.  I decided on implants and arranged for a consultation with a highly respected plastic surgeon.  I had concerns, but the surgeon said the implants would last a lifetime.

I was fairly pleased with the results, but it didn't last long.  I began to experience loss of nipple sensation and numbness.  Then I experienced capsular contracture in the right breast which is very painful.  As the years went by, I was plagued with sinus and respiratory problems.  My health continued to decline.

Finally, in 1993, I consulted my plastic surgeon.  As he wrote in his notes, "I reassured her concerning this and advised her to leave her implants in place unless there are further difficulties because they look very good."   My body was wracked with joint pain.  I was diagnosed with Raynaud's and suffered from constant muscle spasms.

My breathing and chest pain were severely affected and I had been in the emergency room several times.  Finally, tests indicated remarkable changes in my immunological system characterized by an unexplained decrease in the total CD4 or helper cells.  I was getting sicker and experienced a fatigue so devastating that I would weep from the smallest effort.

My concentration was zilch and I started having huge gaps in my memory like pulling out of my driveway and then wondering where I was going or what I was doing in the car.  It was frightening and it became worse.  Finally, I was forced to quit a career that I loved.  At this point, it became necessary to sell our family home.

I sought out Dr. Feng whose practice was 120 miles away.  She informed me that at least one implant was ruptured.  They were removed in 1995, both ruptured, and I promptly noted relief in the chest pain I had been experiencing for years.  Within a few months, my bloodwork, which for many years had been abnormal, returned to normal except for IgE levels and a positive antihistone antibody which took an additional two years to return to normal.

My overall physical condition has taken a few years to improve.  I still suffer from fibromyalgia-like symptoms although attacks are less frequent and of shorter duration.  At the ripe old age of 61, my mind is much sharper and my body much stronger than it was just a few years ago.

Today, I live a far different life than I could have lived if it were not for ruptured implants and their devastating effects to my life.  I am on disability which was granted on silicone rupture and related problems.  It was a very black day for me to finally accept the fact that working is no longer an option for me.

I am not looking for pity because I have been blessed with many other things in my life.  But, if there is anything I could do to help another woman to avoid these devastating effects of ruptured or leaking implants, then I must try.  That is why I am asking you to not approve silicone-gel implants until the company provides long-term safety data proving that leaking silicone and immune diseases can be avoided.

The company's application for silicone-gel breast implants was rejected in 1991.

DR. WHALEN:  Can you conclude please.

MS. CUNNINGHAM:  Thank you.  They had twelve years to collect data but they only collected three.  They will only do better research if approval is denied until they do.

Thank you.

MS. STANSELL:  My name is Anne Stansell.  I came here from New Mexico at my own expense.  I am a breast cancer survivor.  I was diagnosed at age 39.  The doctors said I needed a mastectomy, radiation therapy and breast implants.  Implants were just part of the treatments.  No discussion.

I trusted the doctors who I felt had just saved my life.  I didn't know there was no safety study of breast-cancer survivors with implants.  That had never been done.  I was all right for the first few years.  Then I became very ill.  I was diagnosed with Graves' disease, fibromyalgia.  My eyes were so dry that one of the retinas tore.

My implants were taken out about two years later.  I had to fight with my insurance company to cover the removal.  Half of one of my implants was gone.  Where did the silicone go?  I don't know.  Even with silicone left inside of me, almost immediately, I began to feel better.  My family really noticed a difference even before I did.

I am still recovering, but I can work some now.  I thought of my own experience when I heard about the Inamed study of breast-cancer patients.  I had many of the same local complications.  I can't even remember how many surgeries I suffered.  Silicone was found in my side where it migrated from a broken implant.

Inamed cancer patients also had an increase in some autoimmune symptoms during the first two years.  I think my symptoms started at the third year or so so it is likely that those signs and symptoms will increase over time, just like mine did.

It doesn't do cancer patients any favor to give them the kind of choice I was given.  The choice isn't a choice if it is not a safe choice.  Informed consent isn't possible if the physicians think the FDA approval means a safe product.  A recent NIH study shows a doubling risk of brain cancer and a tripling risk of lung cancer in breast-implanted women.

Health Canada has now started a study in 1996 to determine the cancer risk in 40,000 breast-implant women who were implanted from 1974 to 1989.  This data should be available later this month.  It is outrageous to continue to allow the marketing of a device that may increase the risk of cancer in women who are recovering from breast cancer.   	Your job is to look at the research, but I hope you will listen to the patients, too.  We illustrate the data.  We are the examples of what can and what has happened to tens of millions of women around the world.

MS. RUST:  (Read by Beth Nichols.)  My name is Beth Nichols and I am reading this on behalf of Rose Rust who is too ill to be here today.  I have no conflict of interest.

"My name is Rose.  I am 57-years old.  After I remarried in 1990, I thought I needed to look younger so I opted to get silicone breast implants made by McGhan, now called Inamed.  I had talked to a plastic surgeon and he told me they were safe and they would last my lifetime.

"Within four months of getting the implants, I noticed I was aching all over and had headaches.  I went back to the surgeon and told him I thought they were making me sick.  He told me I was listening too much to the media and he saw no reason to remove them.  I believed him and left them in.

"I became more ill and the headaches became almost unbearable.  I had so much pain that I had to quit my job.  I had become so ill that I was put on Social Security disability.  In the mid-'90s, my nasal passages started burning and my nose was swelling shut in addition to the horrible headaches.  My GP sent me to a neurologist who did many tests and sent me to an ENT.

"I ended up having three sinus surgeries.  The three different doctors all told me that something was inflaming my nasal passages.  They told me that if I didn't find out what my body was reacting to, they couldn't help me.  I went to an allergist and he found no allergies.  We had our house checked for any toxins.  My husband checked our vehicles.  By now, I was feeling much worse, dizzy and nauseated.

"Over the next few years, I went to many doctors including the Mayo Clinic, but nobody could find out what was wrong with me.  The last two or three years, I have become so ill I am almost bedridden.  My nasal passages burn all the time and my headaches are constant.  My legs and feet burn, swell and hurt so badly, I can hardly walk.  I am dizzy and nauseated.  I have lost much of my short-term memory.

"Over the last few years, I have become intolerant of all chemicals.  When I am exposed accidentally, my nose swells shut, nausea and headaches worsen and my heart acts up.  In June of this year, I finally decided I was going to die if I didn't do something.  In August, I had my implants removed.  My left implant was ruptured and the plastic surgeon said most of the silicone had leaked out, migrating into my chest and into my lymph nodes.  She described it as 'a sticky mess.'

"She also removed the lymph nodes with silicone in them.  At that time, she told me there was probably silicone in my liver, lungs and brain.  The right implant was intact.  I have no idea how long the left implant had been ruptured.  I saw no difference in how I looked.

"Through the years, I have become unable to even take care of my home.  Sometimes, I am not able to fix meals for my husband.  I spend many days in bed.  Long ago, I gave up on social events.  My husband tells me I actually am much worse off than I realize.  It is just a struggle to survive.  If the doctors had not told me my illness was unrelated to my silicone implants, I would have had them out many years ago and perhaps avoided much of the silicone leakage.

"I think of how, in 1991, the FDA decided not to take silicone-gel implants off the market.  It sent the clear message that silicone-gel implants are probably okay.  I feel many of us have been the guinea pigs and I feel we have proven that silicone is not safe because the implants can leak throughout our bodies.  It is too late to do much for women like me, but let's please save the next generation.

"It is your responsibility to protect us and hopefully you can restore my faith in our FDA.

"Thank you."

MS. PATE:  Hello.  My name is Kathy Pate and I would like to thank the panel for the opportunity the speak today on behalf of silicone breast implants.  I am a registered nurse in Albany, Georgia and I work in an operating room there.  I have never had any financial relationship with any medical-device corporation or their competitors.  I am here representing myself and I have paid my own way.

I have breast-augmentation surgery in December of 1983 because I had difficulty buying clothes to fit.  As many other women have stated, after I breast-fed my baby, I didn't have very large breasts to start with and then, after that, the problem was even worse.

I became quite frustrated with this problem so I decided to have breast-augmentation surgery.  At that time, the quality of saline implants was very poor.  The possibility of deflation was very high.  At that time, also, there was no discussion about the serious complications related to silicone implants when I had my surgery.

My plastic surgeon did tell me that there were possibilities of localized irritation if the gel ruptured.  He also said there was a possibility of capsular contracture.  I went ahead and chose the gel because they looked more natural and they were lighter.

When the media put out the information on silicone implants, and they were removed from the market in the early '90s, I wondered what should I do.  Should I have them removed?  What should I do?  So I pondered the thought and decided that the best thing to do was see if anything developed.  I have not developed any problem with silicone implants and really have no concerns about their safety.

I monitor my breasts with visual and physical exam and I have an annual mammogram.  I stay in close contact with my plastic surgeon and we decide we would monitor for any changes in shape or consistency to detect an early rupture.

I realize there are risks associated with all types of surgery and especially with any implantable device.  Although it has been twenty years since my surgery, and I am sure at some point I am going to have to have those replaced, I would definitely do the same procedure again and I will go back to silicone.

I believe firmly that women should have the option to choose saline or silicone and that they should have the risks and benefits outlined for each.  The individual receiving the product needs to have confidence in the product, itself, and also in the surgeon who is performing the procedure.

I believe women who are well informed will make the decision that works best for them.  Given all the research and investigation of these devices, I feel certain that the product has been vastly improved to resist rupture and prevent complications.  Breast-implant surgery has allowed me to be comfortable buying clothes and not reduced to buying bras with heavy padding to fill out my clothes, and even bathing suits.

The real message I want to convey here is that I feel like the women need to have a choice and be informed of the risks and benefits and have the option to choose.

Thank you.

MS. BRUNING:  Good morning.  My name is Nancy Bruning and I have no conflicts of interest.  I was told I had six minutes.  Can I have six minutes, five minutes?

DR. WHALEN:  You are here with Ms. Peterson?

MS. BRUNING:  Bruning.

DR. WHALEN:  I know your name is Bruning.  You are listed as being here with Ms. Peterson.  That is why the two of you together were given six minutes.

MS. BRUNING:  I see.  Well, I am wearing three hats, so--thank you.

I am wearing three hats.  I am here today as a breast-cancer survivor.  I am here as the author of a book about breast implants that is currently in its third edition.  And I am here as a graduate student completing her Masters degree in public health.

I became interested in breast implants because of my own experience with breast cancer and its aftermath.  You would probably never suspect it to look at me or to hear me, but I have had my chest cut open seven times, twice for breast cancer and five times to fix what the breast-cancer surgery wrought on my mind and body.

My most recent surgery was to remove the gel implant that had ruptured along with the entire scar capsule that surrounded it, and it was replaced with a saline implant.  I am here to tell you that, although this could be dismissed by some as merely local complications, this surgery caused me more pain and disability than any of the other surgeries including the mastectomy.

I am here because, if I had to do it all over again, you would have to tell me that, based on the studies that exist including the new Inamed study, breast-cancer patients have a very high probability of these things happening including a 46 percent chance of reoperation in the first three years.

But can you tell me what the rate might be down the line?  No; you cannot.  I am here because, if I had to do it all over again, I would like you to tell me the odds that any silicone that escaped the scar capsule will cause systemic disease down the line.  Can you do that?  No; you cannot.  Can you tell me the odds of that happening with the new Inamed implant?  No; you cannot.

The subtitle of my book is Everything You Need To Know About Implants and one of the first things I tell my readers that they need to know is that we don't know everything.  Far from it.  This was true when I first wrote the book in 1993 and it is only a little less true now.  And what we do know is not good.

During the process of writing and twice revising the book, I interviewed scores of plastic surgeons, breast-cancer survivors, implant patients, researchers and advocates.  The one thing that is clear is that silicone breast implants were sold to almost 1 million women before any research on women was published in the peer-review journals.

Currently, over a quarter of a million American women have breast implants every year.  This could be a disaster waiting to happen if silicone implants become available without proper long-term study.

We now know that silicone-gel implants break after about ten years, on the average.  We now know that leaking silicone can migrate to the lymph nodes, lungs and all throughout the body.  We know that, in several studies, women with breast implants were more likely to die from certain types of cancer.

And we know that, in three studies, women with silicone implants are three times as likely to commit suicide.  In the Inamed study, quality of life, including social interactions and self esteem, actually declined within two years of getting their implants.  This certainly raises questions about whether implants make women feel good about themselves as promised.  Results amy be even worse than the studies indicated.  Let's remember that most women do not like to admit mistakes or to criticize their doctors to their faces.

Regarding the research conducted by Inamed, in light of what we know already, I wonder why they thought two or three years of research was enough.  Although the diagnosed rupture rate was low, 6 percent after three years reconstruction, we must assume that they are higher since three out of four ruptures are not apparent without an MRI.

If you share my concerns that only half the breast-cancer patients were followed for even three years, especially since Inamed has been selling gel implants to cancer patients since the late 1990s.  Breast-cancer patients deserve better.  I am the living proof that they are living long lives.

We deserve implants that don't result in additional surgery for 46 percent of us in the first three years.  We deserve long-term studies that can provide accurate information about the risks of developing serious implant-related health problems.  We deserve to have a choice.  Unfortunately, we do.  Most breast-cancer patients can now choose lumpectomy, saline implant, silicone-gel implants or autologous construction.  Choice is good but we deserve to have safe choices.  Silicone implants should only be approved if they are proven to be a safe choice for all women.

If there is a problem with the way studies are conducted and there is a problem with recruitment and compliance, then the answer is to fix that, not to make silicone-gel implants available.

Since I received my diagnosis of cancer over 23 years ago and experienced the first inklings that breast implants were not as safe as I was led to believe they were, I have been living with two kinds of uncertainties.  One, will I get breast cancer again and, two, will my implant cause me harm?  I have seen studies that address the first question.  Where are the good studies that address the second?  Inamed has not provided them for the women who currently have silicone-gel implants or for the millions more who would if the FDA approves them.

I have suffered more than enough and don't want other women to follow in my footsteps.  We deserve to have better science than the implant companies have given us and it is your responsibility to use your power to make sure that they do just that before approval because it won't happen afterwards.

Thank you.

DR. WHALEN:  Just to make it clear to the other speakers, to make sure that everything seems fair, the reason somebody would have six minutes as opposed to three is because two speakers who have preregistered with FDA have combined their time and one of those two individuals is addressing us.  So, if you are wondering why you only have three and somebody has six, that is the reason.

MS. FALLOWS:  Good morning.  My name is Jill Fallows.  I am a nurse as well as an attorney.  I do not have any financial relationship to Inamed or any of its competitors and, as a note of levity, as a trial attorney, I rarely find myself speaking before a group of such receptive physicians.  I am appreciative of the opportunity.

I will address the issue of choice this morning.  It seems as though Inamed has framed its pro-silicone-breast-implant argument in terms of a woman's right to choose the device.  But simply offering a woman the right to choose is a far reach from providing relevant, unbiased, sound medical information upon which a reasonable woman can make an informed choice.

It is irresponsible to argue that a woman is entitled to choose a medical product while simultaneously glossing over the relevant safety data in the device's promotion and labeling.  We want women to have choices.  But we have learned from the Dalkon Shield and DES that they are not safe products and should not be offered to women simply because we cherish the notion of an individual's right to choose.

This is particularly true at today's meeting because the silicone-gel breast implant does not purport to relieve the discomfort or pains commonly associated with diagnosed illnesses and conditions.  Indeed, we don't want to offer women a choice of an implant knowing that it could rupture and leak silicone gel into their lymph nodes, lungs and other body parts where it cannot be removed.

We don't want it leaking into their breasts so that removal results in a mastectomy to remove silicone rather than cancer.  We should be mindful that breast augmentation is a completely elective cosmetic procedure and the risks to the patient should be clearly outweighed by its benefits.

We need to empower women and their physicians with adequate, accurate, unbiased information broad enough to address the diversity among the 50 states informed-consent laws and yet specific enough to affirmatively warn women that they are buying a product with a litany of known serious risks.

We take, as a starting point, the notion that a doctor or a patient must be apprised of information upon which a reasonable person can rely and make an informed choice about the safety of the implant.  It would not be reasonable to minimize or hide the risks from the physician or the patients.  The manufacturer's product literature should especially make mention of the rate of risks of complications and a black-box warning would be minimally acceptable if it stated that the manufacturer has not adequately completed five-year safety studies.

I am most appreciative your time.  Thank you.

MS. SANTORO:  Good morning.  I drove here on my own expense and have no conflicts of interest.  My name is Elizabeth Santoro.  I have been a nurse for three years and I recently received a Masters of Public Health from Johns Hopkins.  I will be discussing problems with data quality.

A major area of debate is whether silicone breast implants cause autoimmune diseases such as scleroderma, lupus, rheumatoid arthritis and fibromyalgia.  In 1999, the Institute of Medicine reviewed the 17 studies available at that time and found that most research found no significant link.

A year later, a meta-analysis of the same studies agreed with that if the largest study by Hennekins, et al., at Harvard was excluded.  If the Hennekins study was included, there was a significant link between implants and these diseases.

I would like to highlight some of the flaws in the studies that were both in the IOM report and the meta-analysis.  First, not all the studies were published in peer-review journals.  Second, many of the studies were based on hospital records.  This is problematic because a lot of people who suffer autoimmune diseases are not hospitalized for symptoms such as pain and fatigue.  Therefore, data from hospital records will not identify them.

Medical records were also used as a primary source of information instead of a comprehensive medical exam.  The absence of a medical exam creates the opportunity where  valuable health indicators of adverse events can be missed.  Often, vague and nonspecific symptoms can be present in the beginning of an autoimmune process.  As a result, medical records may not be the best way to detect this class of disease.

Most of the studies also included women who had implants for a very short period of time.  Autoimmune diseases can take at least five to ten years to develop.  So, including women with implants for only a few months or years cannot accurately capture all of the autoimmune diseases that might develop.

I would also like to comment on the importance of sample size.  Many of these studies had women with just a few hundred.  That does not sound bad but a larger sample is needed when rare diseases such as scleroderma is being assessed.  The research that has been conducted in the past and reviewed by the Institute of Medicine is not sufficient to draw conclusions about the safety of breast implants regarding autoimmune disease.

And, by the way, those studies did not look at morbidity and mortality from other causes or from disorders that were not diagnosed diseases and therefore are useless for drawing any conclusions at all regarding cancers and other disorders.

Also remember that the first Institute of Medicine on Agent Orange found no link to cancer.  The first Institute of Medicine report on Gulf War veterans found no link to illness.  Subsequent research found clear evidence of statistically significant increased risk of serious diseases that were not apparent to the IOM in their first analyses.

I would also like to note how the data that was made available by Inamed is suggestive of a connection between silicone implants and autoimmune disease.  Upon reviewing these data, FDA scientists found an increase in muscle pain, joint pain, hair loss, rashes and fatigue within two years of patients getting implants.  During this same two-year period, it was also found that many patients reported decreases in both physical and emotional health.  These signs and symptoms are characteristic of several autoimmune diseases in their early stages.

In addition, FDA scientists noted previous studies found an excess of some cancers and leukemia among implant patients.  These studies were published after the IOM report.  The fact that implants may delay cancer detection via mammography is also troubling.  Moreover, the FDA's own study found that silicone implants break within about ten years often with no warning signs and the clinicians and pathologists have reported numerous cases where silicone gel had migrated to the abdomen, groin, fingers and axillary lymph nodes.

These findings raise grave questions about the safety of silicone-gel implants.  They provide evidence that Inamed's implants are associated with many serious complications and declines in both physical and emotional health status even within the first two years.

The data supplied by Inamed, combined with the limitations of the studies in the IOM report, show that there are many unanswered safety questions in regard to silicone-gel breast implants.  As a result, they should not be approved by the FDA.

Thank you.

MS. PETERSON:  (Read by Jennifer Brooks.)  Good morning.  My name is Jennifer Brooks.  I am reading testimony for Evon Peterson who cannot be here today due to an illness.  She has no conflicts of interest and neither do I.

"My name is Evon Peterson.  Ten years ago, my silicone-gel implants were removed.  I was an R.N. and had implants for reconstruction for following bilateral mastectomies.  The doctors acted as if reconstruction was a given, something I should do.  It wasn't presented as a choice.

"When I was explanted in 1993, the right implant was found to have ruptured and the left implant was severely leaking.  Three additional surgeries in 1994, 1995 and 1996 followed to remove bilateral siliconomas.

"The first symptoms to appear after implants were the first to resolve or diminish when the implants were removed.  A few weeks after reconstruction--what an oxymoron that is--the firmness and contractures began.  The shape of my breasts became distorted and the surrounding tissue became inflamed.  Lying on my right side or freely using my right arm was restricted.  Searing pain invaded by back and neck.

"After that, more obvious systemic symptoms developed.  By 1989, at age 45, I was totally disabled in my work and personal life.  I experienced profound fatigue, insomnia, seizures, tingling and numbness of extremities, dizziness, nausea, skin discoloration and hardness, migrating joint pains, chemical sensitivities and balance disturbances.  I felt like I was in a mental fog.  I had contemplated suicide as an option to stopping the pain.  I know if I had kept the implants in any longer, I would have committed suicide.

"One of the worst aspects of this experience has been in the continuing denial of manufacturers and the medical community.  They fail to study implants and their related complications over a long period of time and then arrogantly claim that the products are safe for long-term use.

"Prior to explantation, I faced a paradox.  I was too ill to have surgery and too ill not to have surgery.  I am glad that, on the recommendation of three of my many physicians and also based on my increasing knowledge, I chose to improve the quality of my life by removing my implants.

"Despite living with multiple autoimmune processes and complications, my health has improved from about 10 percent functioning to 60 percent functioning.  Considering that I have also aged ten years, I am delighted.  I remain convinced that, had I not been explanted, I would have died within the year.

"I am unable to physically present today because of pneumonia.  However, I petition you to require unbiased long-term studies before any decision is made about approval.

"Thank you."

MS. WEBER:  Good morning.  My name is Elizabeth Weber.  I live in Southern Maryland and I was not paid to come here today.  The opposite is true.   I asked to come today to share my story.

My mother had breast cancer and had a bilateral mastectomy.  I, too, had a bilateral mastectomy in June of 1999.  I was aware of the controversy of silicone breast implants and, therefore, opted for the saline implant.  I am a very active person.  I trained for the Olympic Games since age of nine in the equestrian sport.

Pain and perseverance is not new to me.  The saline implants were very uncomfortable to me.  Dr. Scott Spear did a wonderful job of my reconstruction and the outward appearance of the saline implants looked good, but I hated the way they felt inside.  They were hard and hurt from the inside.  I felt like I had a mass pushing against my chest.

I stopped doing the sports I used to do and it was hard to sleep at night.  One year later, I went to Dr. Scott Spear and asked to have the saline implants changed and replace by silicone implants.  After I woke up from my surgery, I could feel such a huge difference.  After I healed from the surgery, I felt like myself, a whole, natural and complete woman.

The outward appearance was impressive and beautiful.  But, even better, my breasts felt like mine from the inside.  I now do everything I used to and more.  I hope you give other women the same choice.

Thank you for your time.

ED BRENT:  Good morning.  My name is Ed Brent and I have received no funds nor do I have a sponsor that has helped me to fund this trip.

Three years ago, my wife stood before you and did her best to convince you, the FDA, not to allow saline breast implants on the market.  My wife had double-lumen implants, my daughter, Catherine, leg braces, and that was because of the silicone implants, carried one of my wife's implants to your table showing the black fungus growing inside of it.

My wife was P.J. Brent.  You may have seen her on CNN or other t.v. networks discussing breast implants.  She was distraught after being here and learning that everything that she had said was anecdotal evidence and not to be considered by the panel.  We were here March 1, 2000.

On May 29, 2000, my wife committed suicide.  She left behind seven children and I am here on behalf of my wife and children to urge you not to allow silicone breast implants on the market.

There are three studies showing that women that breast implants are more likely to commit suicide compared to other women.  Some experts think that these suicides mean that women with breast implants have lower self esteem before they got their implants as well as after.  But the National Cancer Institute study, which compares implant patients to other plastic surgery patients, shows that the problem is more likely to be from the implants.

My wife was not a woman with low self esteem.  She was a vibrant, loving wife and mother.  P.J. loved the way she looked the first few years after the implants.  Then she started to get sick.  Her joints hurt.  Her fingers would swell so she could not wear her rings.  She did not know what was happening and sought medical help.  She had lupus-like symptoms and was diagnosed with fibromyalgia.

She was explanted in 1992.  P.J. breast-fed the two daughters she had after the implants.  Catherine was born in 1985 and was diagnosed with chronic inflammatory demyelating polyneuropathy, CIDP, as well as esophageal-motility disorder.  She spent years in leg braces to allow her to walk.  Now, the braces are gone and have been replaced with a wheel chair and she has a special car with hand controls to help her drive.

Christine was born in 1986 and also had esophageal-motility disorder and signs of the neurological disorder.  After P.J. committed suicide, an autopsy was performed from her.  Samples were extracted.  Large amounts of platinum were found in her body.  This came from the silicone.  The high platinum levels, a doctor at CDC saw them and said, "No wonder she committed suicide.  She could not have been in her right mind."

Hair, nail and tissue samples were taken from our daughters who breast fed and it was found that they, too, had elevated platinum levels.

You have an awesome responsibility before you, a responsibility of conscience and doing what is right for the women and yet unborn children of these women who may be forever affected.  Don't let this be a money issue, but a moral issue.

Thank you.

DR. CUNNINGHAM:  Good morning.  My name is Dr. Bruce Cunningham.  I am Professor and Chair of Plastic Surgery at the University of Minnesota.  I am speaking as an individual and have personally paid my way for the appearance today.  I own no stock in any implant manufacturer and receive no compensation of any kind.

I have, however, received peer-reviewed research grants from implant manufacturers and I have served as a medical director of a PMA breast-implant submission.  I do use implants in my clinical practice.

ASPS has recognized and acted on its responsibility to plastic surgeons and their patients.  The American Society of Plastic Surgeons has already established a breast-implant registry within a larger outcomes data-collection activity.  I worked on the development of TOPS, or Tracking Outcomes in Plastic Surgery, which collects plastic-surgery procedural data and clinical outcomes and can include collecting satisfaction data from patients, themselves.

A breast-implant registry is embedded within the Internet data-collection tool of TOPS and can track things such as the number of implants placed or removed, clinical indications and type of facility, anesthesia and short-term complications.  NBIR, or the National Breast Implant Registry, was designed to allow physicians and their highly mobile patients to track implanted and explanted devices.   NABIR was, in fact, so successful in its design that it attracted international interest.  It has served as the template for IBIR, the International Breast Implant Registry, which is poised to become the standard for the European community, Australia and South America.

In the initial piloting phase of NBIR, this data-collection tool over the Internet, 9,000 surgeries were recorded including 14,912 implants and 2,084 explants.  TOPS and NBIR, however, are hindered in their ability to provide benefit for patients and physicians by the stringent confidentiality constraints of HIPAA.  They could be improved significantly to provide transparent tracking of devices, surgeon records and implant-related events.

To do this, however, the breast-implant registry that we have established requires public-health authority status as a part of a Governmental Health Operations Act.  The management, government and composition of the leadership of the NBIR could be changed to reflect the involved stakeholders by including the FDA, the public and industry representative members.

Finally, TOPS could be enhanced to provide an interactive on-line patient-information and education process that could be linked, in turn, to verification and documentation of patient informed consent.

In conclusion, ASPS, the American Society of Plastic Surgeons, is dedicated to providing the gold standard of outcomes data to benefit patients, the public and the profession.  We look forward to a partnership.  TOPS and NBIR can become a valid method for a registry to trace implant-related data and outcomes.

Thank you for the opportunity to present.

MS. McGRAIL:  Hi.  My name is Margaretha McGrail and I am not being compensated or paid in any way to be here.  I live locally.

Four years ago, I made a decision to have breast-augmentation.  It was a personal choice.  I didn't need the surgery.  Years of running and age had caused my breasts to sag and shrink and it was affecting my self-esteem.  Someone referred me to a local surgeon who specialized in breast reconstruction and I went in to speak with him about possibly having the procedure.

After talking with him, I decided the procedure was harmless and that it was worth going forward with.  While discussing the procedure with my doctor, my doctor mentioned that there was a study taking place that I could participate in if I wished.  It involved the opportunity to have silicone implants rather than the typical saline implants.

He explained that silicone implants were much more natural looking and feeling than the saline.  Remembering vaguely that there had been some issues with silicone implants, I asked him about the hazards of silicone implants.  We discussed the pros and cons of silicone and I was given a couple of articles about the topic.

My surgeon felt it was a very safe procedure.  It was my understanding that silicone implants had significantly improved over the years, that leakage was no longer an issue and that the past problems had been due to leakage.  The articles I had been given were very positive and the benefits of my taking place in a study seemed to outweigh any possible problems.  The study provided for regular checkups, periodic MRIs and I thought that was a benefit.

I opted for the silicone implants.  I can attest to the fact that silicone implants are, indeed, dramatically more natural looking and feeling than the saline implants.  I have had no problems at all these past four years.  I have been extremely happy about my decision to have breast-augmentation.  I am very health conscious and take very good care of myself.  I believe in a preventative and holistic approach toward my health, so the issue of possible problems with silicone implants is a concern of mine.

I would do nothing knowingly to jeopardize my health.  I have only had the implants for four years, so I cannot address the potential for problems after ten or fifteen years.

MS. HAMILTON:  (Read by Elizabeth Curtis Hamilton.)  Good morning.  My name is Elizabeth Curtis Smith and I am here of my own will and I am reading the testimony of Gail Hamilton who was too sick to be with us this morning.

"My name is Gail Hamilton.  I got silicone implants when I was 24-years old.  I had had two children and was very disappointed with my appearance after nursing.  I was slim and healthy and was told by the plastic surgeon how wonderful these implants were and how they would improve my appearance and they would last a lifetime.

"Ten days after augmentation, the wounds opened up and there was no appreciable healing.  The implants were removed along with very large blood clots.  The implants were cleaned and reinserted.  This time, the wounds healed.  My body formed scar tissue around the implants causing them to feel like baseballs.

"Seven years after my initial surgery, the implants were removed.  The troublesome scar tissue was taken out and the same implants were reinserted.  By 1989, I was experiencing very uncomfortable pain in my right breast and visited the plastic surgeon again.  He could find nothing wrong so I went home and tried to ignore the sore area under my right breast.

"In the summer of 1993, I started to feel a burning sensation behind the implants against my chest wall.  It started with one side and the other one soon followed a couple of weeks later.  The burning became worse and worse every day and I went back to the plastic surgeon again.  I told him of the burning and how I was experiencing a shooting pain down my arms and legs and through my back.

"The surgeon told me that this pain could not possibly have anything to do with my implants but I would be able to have them removed if I wanted.  I know this sounds dramatic and disgusting, but to try to explain the depth of this pain would say only that I felt like there was rotting inside my chest while I was still alive.

"The pain was so severe I felt I needed to have these things taken out of me.  Nearly 20 years after I sat in the surgeon's office as a young, healthy woman, I was now desperately ill and in unimaginable pain.  On November 16, 1993, I had another surgery and this time the implants were removed and not replaced.

"I had asked for the implants to be returned to me.  Eventually, when I was able to look in the container, I found only one implant.  The other one, the right implant, I was told was too ruptured and could not be returned to me.  At home, I began having one crisis after another.  I wasn't able to get out of bed myself or even sit up.

"I started asking questions.  Why, after this surgery, had I been so ill?  I had had numerous surgeries before and recovered in due time.  Lying in bed, I wrote for the hospital records.  I questioned my surgeon.  One day, as I was lying there reading over the paper, something caught my eye.  My surgery had begun at 1420 hours and was completed at 1443 hours.  Those 20-year implants that I had been telling my surgeon had been giving me so much grief were removed, from start to finish, in a mere 23 minutes, less time than it takes for me to get a haircut.

"How could that be?  More questions to the surgeon revealed that implants had ruptured and had been somewhat adhered to the chest wall.  None of this was in the original OR reports and came out only after I continued to ask questions.  Could some of that silicone have been left inside of me, I was so desperately ill?

"I owned a thriving hair salon and, no matter how hard I tried, I could not get rid of the terrible pain throughout my body.  I eventually had to close my business, sell everything I had worked for my entire life.  Life, as I knew it, was over.  I was 43-years old.

"In February of 2002, a different plastic surgeon operated on the right breast and found a very large mass containing silicone.  The pathology reports confirmed it was silicone within the scar tissue.  Curiously, I asked about the size of it.  Was it the size of a walnut?  He said, no.

DR. WHALEN:  Ms. Curtis Smith, if you could conclude, please.

MS. CURTIS SMITH:  Thank you.

MS. MORITZ-CIANCUTTI:  Good morning.  My name is Audrey Moritz-Ciancutti.  I have come here from Pittsburgh on my own and I have no conflicts of interest.

I had a bilateral mastectomy at Mayo Clinic due to fibrocystic disease with reconstruction with silicone-gel implants.  My first rupture was two years later and both implants were replaced.  Over the years, I started to become ill.  I had a continuous burning in my chest and developed a very high sensitivity to any kind of heat and sun.  Fourteen years after that rupture, and MRI revealed both implants were ruptured.

The surgeon that removed the implants left some residual tissue behind.  I think he thought he was doing me a favor because I opted not to have an additional implant put in that I would have some sort of breasts.

My rheumatologist at Cleveland Clinic told me that, unfortunately, the surgeon didn't do me any favors, that in that tissue that was left behind were globules of silicone and there is silicone in my lymph nodes.  She felt that this silicone has caused my condition.  The silicone in my body explains why, immediately following the surgery, I began to have severe pain in my right arm and continuous swelling.  I developed a cyst on my hand which had to be removed surgically and, after two biopsies, I was diagnosed with scleroderma.

I continued to suffer with extreme fatigue and depression and my immunologist diagnosed me with atypical-tissue disorder, arthritis, cognitive dysfunction as well as asthma and allergic rhinitis.  I suffer from recurrent bacterial infections and I go through monthly treatments of I.V. gammaglobulin.

My body can no longer tolerate heat or sun and, two-and-a-half years ago, I had to leave Florida and move back to Pittsburgh because I could only tolerate a colder climate.  Today, I am experiencing new scleroderma lesions on my chest, abdomen, my face and the top of my head and, once again, the chronic fatigue is reoccurring.

The bottom line is that I have silicone throughout my body because these implants ruptured and the doctor is sure this has caused my condition and this silicone cannot be removed.  There have never been any studies of the health of women with leaking silicone implants except for the one study conducted by the FDA.  That one study found painful disorders such as fibromyalgia among women with leaking breast implants compared with women with implants that didn't leak.

The results might have been even more dramatic if they had compared to women without implants but we don't know that because no such study has ever been done.  If Inamed really wants to prove their implants are safe for women in the real world, they need to study the health of women with broken gel implants.  Instead, they avoided that by studying women who only had implants for a very few years.

Thank you very much.

MS. TEAGUE:  Ladies and gentlemen, my name is Cynthia Teague and I am a registered nurse from Baton Rouge, Louisiana.  Inamed had sponsored my trip allowing me the opportunity to be here today.  However, I am here because I am an individual recipient of silicone implants.

I would like to address the panel regarding the use of silicone implants and encourage you to allow women to make the choice between silicone and saline implants for themselves.  I would like to offer my personal testimony as to the safety of silicone implants and to their importance in improving my quality of life.

I decided on breast-augmentation surgery due to the postpartal change in my body and after breast feeding my three children.  I was feeling self-conscious, dissatisfied and eager for a change.  My self confidence was suffering and my body image was low.  I sought the consultation of a board-certified plastic surgeon and spoke with him at length about my concerns and desired outcome.

My physician was very forthcoming with the benefits and risks of this surgery.  At this time, women had the choice between silicone and saline implants.  Ladies and gentlemen, I freely chose silicone implants.  That was over thirteen years ago and, since that time, I have had absolutely no complications or regrets about that decision.

My surgical results were excellent and I remain as healthy now as when I had my surgery.  I attribute the remarkable improvement in my self esteem to the augmentation surgery.  I attribute the remarkable surgical result largely, in part, to the use of silicone implants.  In my consultation, I was able to handle both types of implants before making my final decision and I was also shown postoperative pictures of patients with both types of implants.

I chose silicone because of their more natural look and feel.  My excellent surgical results further validated my decision to use the silicone implants.  Having the ability to choose what type of implants were placed in my body, was very important to me.  I felt that I was in control and in charge of my own destiny.

Not all women have that opportunity today.  It is my belief that this choice should be extended to all women.  Being able to make important decisions for oneself is crucial.  Women in American have struggled for decades for this right.  It is my request that the FDA reconsider its position on silicone implants and restore this choice to each individual.

MS. DOWD:  Good morning, ladies and gentlemen of the panel.  My name is Pamela Dowd and I have traveled from Boise, Idaho yesterday to come to you at my own expense because my family felt it was vital that the panel hear from a 28-year survivor of breast cancer.  Today, I stand before you totally breastless and I am proud.

At the age of 27, I became a breast cancer statistic.  I never had chemotherapy.  I never had radiation.  I was never sick before or after the mastectomy.  But, in November, 1980, almost 22 years ago--23 years--I began the road through reconstruction hell with a bilateral breast reconstruction that included a failed latissimus dorsi flap and three ruptured silicone breast implants.  The first replacement was within 90 days.

In June, 1987, when I experienced the third rupture, it sent me to the ER with a pain so excruciating, it felt as if thousands of needles were pulling and pushing through my body and I was literally pulling at my hair and screaming.

After getting breast implants, I developed a constant cough and, at times, I choked on my own saliva.  I have peripheral neuropathy and burning in the spine.  I have IBS and photosensitivity.  I have been diagnosed with Sjogren's and Raynaud's and considered lupus suspect.  I have MS-like symptoms that cause me to fall at the drop of  a hat.  The silicone in my chest and vascular system will never go away.   My chest constantly itches and burns from the residual silicone.  My bones still scream with pain.

In 1995, when I underwent explantation, my chest had to be scraped and cauterized to clean out the residue of the previous ruptured implants.  The American taxpayer will, at some time, be responsible for the cost of my healthcare as they are now responsible for the healthcare of other victims of silicone-gel-filled breast implants.

We are being asked one more time to trust the manufacturers.  Their story is based on 41 years of corporate lies and we have no reason to believe they are telling the truth now.  The manufacturer-funded New York study in the early 1990s didn't ask if anyone developed breast cancer autoimmune disease after breast implants.  It asked if participants had ever heard of anyone who did, proof of nothing more than do people listen to gossip.

In a federal depository in Birmingham, Alabama, lie 80 million documents of implant manufacturers, the ASPRS, FDA and others, telling the truth of this massive corporate coverup of the dangers of silicone-gel breast implants and I am a freelance researcher of those documents.

I have survived breast cancer for 28 years and do not believe I deserve to have a faulty toxic product foisted onto me by corporations without a conscience.  The implant industry has taken away my health but they will not take away my hopes and dreams.  This, too I will survive.

I ask you to please keep the ban on silicone-gel breast implants in place, not for my sake but the sake of generations of women to come who can and will become ill and disabled due to the toxins in the silicone gel.

Thank you for your time.

DR. WHALEN:  If any of the speakers have handouts, could they bring them directly to Dr. Krause rather than to any other panel member.  Thank you.

MS. BORDELON:  Good morning, ladies and gentlemen of the panel.  My name is Lisa Bordelon.  I am a registered nurse and I have practiced in the field of plastic surgery for fourteen years.  I have also been actively involved as a study coordinator with both McGhan, Inamed, and Mentor adjunct studies from the inception.

I have had silicone implants for some 20-plus years also.  I am not here today because Inamed has sponsored my trip nor because of any business affiliation we may have but as an individual recipient of silicone implants.  I am here to share my experience with and opinions of silicone implants.

My decision to have implants initially was not unlike many others.  I simply wanted to be more proportionate for my body and to appear more feminine.  After the birth of my first child, I began to explore the option of breast-augmentation surgery.

At the time of my initial consult in 1982, I was given a choice between silicone and saline implants.  I chose silicone because of their more natural appearance and feel.  The resulting surgery had a tremendous impact on my self esteem.  I was more proportionate.  My clothing fit better and I felt more feminine.  Basically, I felt better about me.

Many patients today do not have the choice of silicone implants for their breast surgery unless they meet certain qualifying criteria.  For those select few that do meet these criteria, they often choose saline because of the prohibitive cost of silicone and because of the negative publicity received in recent years.

As a plastic-surgery nurse, I see many patients with unacceptable results that may have been avoided by the use of silicone implants.  Therefore, many patients are undergoing multiple procedures to achieve a result they may never reach with the use of saline implants.

I cannot say that my experience with silicone has been totally uneventful.  I have experienced encapsulation that has required surgical intervention.  But this is a condition associated with saline implants as well.  This response cannot always be predicted beforehand and makes no difference as to what type of implant is used.

Great advances have been made in the manufacturing of silicone implants in recent years.  I do not believe that the silicone implants are more dangerous than any other medical device but, like any other medical device, there are inherent risks involved.  It then becomes the responsibility of the physician to inform their patients of these risks.

I believe the decision to choose silicone or saline implants should be left up to the individual patient based on physician recommendations as to which result in a better outcome.  For me, the correct choice was to have silicone implants.  I would make this same decision again today.  I believe silicone implants offer a good safe option to those considering augmentation surgery and I hope you will see fit to, once again, give all women this option.  		Thank you.

MS. FORMAN:  Good morning.  I am Jessica Forman.  I am a cancer survivor and I have no relation to any of the implant companies.  I was here about ten years ago when FDA was holding its first round of hearings on silicone breast implants.  I was then a breast-cancer patient scheduling surgery to replace the tissue expanders for permanent implants.  I participated in that lobbying effort to persuade members of Congress and the FDA panel that silicone implants should remain on the market.  I feel very differently today.

I had my first mastectomy in 1990.  I opted for immediate reconstruction and was actually in the hospital for surgery to equalize the normal remaining breast.  It all came to a screeching halt because my oncologist re-read the latest mammogram and found another suspicious spot.  Two weeks later, I had the second mastectomy and still wanted immediate reconstruction.

Together, those two surgeries were a horror story,  the first, for emotional storms, the second for physical problems.  I had so many bad side effects that my doctor said to me, are you sure you are not an M.D.?  This is only supposed to happen to M.D.s.  Well, I am not.  But all those setbacks explain why, two years later, I was still scheduling surgery.

Ten years ago, I didn't care about the future.  I wanted to look and feel as if I had never had breast cancer.  I wanted to wear a bathing suit, wear sleeveless tops.  I wanted to shop at Loehmann's in a group dressing room and not have anyone stare at my scars.  And I wanted it right away.

It didn't happen.  It could not happen.  Some days, I look in the mirror and I hate my scars.  Other days, they are all right.  I can't wear some clothing styles.  They are simply too revealing.  Sometimes, I can't lift my arms over my head.  And I still can't find a bra to fit both parts of me.

I know that the company testifying today provided data showing women with silicone or saline breast implants have a higher rate of complication than augmentation patients and I am upset to hear about that.  I have an older sister who also had bilateral breast cancer.  She got silicone-gel implants.  She is suffering so badly from these connective-tissue immune diseases that she almost cannot work.  She sacrificed everything to be able to work five hours a day.

We don't know.  Is this from her cancer or is this from the silicone implant.  There is no way to find out.  For myself, I know it takes longer to heal.  I know I hurt a lot more and the fatigue factor has gone up tremendously.  I can't get in to work like everyone else in my office.  If I commute downtown for two or three days, I am exhausted.  I am too tired to cook dinner, remember anything, can't solve problems, don't know what I am talking about.  It is a lot easier if I can do this at home and cut out a two-hour commute.

Hindsight is always twenty-twenty.  If I had cancer today, I wouldn't opt for immediate reconstruction.  I am not sure that I would have reconstruction at all.  I would, however, take a long time to investigate my choices.  There may be a better way than implants.  There are tissue-flat reconstructions.  They are not always appropriate.  Given my weight and medical health, my doctor refused to do them.  More power to him.  I think it would have caused more problems.

Recommendation; I think we need a big caution flag waved.  There is a lot of information we don't have about cancer, the immune system and the safety of breast implants.  I would like to see much more long-term research on silicone before your restrictions are lifted.  I would like to make sure that they are long-term safe, not two years, not five years, but ten, fifteen, twenty years.

I am not actually sure that gel implants need to be available since saline is.  But, if they are available, I would ask you to restrict them and to insist on truly informed consent.

Thank you.

MS. FJELSTAD:  (Read by Kathy Sachs.)  Good morning.  I am Kathy Sachs and I reading the story of Diane Fjelstad who is too sick to be here today.

"In 1978, I was encouraged by a team of doctors, including an obstetrician, a plastic surgeon and radiologist to have bilateral mastectomies to prevent breast cancer.  I believed these doctors when they told me that the implants were perfectly safe, would last as lifetime and I would go to the grave looking better than most women.

"Although I had early warning signs, I didn't suspect the implants.  There were unexplained blood disorders, a strange rubbery substance oozed through the skin of my foot, and a rash of new allergies developed which did not respond to treatment.  I could not understand my continual fatigue in the late '80s and '90s.  I was unable to drive to the next town, only fifteen miles away, without having to stop for a nap.

"This was very foreign to me as I had been a competitive athlete.  My skin began to hurt.  Worst of all, I feared I couldn't continue to work as a high-school teacher due to my memory lapses, fatigue and continual illnesses.  In '95, I noticed a dyslexia-like problem with words and numbers.  Finally, I noticed that the left implant had seriously shifted and was pressing against my lungs.  Breathing became difficult.

"I attended a conference on silicone exposure and discovered that many of my symptoms were classic; chronic inflammation, myofascial-pain syndrome, swollen glands, numbness and tingling, and, later, severe pain, early signs of Sjogren's syndrome, chronic-fatigue syndrome, muscle atrophy, low-grade fever, severe headaches, chronic sinus infections, lung problems and gastrointestinal problems.

"My condition deteriorated very quickly in the fall of '95.  I was explanted in late January, '96.  The explanting surgeon was furious that I had been allowed to get so sick and that the doctors had not picked up on my symptoms and treated them earlier.

"The simple explanting surgery turned out to take over five hours with black substances and school-bus yellow oil running all over the operating table onto the floor.  What a stinking mess.  The implants were filled with fungus and no sign of silicone gel remained in the ruptured shells.

"Since that time, I had a routine bunionectomy and asked the operating physician to do a biopsy testing for silicone.  He was totally shocked to discover that there were, indeed, silicate crystals in the bunion.  I also know that I have lesions on the brain from the silicone as revealed in an MRI.  Who knows where else the silicone has settled since the body has few ways to rid itself of it naturally.

"My life has been disastrously impacted by silicone-gel implants.  I have no breasts.  I was no longer able to work as a high-school teacher.  I had to go on disability.  My husband of 40 years divorced me.  My financial income has been reduced to one-fourth of what it was.  My health has been compromised in many ways as I have mentioned.

"You will hear from some women who insist on being given a choice.  Don't do us any favors by giving us the choice of an implant whose safety is unknown after just a few years.

"Thank you."

DR. VASAY:  (Read by Dr. Diana Zuckerman.)  Hello.  I am Dr. Diana Zuckerman.  I am going to be testifying on behalf of myself tomorrow.  Today, I am just reading the testimony of Dr. Frank Vasay who asked me to briefly describe his research.  Dr. Vasay is Director of the Division of Rheumatology at the University of South Florida College of Medicine.

Dr. Vasay did provide some testimony ahead of time.  I hope you have it.  He has personally evaluated approximately 2,000 symptomatic women, most of whom have silicone-gel-filled breast implants.  He has been doing this work for over twenty years.

He wants me to say for him that he has been satisfied that some women's immune system does not tolerate silicone breast implants.  The syndrome remains undefined but recent information suggests that 10 to 25 percent of women are so ill approximately thirteen years after getting implants and that recent epidemiologic studies show statistically increased symptoms in women with breast implants compared to control groups.

He has done some research and he asked me to describe it for him.  This is published in various places.  I had hoped his coauthor could do this for him, but she works at the National Cancer Institute and I guess that is a conflict of interest.

He studied approximately 100 women who had rheumatologic symptoms as diagnosed by rheumatologists and who had silicone breast implants.  Approximately half of them chose to have their implants taken out as was suggested to them and the other half, although it was suggested to them, and the other half, although it was suggested, decided to keep their implants.  So that was his study group, obviously not randomly controlled.

He followed them up for an average of 18 months.  What he found was a remarkable number of women who had their implants taken out and not replaced showed slow steady improvement.  The symptoms that they had were similar to fibromyalgia; muscle pain, chronic fatigue with multiple trigger points.  He points out that, unlike classic fibromyalgia, exercise aggravated the muscle pain.  So that was different from what one would normally see in fibromyalgia.

He also found that some women got better and some did not, not surprisingly, but that most of the women who had their implants taken out did get better and most of the women who did not have their implants out got worse.  By "most," that was over 90 percent.

There was a way to at least predict who would get better and who would not.  Women who tended to have a bad prognosis and ruptured implants with silicone debris in their chest wall so that the silicone had not been entirely removed and could not be removed.  He also found that women who had delayed removal where their implant had ruptured some years before but they hadn't gotten it out for quite a while also had a poor prognosis.

He also found that the women who didn't do so well after getting their implants taken out tended to have an onset of problems very soon after getting their breast implants, usually within two years.

He has, as I said, published some of this work and it will be made available to the panel.

Thank you.

MS. GROSS:  Good morning.  My name is Marcy Gross.  I have six minutes having combined time with another potential participant.  I am a health-policy consultant and a member of the State of Maryland's Women's Health Promotion Council.  I retired in November of last year after 25 years at HHS most recently as the Senior Advisor for Women's Health at the Agency for Healthcare Research and Quality.

During part of my tenure at HHS, I also participated in the Secretary's ad hoc working group on silicone breast implants.  I give you this brief resume just to establish my familiarity with the issues you are considering today but I am here as a private citizen and I base my remarks on my personal views and on information available to the general public.  I am not paid.

One legacy of my six-year tenure at ARC is a good appreciation of the need for women and policy makers to have a strong evidence base for making decisions on healthcare issues.  My concern today is that an adequate evidence base for the premarket approval of silicone-gel-filled breast prostheses does not exist and that the short-term assurances drawn from the applicant's study will override continuing gaps in the research on long-term risks.

I can see from the information posted on the panel's website for the hearing that the FDA has done exhaustive background work including literature reviews, assessments of toxicology and chemical testing and extensive consideration of the clinical data.  All this and more will come to you over the next two days.

But some facts are simple and have not varied in the twelve years since silicone breast implants were pulled from the market.  The studies available on the health aspects of silicone-gel implants are short-term analyses often involving mere handfuls of mice studied over the course of a few weeks in the case of toxicology and chemical testing and over a two-to-three-year period in the case of Inamed's core study being cited in this proceeding in support of a case for reintroducing silicone breast implants for unrestricted use.

Yet, even these short-term assessments indicate that the rate of ruptures and other complications which lead to a reoperation remains unacceptably high, 20 percent plus for patients using implants to augment their figure and 45 percent plus for reconstruction patients.

Thanks to our longer life spans, breast implants will stay in a woman's body for the rest of her life which can be 50 years or more.  A two- or three-year study simply doesn't offer the kind of assurance of safety needed.  It is a setup for future medical problems for women who have this type of device implanted, especially when retrospective studies show that, by ten years, most women with silicone-gel implants will experience at least one broken implant.  At least these women know they have a problem and have further surgery.  Silent ruptures are also a documented and recurring problem.

Sometime in the next few days, I feel certain you will hear someone say that none of the several important studies of the possible risks of implants have shown that they cause long-term health problems.  You may even hear it said that, well, we use silicone in other implanted devices such as heart valves, knee joints, hip replacements, so why should silicone from breast implants be different.

My response is that having bubbles or globules of migrating gel floating through a human body into and around organ is, on its face, a health risk.  As for silicone and other implanted devices, one doesn't expect bits of a heart valve to break off and migrate to the brain, liver, other organs.

Further, patients who get hip or knee joint replacements are routinely warned that a small number of patients will have an immune-system reaction to the silicone fragments shed by newly implanted joints.  How is it that orthopedic surgeons know this problem but the breast-implant community continues to doubt.

There is no health imperative behind the push to reintroduce silicone breast implants.  On the contrary, they are used in elective cosmetic procedures that often cause serious health complications associated with ruptures and the surgery, itself.  Although there are other options available to breast-cancer patients that carry fewer risks, for some, silicone-gel implants is felt as important to their recovery.

We know that this choice will make them more vulnerable to future illnesses and complications and will obscure future cancer screening procedures.  But, again, we lack the kind of long-term studies that would allow their choice to be fully informed.

This FDA panel will be making a decision that affects the lives and pocketbooks of women quite substantially.  There are emotional issues involved in the decision and results.  Implants are an expensive procedure especially when the cost of care for complications, reoperations, infections and other medical side effects are added in.

About 250,000 new procedures were done last year at a cost estimated at approximately 5,000 each on average for a total cost of $1.2 billion to the healthcare system.  While much of this is paid for by the patient, the cost for reconstructive procedures that many cancer patients elect to have is most often covered by insurance and, of course, follow-up care also would be covered most insurance.

What is the policy recommendation here?  The FDA found the middle ground twelve years ago.  While I would still prefer to not have seen the fourfold increase in implants that has since occurred, at least the current policy sends a strong cautionary signal to women with choices.  This should not change until we have more confidence that these implants are safe for long-term use and that will take time and more research, the kind of research that should have been initiated ten years ago.  Let's not miss another opportunity.

I thank you for your time and attention this morning.  I see that you have an agenda which runs well into this evening so I wish you luck and stamina.  Thank you very much.

DR. MASSEY:  Good morning.  I am Dr. Marga Massey.  I am an assistant professor of surgery and the Director of Basic and Clinical Science Research in the Division of Plastic Surgery at the University of Utah Hospitals and Clinics in Salt Lake City, Utah.  I have no affiliations or conflicts of interest with any implant device manufacturer and have paid my own travel expenses to testify at the hearings today.

I use breast implants in performing both cosmetic breast-augmentation and breast-reconstructive surgery.  I am on a fixed salary as a University of Utah paid employee and receive no financial gain whether I perform breast-augmentation or reconstruction procedures utilizing breast implants.

As an academic plastic surgeon involved in basic science research which occupies greater than 50 percent of my current time, I have restricted the majority of my clinical-practice interest to breast patients, both cosmetic and reconstructive.  I have been impressed that cosmetic patients are very informed about their current surgical options to enhance the physical appearance of their breasts.  This level of knowledge likely reflects the current availability of information regarding cosmetic breast augmentation on the Internet and through the mass-media coverage of this topic.

Breast-cancer patients additionally master a very complicated variety of breast reconstructive options to include but not limited to tissue expansion with delayed breast-implant placement, immediate breast-implant placement in the setting of skin-sparing mastectomy and the use of autologous muscle flaps which may be augmented by the placement of the incorporation of a breast implant in order to restore the appearance of a natural breast after a mastectomy.

Many speakers today will provide testimony regarding the safety of silicone-gel-filled breast implants.  My personal review of the published information to date regarding the safety and efficacy of silicone-gel-filled breast implants supports their availability to cosmetic patients for breast enhancement.  This availability would mirror that which is currently acceptable for breast-reconstruction patients.

As a plastic surgeon practicing in an academic setting, along with my colleagues here today, we eagerly await review of this most recent evidenced-based investigation.  We wish to continue collaborative scientific clinical investigations with the FDA and the device manufacturers to ensure patient satisfaction and safety.

When the final review of this most recent evidence-based investigation becomes publicly available, I believe women will be empowered to make informed choices regarding the use of silicone-gel-filled breast implants.

In conclusion, breast-implant device availability in clinical use should be based on sound scientific investigation.  Public education should be based on sound scientific investigation.  Implant-device availability and public education based on sound scientific investigation will ensure patient safety and will allow patients to make informed personal decisions regarding the use of silicone-gel-filled breast implants for cosmetic enhancement and for reconstructive restoration after mastectomy.

I greatly appreciate the panel's time and consideration.  Thank you.

DR. WHALEN:  Dr. Massey, would you entertain just one question?  I appreciate your full discloser at the beginning.

DR. MASSEY:  Yes, sir.

DR. WHALEN:  But is it your contention, in your academic environment, that there is absolutely no relationship between any volume of procedures you do and present and future income?

DR. MASSEY:  I beg your pardon?

DR. WHALEN:  I am in an academic setting, as you are.

DR. MASSEY:  Yes, sir.

DR. WHALEN:  And certainly the number of procedures that I do will impact, at a divisional and departmental level, a revenue structure that does impact upon the income of I and my fellow faculty members.  You said you are on a fixed income.

DR. MASSEY:  I am on an absolute fixed income that has no reflection on the number of surgical procedures that I perform.

DR. WHALEN:  So your department is able to remunerate you even if you do zero procedures or you do 200?

DR. MASSEY:  That's correct, sir.

DR. WHALEN:  Thank you.

DR. MASSEY:  Thank you.

MS. WOLF:  My name is Carolyn Wolf.  I live in Virginia.  I came here at my own expense.  I have no conflicts of interest.

In 1971, I had a double mastectomy and was reconstructed with what were then considered new and improved silicone implants.  That was the time when breast implants were much thicker than they are today and were a rarity.  Few women have had implants as long as I did.  There are no published studies focused on women with implants for longer than 25 years because so few implantees have lived so long.

I would like tell you what happened to me.  At first, I had no health problems and, even when they started, I had no idea they were related to the implants.  But, after nine or ten years, I developed dozens of hard, burning blister-like growths on my neck and boils on my forehead.  I have pain and numbness in my left shoulder, arm, hand and foot.

By 1993, my family had noticed a distinct change of personality and I noticed cognitive changes.  I am diagnosed with rheumatoid arthritis, Raynaud's, irritable-bowel syndrome, silicone-induced MS-like syndrome with neuropathy of the extremities.  I have neurological damage to the left eye.  My silicone level is double normal.  This is two years after explantation.

My brain MRI reveals more than 20 lesions.  In March, 2000, a batch of silicone moved from my left breast into my armpit leaving an elongated swelling in its path with excruciating pain.  On April 13, 2000, I lost vision in my left eye for 45 minutes.  I had constant pain in that eye for six weeks but, after a long, thin greasy glob came out of it, the pain lessened.

I have had two similar strings come out of my left ear in the last year.  That is 30 years after implantation.  Although I did not smoke, I was coughing up hard gold-colored globs and a lung X-ray showed chronic obstructive-pulmonary disease.  Although the type of silicone implants that I had were thicker than those of today, MRI showed both implants ruptured.  I went to three different plastic surgeons but they only wanted to install replacements.

When my implants were removed three years ago, both implants were extensively ruptured, both much smaller than when they were implanted.  The scar capsule surrounding the implants had cysts filled with chronic inflammation and foam cells containing silicone materials.

Chest X-rays show that I still have silicone in the breast area.  I am still exuding silicone from the nipples and have deposits of silicone in both armpits.  Researchers at the National Cancer Institute have conducted the only studies of women who had implants for at least eight years.  They found that women with implants were more likely to die from certain cancers and suicide.

The FDA conducted the only study of women with ruptured silicone implants and they found they were more likely to have fibromyalgia.  If implants are so safe, make sure that the implant makers prove it.  The studies you will hear about today do not do that.  Please remember that, whatever you do, if implants are approved by the FDA, you cannot control the advice given by plastic surgeons many of whom are still claiming that breast implants are perfectly safe and continue to insist that broken implants should be replaced.

We beg you, please protect the younger generation.  Thank you.

DR. WHALEN:  Thank you very much.  I think that is an appropriate time that we will break for about thirteen minutes until 10:30.  I need to remind the panel members, as difficult as it is, that any discussions upon the topic at hand need to be public, in front of the microphone, so refrain, outside of this panel room, from discussing anything at hand.

[Break.]

DR. KRAUSE:  Let's continue, please.  Thank you.

DR. WHALEN:  To try to keep on time, we are going to resume.  Actually, the next public presenter is going to be via a conference call from the Hofnagel Group.  They will have five minutes.  Probably there will be people filtering back in because apparently the rest rooms were engineered by the same people who did the elevators.  So we will have to have some forbearance for that.

So the microphone is already rigged towards the speakerphone and we will initiate with the Hofnagel Group's discussion.

DR. HOFNAGEL:  Hello.  We are speaking to the committee.  We had difficulty hearing before.  My name is Dr. Vicki Hofnagel.  I have a group called the Hofnagel Institute.  We came before the FDA previously on multiple issues.  In 1994, I presented documents to the FDA showing that saline silicone breast implants still had migratory silicone into the chest wall of women.

At that point of the purpose of that discussion and conference was to discuss the fact that surgeons who were removing implants were replacing implants and not taking tissue to study it from migratory silicone.

At this time, we supported the decision by David Kessler because he was looking at life-long chronic complications not life-threatening complications.  The implants that are being presented have not changed in their design, the manner in which the protocol for placement or the protocol for follow up has occurred.

In this hearing, the availability to give public testimony is wonderful as it is in all hearings.  However, that public testimony is usually not taken into consideration whatsoever as seen in the May, 2000 uterine-artery-embolization discussion before the FDA.

We will be petitioning to the FDA to request that all FDA hearings will now be made public with public-television and also on the Internet.  It is extremely important that this meeting be made known to women across the nation and throughout the world.

The things that prompted the Dow Corning case are still existing today.  I still see women every day with silicone breast complications.  Virginia Gallagher will be speaking to that, the crystallization of silicone causing glass to cut into the chest wall, the implantation of migratory silicone onto the nerves, and things that have not been made public such as the common condition of chest-wall deformity from all breast implants being used today and that this chest-wall deformity is the cause of many of the complaints of pain, itching and heat discomfort that women have.

This has not been known to the public.  You can take chest X-rays and see the actual rib deformations in all of these women.  Yet it is not a baseline study.

There is much to be done.  We need a national registry that will actually take in all of the complaints and look at them from the past if not looking on to the future.  C-reactive protein, ANA, LE prep, complement fix, all of the routine studies still are not part of a strong protocol.

We need to move forward and we should have seen a new design.  We have technology to increase tensile strength.  We have technology to encapsulate even a silicone gel.  But the technology has not been shown in the new implants.  Instead, we are given an implant now with a two-year shelf life.  That two-year shelf life and the aspect of the life of an implant is still not given to patients.  Informed consent and complete disclosure is absolutely a sham still today.  We need a much better educational program for women.

If we are going to move forward and reallow silicone to be used, then we need to change the design of the implant, make sure that it has radiopaque material to see when it leaks because so many women have no idea that they have leaks, by they have symptomatology and they are basically told, honey, go home.  It is a little ache and pain.  You are crazy in the head.  And then, when we operate on them, we find that their entire chest wall has been destroyed.

I am going to move on to Virginia Gallagher who is going to speak.  We had two other women who wanted to speak.  Those women are young and beautiful and, when they went to their physicians complaining of itching and burning and other difficulties, they were patted on the head because their breasts looked great.  They looked great.  These women were attractive.  They were scared to death to have their implants removed because of chest-wall deformity.

Our institute notified the FDA in 1994 and we have continued.  We now have developed a plasma-state surgery using argon beam and liquids to actually prevent chest-wall deformity and the public is unaware of this.

In the panel and committee, you should also have psychotherapists because of the reasons that women have implants and the reasons they do not present afterwards such as the two young women within our group and we will be forwarding documents.

I conclude because of the time limitations.  We will be sending a formal document because we cannot put all the data in such a brief discussion--and move on to Virginia Gallagher.  We will close with Denise Bennett who is the attorney and counsel for the Hofnagel Institute on illegal issues.

Thank you.

MS. GALLAGHER:  Hello.  My name is Virginia.  I can't tell you in this short time the multiple problems that I developed, but I suffered many health problems and, within two to three days of surgery of having the implants put in, my back and chest area was completely erupted with red sores that left multiple scars all over and remain to this day.

My implants became hard and lumpy, not only painful but socially embarrassing.  Can you imagine what it is like to meet an old friend at your alumni and have them give you a hug, jerk back and look at you in shock at your chest because they have just hit a rock wall.  That happened to me enough times so that I began to withdraw.  It was embarrassing and humiliating.

My lymph glands were very sore and felt like they burning.  Within one or two months of the time my implants were put in, it was almost impossible--well, was impossible--to maintain totally normal range of motion with my arms.  It was impossible, also, to find a comfortable

DR. KRAUSE:  Excuse me.  Could you wrap up, please.  We need to go on.

MS. GALLAGHER:  It was uncomfortable to find a position to sleep in.  I asked the surgeon to remove them. He said, no, I would be psychologically damaged and that it would be an irreversible depression.  The social situations were humiliating but the pain was terrible.  My left breast grew to protrude from my body several inches more than the right breast and was inflamed and red.

I was told by another surgeon that the implants didn't need removal, although they could be replaced with new ones.  At the time of my removal, the burning lymph glands were free of pain when I came out from under anesthetic for the first time in 26 years.

I am outraged that doctors have no concern for our health, that they deny the health problems that results, that insurance doesn't cover the removal of the implants and for teaching--(operator interruption)--is inorganic.  It is not natural unless it is already there normally.

I was told that these implants--I was not told that they were not a lifetime procedure, that they would need to be replaced at any time in the future.  When I got them.

MS. BENNETT:  This is Denise Bennett wrapping up.  For a long time, manufacturers and regulatory--

DR. WHALEN:  I am afraid that we need to move on because you are so far beyond the allocated time and I apologize for cutting you off.

Let me just say to the rest of the people in the room as well, we try very hard to balance between the number of people who wish to speak to us and to try to allocate the time appropriately.  So we are not trying to efface--ignoring what people are saying by cutting people off but, rather, trying to balance that with justice towards everybody who wishes to be heard.  Indeed, there are many people who do wish to be heard.

On the contrary, I have to interject at this point to what I have heard now at least twice this morning.  Having chaired the panel where we looked at saline breast implants, I can totally assure you that I, and our panel members, do not ignore what the public is saying.  We very acutely listen to what the public is saying and we take that into account.

The next speaker, please.

DR. JEWEL:  Thank you.  My name is Dr. Mark Jewel.  I am a board-certified plastic surgeon with over 25 years of experience with breast-implant surgery.  I have no financial ties to industry or health-professional societies.  I have been a clinical investigator for Inamed's core gel study and an investigator for Inamed in Mentor's cohesive studies.

I am Vice President of the American Society for Aesthetic Plastic Surgery which will reimburse my expenses for travel and hotel.  Today, I am speaking as an individual.  I thank the panel for this opportunity to address the topic of patient informed consent and the management of expectations.

I have a particular interest in this subject having researched and written the Patient Consultation Resource Book which is used by more than 40 percent of all board-certified plastic surgeons.

With the Internet patients who see more information today than at any time in the past regarding breast-implant surgery, many arrive at the physician's office with a great deal of information including pictures of results from the websites, yet have an incomplete understanding of what will actually be their personal result of after breast-implant surgery.

I believe we need to move beyond traditional approaches of informed consent to empower our patients to be actively involved in all aspects of the clinical decisions which precede breast-implant surgery.  We must effectively educate patients regarding potential problems that may occur subsequently.  In this way, the patient can better understand surgical risk, expected outcome, potential dissatisfaction and financial responsibilities of breast-implant surgery.

Because the patient takes an active role in the entire process, she and whoever else she chooses to involve in her decision, are invested with a degree of responsibility and control over the outcome.  Physicians have discovered that, even though patients may sign a consent form, that does not necessarily mean they will in the future remember being informed about potential complications or the possibility of dissatisfaction with results.

Research into effective methods of achieving informed consent is ongoing and plastic surgeons have funded such research to help determine the extent of informed consent in breast-augmentation procedures.  Informed consent means that the patient must be educated as to surgical alternatives and the probability of any unwanted outcome.  Risk must be quantified in ways that patients understand and communicating risk effectively is a complex task.

Patients must understand the limitations of surgery and the process involved with recovery.  Patients need to understand this involves the potential for complications of the surgical procedure, itself, inherent risk of implanted devices and additional advisory information pertinent to the use of breast implants.

It is essential for patients to understand that breast implants may need to be replaced over time.  I have been involved in developing materials for breast-implant patients that I believe successfully meet all the criteria that I have discussed today.  The FDA's ongoing regulatory process will ensure that sound science will determine what choices may be available for women going forward.

Plastic surgeons welcome the current hearings which will determine whether silicone-gel breast implants can be available for general use.

Thank you.

MS. RICHARDS:  Dr. Krause, and members of the panel.  My name is Mary Ann Richards.  My transportation costs and hotel accommodations have been paid for by Inamed.  I have not been offered, nor would I accept, compensation for the time I have taken from my employment to present with you today.

I am a four-year cancer survivor who chose to have a silicone implant after a mastectomy.  Between the time it was determined I would need a mastectomy and actual surgery, itself, I did considerable research on the various types of breast construction available to me.  After extensive research, I felt an implant would be my best option.

The initial expander was put into place following my surgery on January 25, 2000, by Dr. Caroline Glicksman, an excellent board-certified plastic surgeon.  It was a saline expander and, as it was filled with the saline solution, I found it extremely hard and uncomfortable, a constant reminder of my disease and loss.

As part of a cancer support group, I knew several other women who had opted for the saline implants and they, too, found them hard, uncomfortable and disappointing.  When I mentioned this to Dr. Glicksman, she said that I might be a candidate for a silicone implant.  She gave me some literature and I did extensive research.  After I did my homework, I felt that the silicone implant was the way I wished to proceed.

I discussed this at great length with Dr. Glicksman and she gave me additional literature with both the advantages and disadvantages of silicone implants.  I decided to proceed and, on May 26, 2000, I had a McGhan silicone breast implant inserted.

Losing a breast because of cancer is not easy, but facing daily the reality that the cancer might reoccur was, for me, far more difficult.  With my silicone implant, I am not constantly reminded of that possibility.  My silicone implant is comfortable, soft, and feels very natural, unlike my original saline expander.

I feel I am a responsible, informed consumer.  I see Dr. Glicksman yearly.  I am aware that silicone implant have a limited life expectancy, much as a pacemaker or artificial knee or hip and will need to be replaced in the future, much as they would.

I have never, for one minute, regretted my decision with regard to my silicone implant and feel very strongly that other women who have to face the horror of breast cancer and the mastectomy should have the same options available to them that I have had.  It is for this reason that I am willing to present before your committee about this very personal matter.

Thank you for the opportunity to speak before you.

MS. BANCARZ:  Good morning.  I'm Lisa Bancarz, a 1982 graduate from Pacific Union College and resident of California who had silicone-gel-implant surgery in 1987.  Following two pregnancies and breast-feeding both of my children, I found my breasts were asymmetrical.  I was not happy with this change in my appearance.

I decided to have breast-augmentation surgery and scheduled an appointment with a board-certified plastic surgeon.  My surgeon discussed at length the pros and cons of the surgery.  I spoke to several of my friends and acquaintances who already had this surgery.  Some mentioned they were slightly more firm because of capsular contracture but they were happy and did not want them removed.

I decided to proceed with the surgery and have never regretted it.  I remember several articles and media programs regarding the issue of silicone and breast implants from the early 1990s.  I spoke to my surgeon and he talked to me at great length regarding the silicone issues.  He explained that, since I was not having problems with the implants, I should follow the recommendation of the American Society of Plastic and Reconstructive Surgeons and not have them removed.

I obtained more information from articles given to me by my physician and what I personally read in the newspaper and magazines.  It became evident, through my readings, that silicone had been, and still is being, used successfully for other types of implant procedures.  I am offended that a woman's option to choose a silicone-gel implant was taken away.

I feel comfortable with my decision to have silicone-gel implants and have never regretted it.  Because of the controversy surrounding the silicone implants, I have attempted to keep up the studies and their results.  Again, I feel I made the correct decision.

I have several friends who have had the surgery.  After seeing their results with saline implants, I am thrilled to have my silicone-gel implants.  My breasts are soft and feel like natural breast tissue.  I am extremely satisfied with the surgery and would not hesitate to have the procedure again.

I would strongly recommend silicone-gel implants to friends and, yes, even my daughter if she expressed an interest at an acceptable age.  I am very appreciative and glad that I was given the opportunity to choose silicone-gel implants.  It is my strong belief that every woman, after consultation with her physician, should be able to determine which type of breast implants are best for her.  Silicone-gel implants should be an option.

Thank you.

MS. WRIGHT:  Good morning.  My name is Eileen Wright and I want to thank the panel for allowing me to comment on my nearly 20-year experience with silicone breast implants.  I currently work as a nurse manager within the Johns Hopkins Health System.  I do not have any financial interest in the manufacturer or its competitors.

I am a breast-cancer survivor.  After my diagnosis and subsequent mastectomy, I chose to have breast-implant surgery.  At that time, I was under the mistaken impression that my implants would last forever.  I lost touch with the surgeon who performed my implant surgery very shortly after the procedure as he left the area to practice in another state.  I truly felt any further follow up by a surgeon would be unnecessary.

It was not until nearly 16 years later that I noticed, one, my implants had appeared to rupture; two, I discovered implants were not designed or meant to last forever.  It was only through a conversation I had with a colleague who also had silicone implants that I learned implants, like any other medical device, may need to be replaced periodically.

Needless to say, when I visited a local plastic surgeon in 2000, he advised me that the implants I had for 16 years needed to come out.  He also told me about the various surgical options available to me after the implants were removed.  As I look back on this process, it is clear to me that the science, along with the patient-education process, has evolved considerably since my first implant surgery in 1984.

My plastic surgeon took the time to explain the range of reconstruction options available to me.  He noted that I did not need to replace my implants with silicone-gel implants if I did not want to.  He showed me, and allowed me, to handle examples of both saline and silicone implants and he detailed an exhaustive list of potential complications associated with the devices.

When I expressed an interest in having the old gel implants replaced with new ones, he explained that rupture could occur again.  He also confirmed that I understood that the implants might still need to be replaced at some point in the future.  Until the time of the implants' rupture, I had been very satisfied with the outcome of my initial surgery.  Thus, I felt very comfortable replacing the old implants with new ones.

It has been three years since my most recent implant surgery.  I am very satisfied with the result and now consider myself a well-informed patient.  It does not make sense to me that silicone-gel implants are only available to women for breast reconstruction.

I think it is important that all women have the option of silicone implants available to them.  I believe they provide a look and feel that is most natural.

Thank you for your time.

MS. LOWENSTEIN:  Good morning.  I am Lisa Lowenstein from Fremont, California.  Inamed paid for my travel and accommodation to attend today's meeting.

After breast-feeding and raising four children, I was not happy with the way my breasts looked and the way my clothes fit.  I decided that I wanted to have breast augmentation to restore the look and the feel of my breasts.  In 1997, I researched all the implants that were available to me.  In August of that year, I met with my plastic surgeon and he informed me that my only choice was saline-filled implants.

I was not happy with that choice due to a number of reported rupture incidents and an unnatural look that it had a history of.  My next step was to have a mammogram done before the procedure.  Within 30 minutes of having the mammogram done, I received a call from my doctor stating that he wanted me to retake the mammogram and that he wanted me to see him at his office.

During that appointment, my doctor explained to me that what he and the X-ray technician found were decalcified starburst spots on both of my breasts.  Considering the fact that my mother died from breast cancer, this was alarming to both me and my doctor.

After discussing my options and the procedures, we decided it would be best to do a bilateral subcutaneous mastectomy, something that is rarely performed these days.  The surgery took place in September of 1997.  I was so scared of the possibility of having cancer.  My doctor was quite pleased when the surgery was over and the tissue expanders were put in for healing time.  The biopsy report came back negative.  What a relief.

My doctor knew that the silicone-gel implants would be out for testing and maybe in time for me to have them.  He explained to me that if I was his wife, that he would want her to have the silicone-gel implants in this situation because they were, in his opinion, much more natural, softer, tier-drop shaped like real breasts, and feel like real breasts.

He believed that there was less chance of leaking or rupturing and I trusted his opinion.  In May of 1998, after qualifying to be a case study, I underwent my final surgery to insert the silicone-gel implants.  My doctor informed me that these would be so much nicer than the tissue expanders and that, in time, would soften and drop into position like real breasts.

I was ecstatic when the surgery was over.  I hardly had any pain and my breasts looked beautiful.  They were perfectly shaped, soft to touch and looked great in clothes.  I never had any side effects or problems of any kinds with these implants and, as time went by, they softened even more and looked more natural than my girlfriends who had saline implants.

Even though I had all my breast tissue removed, my implants looked much more natural with no wrinkles and no hardness.  I am a very athletic women and work out at a health club very strenuously.  I wanted to make sure that these implants could keep up with and not limit me in any way.

Nothing limits me in my life by having these implants.  I forget that I have implants most of the time and no one knows that these are not my own breasts and I never give my secret away.  The best part for me is that I am healthy and these implants look great.  I love being able to find clothes now that I fill out the top perfectly and evenly.  It is important to me that I am proportional.

The worry about breast cancer is erased from my mind.  I would make exactly the same decision today if I had to.  I wouldn't change one thing about my experience.

In the dental field, which is my background, we let the patients choose what kind of filling to put in their mouth, whether it is a lesser material, veenamalgam*, one-color silver, unnatural, or higher-quality materials such as a tooth-colored composite that is more natural and unnoticeable.  Based upon my experiences, patients that care about their appearance will choose the higher-quality composite because it is more natural and aesthetically nicer.

The key word here is choice.  Women should have a choice of what kind of implant they want in their body from a list of all products deemed safe and effective.

DR. WHALEN:  Can you conclude, please, Ms. Lowenstein.

MS. LOWENSTEIN:  At the end of the day, I don't consider these an implant but part of my body.  Thank you for your consideration.

MS. NYE:  Good morning.  My name is Kathleen V.F. Nye.  I rode in from Pennsylvania with friends and I have conflict of interest.

I know from experience that silicone-gel implants will adversely affect the lives of many people if they are approved based on just a few years of safety data.  Not only are the women who are implanted at risk but, also, there is a trickle-down effect affecting the whole family.

Inamed says that their new and improved silicone breast implants are safe.  I would like to relate my story of new and improved silicone breast implants.  My first experience was in 1968.  I was 22-years old and had a bilateral mastectomy.  These new devices started to get hard in about six months.  I also started having fatigue and joint pain and, in '76, I found a doctor who would remove them.  The new ones also hardened and were removed in '81.

At this time, the doctor told me there was a new and improved version and they would not get hard.  I had the new and improved inserted in '81.  This new and improved silicone implant started to harden again in about six months.  After three years of much pain, I went to a plastic surgeon and told him just to remove the implants.  I did not want another set.

But, again, I was told about a new and improved silicone breast implant.   Again, the new and improved implant started to get hard in about six months.  I started to feel lumps around the edge of the implant.  There were three masses.  The cancer was sandwiched between two masses of foreign material with giant cells.

This started me on the road of more implant disasters.  I had a total of 13 implants including devices made by McGhan, a company now called Inamed.  In the early '90s, the implant came through the skin and popped out on its own.  I have a photo that I included in my presentation.

I am sure you understand my concern when I heard the manufacturers are now trying to promote a new and improved silicone breast implant.  Talk is cheap.  It is easy to call an implant new and improved, but you won't know if it is really better until it has been inside a woman's body for five, ten or fifteen years.

I urge you not to be as naive as I was when I was told that they are better than the old ones.  Please do not put these new and improved implants, silicone-gel implants, back on the market unless they have been tested for a long enough time to really prove they are safe.

Thank you.

MS. BUIKEMA:  (Read by Kirstin Hoskison.)  Hi.  My name is Kirstin Hoskison and I am reading for Betty Buikema.  	"Every day has been a silicone day in my life for the last ten years.  In 1993, I became a support-group leader with two other special women, Kay Dlugopolski and Peggy Pardo.  We had the same thing in common, silicone breast implants.  One was for reconstruction after cancer, one to keep from having cancer because of recurring cysts, and one for cosmetic reasons.

"We each became ill, have problems with the implants, had to have them removed and be reconstructed with our own tissue.  Our medical records and history are similar to those we have heard over and over from so many women.  Because we believed in the cause of fighting what happened to us, we began, at our own expense, to inform and educate ourselves and as many women as possible.

"We have had the opportunity to meet and talk to hundred, perhaps thousands, of women across America and the world.  What we all had in common were silicone breast implants and the illness caused by them.  That many women couldn't be a coincidence.

"When I heard that there are now three studies showing that women with breast implants are three times as likely to commit suicide compared to other women, I wasn't surprised.  We knew women who have committed suicide, something I, myself, had considered when the pain was unbearable, when it doesn't seem like the pain will ever go away, and doesn't seem like you ever get better and, when you have mutilated breasts and you are broke, in addition to the pain, it can be hard to feel that life is worth living.

"The cost of having implants inserted is nothing compared to the cost of having them removed, not just in dollars but in health.  This has cost me all of my savings for retirement.  At a time when I should be enjoying life, all I can do is pay for pain-killing medications to survive.  	"My own story started when I was 40 and had a double mastectomy because of chronic cystic mastitis and reconstruction with silicone implants.  I immediately had a hematoma.  A year and a half later, I had a small tumor near my rib cage which was removed a few months later.  By 1990, my left breast was deformed, probably due to rupture.  My implants were removed in 1992, but I was diagnosed with Sjogren's syndrome and later with atypical connective-tissue disease.

"Are Inamed's current silicone breast implants substantially different from implants I had.  Will they leak or rupture, cause contractures?  Will some women be allergic to silicone?  Can you guarantee they will be safe this time?  Are women going to be guinea pigs again?  You know that silicone implants leak as they age.  Why would anybody think a two- or three-year study is long enough when everyone knows that most women's health problems don't begin until seven to ten years after they get implants.

"I am not here to tell you what to do.  I am here to beg you to base your decision on whether you would want your 17- or 18-year-old daughter to get these poorly tested implants.  Remember that if silicone-gel implants are approved, teenagers will be first to line up.

"Thank you for listening to the women, but listen with an open mind and be sure they are safe before restrictions are lifted.  You owe this to the women who are depending on your judgment."

DR. GLICKSMAN:  Distinguished panel members, I have no financial affiliation to any corporation.  I have come to testify before you today as both a plastic and reconstructive surgeon and as a woman, wife and mother.

I began my experience with breast implants in the late 1980s where, at Memorial Sloan Kettering, I performed close to 350 breast reconstructions and augmentations using both saline and silicone implants.  In 1992, I began using only saline implants for augmentation and reconstruction and was an investigator in the McGhan saline study for five years.

For seven years, I limited my patients to only saline implants while I, like my colleagues, followed the national and international studies on silicone and systemic illness.  When, in 1998, the European nations declared silicone implants safe, I became an investigator in the FDA adjunct silicone study with Inamed.

To date, I have implanted 130 women with over 300 silicone implants.  I have followed these patients for close to five years with no reported systemic illness and a capsular-contracture rate of less than 1 percent.  You have before you the data, the science, but I would like to talk to you as a physician and a woman.

I, like many of my patients, am a 40-something-year-old mother of four small children.  We have all heard the arguments of the National Organization of Women.  Women should feel good about themselves.  Women don't need to diet and exercise.  Women don't need to feel attractive to men.  Well, I'm sorry, but NOW does not represent me or the thousands of my patients whom I have treated over the last twelve years.

We know the early detection of breast cancer and mastectomy saves women's lives.  If we are to encourage women to undergo such procedures, we must be able to provide them with an aesthetically natural, soft and comfortable breast reconstruction.

The majority of my patients seeking breast augmentation are in a similar position to myself, in their 30s and late 40s, have had several children and desire the fullness they have lost to pregnancies and breast feeding.  They are not looking for the Baywatch look.  They are well informed, educated, mature, American women who are seeking to restore their self esteem and improve their self image with a soft, comfortable breast implant.

If you, as a panel, carefully review the last ten years of rigorous epidemiologic peer-reviewed studies, you will see no proven connection between silicone and systemic illness.  Therefore, the final decision as to which implant to use should be made between the well-informed patient and her physician.

No medical device can guarantee a lifetime of safety.  Capsules, often tender and uncomfortable, form around neurostimulators, pacemakers and heart valves, and total joints do not last a lifetime.  When I sit with my patients, we weigh the risks and benefits of any surgical procedure.  Total joints may increase mobility and relieve pain and heart valves may prolong life, but both procedures carry considerable morbidity and mortality rates.

But what are the benefits of breast implants?  Who determines that the benefits of breast implants are not valid?  Is self esteem and improved quality of life only reserved for men on Viagra?  Are we protecting women in America or denying them the free choice that women in the rest of the world already have.

I urge you to let the science of medicine speak, not emotions, hysteria, money or politics.

Thank you.

DR. WHALEN:  Excuse me, Doctor.  I don't believe you stated your name.  If you could, please?

DR. GLICKSMAN:  Dr. Caroline Glicksman.

DR. WHALEN:  Thank you.

MS. GILBERT:  Dr. Glicksman, I have a question.  Can you tell me how you followed these patients?

DR. GLICKSMAN:  I follow my patients at six-month intervals.  I bring all my patients in, whether the study guidelines are at six months or a year or the three years and five years.  I follow all my patients every six months.  I bring all my patients back on an annual basis regardless of the study guidelines.

MS. GILBERT:  Is this done through MRI or is it physical examination?

DR. GLICKSMAN:  Physical examinations, mammography, annual mammography for those patients who are appropriate, unless there are problems, then ultrasounds and MRIs.  But I have not had any problems that required those testings as of yet.

PROF. DUBLER:  Dr. Whalen, since the panel is arranged so that we can't see who is talking, could you ask panel members to who ask questions to please state their names.

DR. WHALEN:  And who were you?  (Laughter)

DR. EDWARD MELMED:  I am Edward Melmed.  I have no financial payment made to me but my trip was supported by NOW.  I am a board-certified general surgeon and plastic surgeon.  I am board-certified in England, South Africa and the United States.  I came to the United States in 1973 as assistant professor of plastic surgery at the Southwestern Medical School in Dallas and I am currently in private practice.

I did an enormous number of breast implants.  I actually started in plastic surgery doing implants before there were implants.  We used to use buttocks.  I have seen the evolution of implants and I wrote extensively about it.  In 1992, I changed and I started seeing women with problems.  And now, when I counsel women and I want to tell you I have now done over 600 explantations, and if you include the miserable Pep implants, probably about 650 because every one of those ruptured.

When I meet with women I tell them there are three problems with implants.  The first is they rupture.  Now, when you see an implant that is three-years old and you take out an implant that is three-years old, you are going to see a perfectly good gel implant that is not sticky.  It does not cause any problems.  The likelihood of capsular formation is minimal.

When you get to four-and-a-half years, and these are implants that I have explanted fairly recently--when you get to four-and-a-half years--I beg your pardon.  I don't do that quite as roughly when I explant.  (Laughter)  I will get to it eventually.  I apologize--the implant is now starting to get sticky and if I wasn't wearing gloves, I'll have a residue on my hands. Implants are porous.  They have silicone bleed from the Day 1.  But when they get to four-and-a-half years, we started to get significant tackiness to implant.

A few months ago, I explanted a lady who I had implanted who had implants that I had put in seven years ago.  What came out was the typical explanted mess.  Now I have seen this hundreds of times and I am worried.  I am very worried because we have an implantable surgical device.  When I hear the last speaker make an analogy to a hip joint, a hip replacement, we are talking about a 70-year old who has got pain.  We are talking about 17-year-old girls who are faced with a lifetime.

One of my colleagues in Dallas said, This is a matter of informed consent.  I, he says, put the silicone implants into my own 17-year-old daughter.  I have lots of comments I could make on that.

DR. WHALEN:  Doctor, you need to conclude, please.

DR. MELMED:  I was granted six minutes, sir, because of another speaker.  So, my first question to you, as a physician in practice, is tell me what is an acceptable failure rate of an implant.  We accept that hip implants break.  That is one thing.  But here we are talking about putting an implantable device into a young person, and I, in practice, would like you to guide me and tell me what you consider an acceptable failure rate.

We know it is upwards of 70 percent after a given time.  And time is the factor that we are talking about in implants.

The second problem that implants have is they form tremendous scar tissue and, perhaps, it is relevant to show a specimen of what it looks like at 15 years--I am too efficient for the airport.  I apologize again.  This is a capsule that I took out after 15 years.  This you could probably put on the soles of your shoes and wear for a while, it is so thick.  It is full of silicone inside.  This horrible tissue is in a human.

Now, if you extrapolate that you are going to put implants into 17-year-old girls, at the age of 35, you are going to have these problems.  So my next question to you is what do you consider a time element at which you consider implants must be replaced.  Surely, you are going to replace them or you are going to just say, you can have this inside of you; we don't care.

As my FDA, I want you to advise me what you consider a time element, first in terms of product safety and failure and, secondly, time in terms of scar contracture and its resultant problems.

The third issue I don't want to dwell on because you have heard about it, but when it comes to symptomatology, I see an enormous lot of women with problems.  Undoubtedly, a lot of psychological overlay.  Undoubtedly, there are a lot of who real disease.  The truth probably lies somewhere in the middle.  I think our legal profession probably muddied the water by going after established connective-tissue diseases.  We know it is not that.  The Institute of Medicine is 100 percent correct.  This is not a problem that is classical.

DR. WHALEN:  Doctor, conclude, please.

DR. MELMED:  My last question for you is, then, if you are going to reintroduce this product, please tell me that it is safe.  What do I tell me patients?  Do you accept 1 percent, 2 percent?  That is for 100,000 women.

I think you, sirs.

DR. WHALEN:  Doctor, which of the other speakers yielded the time, just for our list?

DR. MELMED:  I beg your pardon, sir?

DR. WHALEN:  Can you just tell me which other speaker yielded to you for our list?

DR. MELMED:  I don't even know.  I was sponsored by the National Organization of Women and they said that somebody had given up their time for me to talk.

DR. WHALEN:  Thank you.

DR. WOOLEY:  Hi.  I'm Dr. Wooley.  I am also sponsored for travel arrangements by the National Organization of Women and I also have six minutes due to a donation from a breast-cancer survivor.

I am a professor of orthopedic surgery, immunology and biomedical engineering at Wayne State University.  One of the studies that I am going to describe to you was sponsored by Lifomatrix*, the manufacturers of Trilucent.*

I am going to take the first minute to describe the findings of the scientific panel convened by the National Organization of Women in July of this year.  This is a collection of scientists including those from the FDA, the NIH and academia.

We are not anti-manufacturer and we are not anti-implant.  However, our review of the literature and recent findings including some not published reveal three concerns to this group.  The first, as you have heard, is device failure.  It appears that elastomer shells are still experiencing a high rate of failure with recent studies suggesting about 17 percent rate of detectable leaks within five years of implantation.

The second concern is local tissue reactions, capsular contracture, which you have just seen a sort of dramatic demonstration of, and tissue inflammation due to silicone exposure remain a viable concern.

The third issue which I am going to speak to scientifically is latency issues.  The study suggests that long-term exposure to silicone may result in an increase in both immunological and autoimmune connective-tissue disease problems, albeit not classic.  That had been implicated during short-term testing.  However, the consideration of the two years of clinical data, it is the panel's opinion, is woefully inadequate.

I am going to move on to preclinical or animal testing and I would like to acknowledge the FDA for publishing, if not the whole technical report, summaries on their website.  I would like to quote from Study 7.  "While a significant increase in total spleen activity was observed in high-dose animals implanted with the shell test article, the increase appears to be related to a low response of the sham control animals."

I have heard far better explanations from recalcitrant graduate students trying to explain away data.  This is a fairly reasonable explanation of what we believe is the adjuvant effect of silicone, something that the manufacturers are quite well aware of.

If the panel would like to look at the handout that I provided them, I am going to contrast two studies, the first of which was done about ten years ago and looked at the implantation of silicone within an experimental model of autoimmune arthritis.

Our finding in our first short-term was there were  no adverse effects of silicone implantation of either the elastomer or gel on the incidence of this model Type-2 collagen-induced arthritis.  There was no adverse influence of the elastomer or gel on the severity of the disease.  However, immunological findings did reveal that a variety of proteins which had adhered to the elastomers were observed and autoantibodies unique to this type of implantation were detected.  But I want to emphasize again, these autoimmune findings did not influence the disease.  This is published in the leading journal of the American College of Rheumatology.

A follow-up study to look at long-term effects was quite different.  Long-term silicone elastomer--by long-term in a mouse, I mean a year which is a considerable portion of a mouse's life-span--significantly increased the incidence of arthritis in the mice and long-term silicone elastomer implantation significantly increased the severity in the immunized mice.

So what I am afraid of is that the panel, by looking at only two-year data or short-term data is going to make the same misconclusion that we did with our initial look at short-term implantation and that long-term implantation must be considered. These autoantibodies that we see included anticollagen antibodies, particularly Type I.  That is a protein that is inevitably going to bind to any elastomer.

What is the significance of animal findings?  Well, extrapolation of animal findings to the human situation is controversial, but this is an essential part of drug and device development and it is the panel's responsibility, of course, to consider that.

The findings suggest that the implantation of silicone could exacerbate autoimmunity in susceptible individuals resulting in the development of autoimmune connective-tissue disease.  We do, of course, predict that there are many other contributing factors particularly genetics which will influence the response to silicone in patients.

I would like to thank the panel for their time.

DR. WHALEN:  Thank you.  Before the next speaker, we apparently are going to take a brief unscheduled break of just five minutes.  So, please, don't stray too far.  Five minutes.  A reminder to the panel, just talk about how the Red Sox are going to beat the Yankees today.

[Break.]

DR. WHALEN:  We would like to resume.  It may have been obvious to anybody who witnessed what just transpired, but, for future speakers, please, if anybody has any specimens that have potential as having biologic properties, please don't bring them to the podium or else we are going to be suspending activities once again.  If we have to do that again, we may need to cut some speakers off and we certainly don't want to do that.

Please.

MS. GERTNER:  Good morning.  My name is Nancy Gertner.  I want to thank the panel for hearing my testimony today regarding my experience with silicone implants.  I am a retired attorney, a general surgeon's wife, and a five-year survivor of breast cancer.  I am here representing myself.  I have paid my own travel expenses to be here and I do not have any financial interest in Inamed or its competitors.

When cancer was found in my right breast in a routine mammogram in 1998, I elected to have a double mastectomy.  My choice to have what I thought was a prophylactic mastectomy on my left breast proved to be a prudent one as the surgery revealed cancer was also present in that breast.

Before my surgery, I discussed various reconstruction options with my husband and my plastic surgeon.  I knew I wanted to have immediate reconstruction following my mastectomy surgery and we discussed the options available including saline implants, silicone-gel implants and tram-flat reconstruction.

My plastic surgeon reviewed the benefits and potential complications of each procedure with me.  He showed me both pictures and actual implant samples.  We discussed the advantages, disadvantages and likelihood of problems or complications with each option and I chose silicone-gel-filled implants because I felt they offered the look and feel most like my own natural breast tissue.

I also knew other women who were satisfied with their silicone-gel-filled implants.  I decided to have tissue expanders inserted at the time of the mastectomy surgery in November, 1998, with subsequent expansion.  The silicone implants were inserted in February of 1999.  I was immediately pleased and continue to be happy with the result.  Because silicone implants were a viable option to me at the time I was diagnosed with cancer, I was much more able to cope with the loss of my breasts.

As a woman, it was also important for me to have a choice and be able to make an informed decision.  While every woman is different, gel implants were the best alternative for me and I am grateful my doctor was able to present them to me as an option.  I felt then, and continue to feel, that reconstruction with silicone implants provided me with the best possible aesthetic result with a reasonable post-op recovery time.

Having options and being explained my choices made me feel empowered as a woman and made me feel like I was not doomed to a single situation.  I am happy with my reconstruction.  I think it looks and feels beautiful and I have never experienced any personal or professional limitations as a result of my implants.

I am frequently asked and glad to show other women my reconstruction and describe my experiences hoping to allay their fears when they, too, are diagnosed with breast cancer.  I feel this because I was given many reconstruction options including gel implants.  Breast cancer ended up being a small problem in my life rather than a big one that my body would have been a constant reminder of.

I sincerely appreciate your time and appreciate the opportunity to share my experience.  Thank you.

DR. GOLDBERG:  I am Eugene Goldberg.  I am a professor at the University of Florida.  I am going to try to provide some scientific information pertinent to this discussion.  I have no financial interest.  I have paid my own way here.  I have worked in the past as a consulting expert in litigation on behalf of both defendants and, as well as defendants, plaintiffs.  I want you to understand that clearly.

I have some charts up there which about a half a mile down there you are not likely to see.  But they are in a handout so that may help.

My concern for this discussion will be--and if my voice breaks up, it is not because of puberty, it is because of bronchitis.  So please forgive me.  My concern is with the risk/benefit issues and, listed here, are some of the safety risks that are associated and are well known from fibrous capsule contracture, pain, disfigurement and I am going to refer primarily to the question of shell failure and additional surgeries because that is really the substance of much of what we have done research on at Florida.

I should mention parenthetically that several of the papers that we have published in Applied Biomaterials and in the Annals of Plastic Surgery are not discussed in your PMA nor is our failure analysis discussed in the PMA.  The 1992 moratorium on silicone-gel breast implants was imposed due to inadequate safety data.  The question is, do you have more safety data now in a relatively short two- to three-year period of evaluation, than you had at that time.

Sound engineering-design criteria for gel implants and the mechanical and biochemical properties have not been available and no peer-reviewed large cohort prospective or retrospective clinical studies have been reported by manufacturers today, even though gel implants have been on the market for 40-ish years.

The only peer-reviewed large cohort failure and safety analysis paper to date is from our laboratory and not discussed in the PMA.  This analysis shows a statistically significant result for almost 10,000 explanted devices from 42 separate studies.  Some people have considered, in the literature and in private discussion, that this may be a biased study.  If it is biased, it is biased in favor of having a large, the largest, cohort available from 42 studies and, therefore, represents a very wide range of different implants from different time periods, all of them gel implants, for which we have significant data in the literature.

The results of that study can be represented qualitatively by some of the things you have just heard from Dr. Melmed where I show some examples from our laboratory examinations ranging from six-and-a-half years which are intact to 13-and-a-half years which is obliterated, literally, hardly there anymore.

The high prevalence of shell rupture and the frequency of additional surgeries is something that we have tried to emphasize beginning around 1995, first in a letter to the New England Journal and subsequently in considerable research in our laboratory.  We observed, as you will see on the failure analysis curve that goes up next, a shell failure prevalence at five years of 30 percent, at ten years, 50 percent and 70 percent at about 17 years.

This explant meta-analysis, not discussed in the PMA, substantially is confirmed by a noninvasive MRI study that has been provided to you by Dr. Brown and coworkers and essentially indicates that this curve is a reasonable representation of even what has been termed silent ruptures.

If you look at two years, at the beginning of that curve, you find that it is extremely difficult to make any sense, whatsoever, out of rupture rates and reoperations.  Our data from this study indicate that the reoperative frequency is of the order of one woman in three within six years and sometimes as many as two or three or four surgeries within the time frame.

Two other points, quickly, I would like to make.  One is that this is a new shell.  This is not the old dimethylsiloxane shell.  It is a phenyl shell and the phenylsiloxanes are associated with carcinogenic behavior as indicated in the PMA and there is also indicated the PMA a higher prevalence of both pulmonary and brain cancer which has been glossed over.

DR. WHALEN:  Dr. Goldberg, you need to conclude, please, sir.

DR. GOLDBERG:  I will conclude.  To conclude, therefore, it seems to us that there is a need for much more scientific information before this new implant should be approved.

Thank you.

DR. WHALEN:  As the next speaker is coming up, just for planning purposes of the speakers who are representing societies, it looks like we will probably be having some of you going after lunch.  We will breaking for lunch at 12:30 for 45 minutes.  If there is anyone towards the tail-end of the society speakers to whom that would impose difficulty, if AnneMarie Williams would raise your hand--you probably already know her if you are speaking, in the back of the room--please see her now so that some juggling of the speakers can be done.

Yes, ma'am.

MS. TAPSCOTT:  My name is Mindy Tapscott.  I am 42-years old and I am a bilateral breast-cancer survivor.  I would like to thank the panel for taking into consideration my experience with silicone breast implants.  I also want it to be known that I am receiving no compensation from any group in any shape for form for my testimony.

I was first diagnosed with breast cancer in 1989 and I knew I would be seeking reconstruction of my left breast following mastectomy surgery.  After discussions with my surgeon, I elected to have reconstruction with a silicone gel-filled breast implant.  I felt this option looked more natural than the other reconstruction options that were available to me.

My surgeon provided and asked me to review and consider the insert that the manufacturer included with each implant.  In addition, she reviewed with me the potential complications associated with silicone implants including capsular contracture and rupture.  She informed me that some women are more prone to complications than others and that should could not reproduce what God gave me.  I, nonetheless, felt comfortable with my decision.

At my initial post-surgical follow-up visit, my surgeon advised me that massaging my breasts would help reduce the risk of capsular contracture but, I must confess, I did not follow through with that recommendation and, after about one year of my first implant surgery, I experienced significant capsular contracture in my left breast.

I decided not to do anything about the encapsulation until ten years later, in 1999, when my implant was replaced.  I requested that it be replaced due to the increased size of my right breast resulting from two pregnancies.  After the procedure, my plastic surgeon specifically told me that the first implant was intact.

One year later, the new implant also became encapsulated and, in the same year, I received a diagnosis of breast cancer in my right breast.  Despite the complications I had experienced with my first implant, I also elected to have a silicone-gel-filled implant inserted after my second mastectomy.

This, too, became encapsulated shortly thereafter.  In January of this year, I decided to have both implants replaced once again.  The latest procedure produced a terrific outcome and I have had no trouble at all since that time.  Despite my history with local complications and the number of reoperations I have had, I still consider silicone breast implants the option of choice for me following my breast cancer.

When you are diagnosed with breast cancer, you feel a loss of control because you don't know what you did to get the disease.  Having choices made me feel like I got some control back.  The many operations I have had do not bother me one bit.  It was all worth it.  Silicone implants have simplified my life.  I am a mother of three small children and I can barely remember my purse when I leave the house, let alone my breast prosthesis.

I don't have to worry if my breast is in before I answer to door or if my daughter has a friend sleep over.  I never have to worry about my prosthesis falling out or coming out of position.  The silicone-gel implants look and feel most like me and allow me to get on with the business of living.

I am a carrier of the Bracca 1 mutated gene.  The odds are great that I have passed this gene on to at least one of three of my daughters and I would want them to have the same choices.

I thank you for your time today.

MS. GANDY:  Good morning.  I am not Marlene Keeling.  I just traded places with her.  My name is Kim Gandy.  I am President of the National Organization for Women.

The National Organization for Women is the oldest and largest feminist activist organization in the United States advocating not only for women's equal rights, their economic and social well being but, also, for women's health.  The National Organization for Women and myself represent only our own 501-C4 nonprofit corporation and no other party.  My appearance today is underwritten by our 500,000 contributing members across the country.

We have also asked three researchers and clinicians who participated in our May 8 symposium and our July 21 press conference to present testimony at this hearing as well.

Thank you for the opportunity to present NOW's comments today on this very important issue.  With hundreds of thousands of women from high school to retirees seeking breast implants each year and the number continuing to increase, a dangerous product will affect the health and safety of enormous numbers of women, individual women as well as their families.

When a substantial number of these women fall ill, and depend on government healthcare systems for their medical care, the safety of their breast implants, a question you must determine here, will have become an important public-health question.

But it is a personal health question for hundreds of thousands of women who need complete information in order to make informed choices about their medical care.  Our organization has been contacted over the years by many women with silicone breast implants who are suffering from serious illnesses and disease that they associate with those implants.

Their stories are heartbreaking.  Many report having lost not only their good health but their husbands and partners, their jobs and even their life savings.  We have also heard cases of suicide by these women, some of them mastectomy patients who became even more ill from their breast implants than they did from breast cancer.

These women have reported a wide range of symptoms similar to fibromyalgia and lupus all the way to rheumatoid arthritis and other serious conditions.  Most of them cannot afford the $10,000 to $15,000 and more that it costs to have their implants removed and they don't know where to turn for relief of their pain and uncertainty.

Informed by the personal stories of these survivors, the assembled membership of NOW at two national conferences have adopted resolutions urging both independent research and long-term following of implant patients in addition to fully informed consent for all prospective patients.

This spring, on learning that the FDA had been asked to consider returning silicone-gel-filled implants to the broad market, NOW sponsored a symposium on May 8 inviting more than a dozen researchers and government health officials from the Armed Services, from the National Institutes of Health, from the FDA and from private research institutes and universities to review their own findings and their own research with regard to the effect of silicone gel in the human body.  Of course, you have heard already from some of those participants and will hear from more.

We did this out of a conviction that government reliance on industry-supplied data, when special interests are involved, may have a result that is good for the industry but bad for the consumer.  In reaching out to potential participants, we also learned that breast-implant manufacturers and plastic surgeons have made researchers and medical writers feel intimidated by spending great sums of money in persuasive campaigns to mislead breast-cancer-patient advocacy groups and the public regarding breast-implant safety.

Women deserve the truth.  They deserve to trust that what this agency tells them is a safe product to put in their bodies is, in fact, a safe product.  When the information they are given is based on flawed or biased research, the result can be tragic.

We know the breast implants leak and rupture and disintegrate in the body over time.  We know that this causes a variety of problems, not the least of which is the promotion of systemic connective-tissue disease.  I will refer you there to the research report which I have brought many copies of but particularly the research by Miller and O'Hanlon of the National Institute of Environmental Health, part of the National Institutes of Health, the Environmental Autoimmunity Group which drew direct connections between silicone gel and systemic connective-tissue disease over the long term.

Our symposium of researchers and clinicians have made numerous recommendations for future research and for improving the FDA's review process.  These are included in the summary document which I would like to append to my comments for your review.

If these preliminary findings do, indeed, lead to more conclusive assessments about silicone gel and immune-system pathologies, then we believe it will be clear that silicone-gel implants are not safe over the long term.

We believe--

DR. WHALEN:  Can you conclude, please, ma'am?

MS. GANDY:  Yes--that if it is in the interest of this industry to develop and promote a safe silicone-gel-filled implant that they would have done so by now.  It is clearly not in their interest to do that.  The only thing we believe that will result in a truly safe implant for women to choose is if the FDA requires that it be safe.

If you permit them to continue to market unsafe implants to women, then they will have no motivation to make a safe implant for ourselves and our daughters.

Thank you.

DR. ANDERSON:  Dr. Whalen, may I ask the speaker a question?

DR. WHALEN:  Yes.

DR. ANDERSON:  Ma'am, you spoke of the symposium on May 8.  When NOW put this together--we are hearing two different viewpoints.  Was NOW trying to present a discussion that included both sides or was the purpose of this to collect data that supported the view that implants are unsafe?

MS. GANDY:  We were trying to find out what was going on because we were getting very mixed messages.  We heard from groups of survivors, groups that represent survivors, that they were told by doctors that these were safe and that there was no problem and that we would be doing a disservice to women by not allowing them.

Then we had heard from so many groups of survivors and individual survivors that had terrible, terrible problems.  The Co-Chair of our National Advisory Committee is a medical doctor and the dean of a medical school.  And we knew that there was some unclarity over the long-term impact.  So we thought it was important to reach out to people who were actually doing the work, who actually were doing the studies, themselves, who were looking at the women and who were engaged in the scientific part of it as opposed to the advocacy part of it.

DR. ANDERSON:  Did you have speakers that advocated for implants as well as speakers that had had these bad experiences?

MS. GANDY:  The speakers didn't advocate one way or the other.  There actually was not a single speaker that advocated for or against the implants.  That was a conclusion that was drawn from all of the research.  What they did say was, and you will see the conclusions there--they found that there were severe painful and debilitating local complications from the breast implants, capsular contracture, the kinds of things that you already know, that there was a very high failure rate and that, over twenty years or so, it approached 100 percent and that there was substantial indication over a variety of the reports of the  silicone gel, in particular, leaking into the human body had potential long-term consequences.

They approached this from a variety of different directions, not making--you are accustomed to dealing with doctors and researchers more than I, but what they did was they said, here are my findings and I have concluded that this seems to do this and we need to look at a longer sample, we need more information.  That was the primary message that came from every doctor, from every researcher that testified, was this bit of information that we have leads me to think that there is a real problem here.  But we don't really know until we have this much information.

Every one of those doctors came to the same conclusion.  As a result, we came to that conclusion, that this body should not be making any decisions on silicone gel until we had this much information and not this much information.

DR. ANDERSON:  Thank you.

MS. GANDY:  Thank you.

MS. WARD:  Good afternoon.  My name is Mary Ann Ward and I am from Coco Beach, Florida.  I am here representing myself and I have paid my own travel expenses to be here, have no financial interest in the manufacturers or its competitors.

I appreciate the opportunity to speak with you today about my experiences as a breast-augmentation patient with silicone-gel-filled implants.  I made a decision in 1978, when I was 29-years old, to have breast-augmentation surgery.  I always knew my implants would need to be replaced at some point.  My plastic surgeon was completely clear with me regarding that.

Last year, that time came.  The decision to have the implants replaced was solely my own.  I was not suffering from any complications.  I was not suffering from any disease or condition.  I just thought, after 24 years, I should have them replaced.

After making the decision to have my implants replaced and have additional surgery, I consulted an experienced board-certified plastic surgeon and we reviewed different options including the type of implants available, the advantages and disadvantages of each option and I received a significant amount of medical information that I was able to apply in examining not only my health issues but my particular lifestyle and personal expectations.

I am very proactive regarding my health as I believe most American women are, and I did research on my own.  The two studies which were particularly helpful to me were the Mayo Clinic Study which was published in the New England Journal of Medicine in 1994 and the Nurses Health Study of 1995.  Both studies found no association between breast implants and systemic diseases.

My physician encouraged me in my personal education and discussed with me every piece of research I found.  For me, having options was critically important not only from a health perspective but the implants needed to fit my lifestyle.  You see, I am a 54-year-old woman who races sailboats competitively in the national and world arena.  I race in bathing suits.  I have to pull lines.  I must be strong.

It was important to me that my strength and ability to continue my sport would not be compromised.  While there are no guarantees in life, my physician made sure that my education on the issues was thorough and that I was making a well-informed decision.

Since my surgery, my physician has stayed close in contact with me requesting feedback on both process and results.  I have been thrilled with the look and feel of my breasts.  The gel implants have allowed me to have a natural appearance for a woman of my age while not interfering negatively or impacting my lifestyle.  I continue to race sailboats.  I work out.  I play with my grandchildren and I do much more.

I believe all women should have the choice I had.  Given the science and research today, it is illogical to me that silicone implants are only available to women for reconstruction surgery.  You have heard many compelling stories today, many of them very sad.  My story is no less compelling because it is a happy one.

Thank you for your time.

DR. CHOTI:  May I ask a question?

DR. WHALEN:  Sure.

DR. CHOTI:  What was your thought process or thinking as to why, after 24 years without a problem, you would get them replaced?  Why not at ten?  Why not at 15?

MS. WARD:  I actually was having no problems whatsoever.  So I think I just kept putting it off.  Finally, I had changed--at the time that the tests were done regarding hormone-replacement therapy, because I had been through menopause and I am on hormone-replacement therapy, when all of that came up and I started having in-depth discussions with my doctor about that, I actually changed doctors at that time.

The new doctor that I went to said, you know, Mary Ann, it has been 24 years.  You should have this looked at more thoroughly than your normal breast exams, than, you know, your annual mammograms, than examining yourself.  It has just been that long.  I think you ought to have somebody take a look at them.  So I did.

DR. CHOTI:  But, at 15 years, you had the same thinking and decided, at that point, no need to replace them?

MS. WARD:  I really didn't.  As I said, I am very proactive about my health so I was attentive to my breasts, most assuredly.  But only really when I changed doctors and it was a new--my new general physician said, I think you ought to at least go see a plastic surgeon and talk to him about it.

When I saw my plastic surgeon, she was quite honest with me and said, you know, we see no real reason to replace them unless you just want to, which I made that decision because, when I had them put in in the first place, they said, these will not last you forever, Mary Ann.

DR. CHOTI:  And when they were removed, there were no leaks or ruptures?

MS. WARD:  They were fine.  I was fine.  I continue to be fine.  Thank you.

DR. WHALEN:  A reminder to panel members of Dr. Dubler's appropriate request that we start off questions identifying ourselves.

MS. KEELING:  My name is Marlene Keeling.  I paid my own way here and I have no conflicts of interest.

I chose to get double-lumen breast implants in 1978 after my plastic surgeon assured me only 2 percent of his patients ever had problems and he could fix anything that happened.

My health problems did not start until about seven years after my surgery.  I developed weakness, fatigue, periodic swollen lymph nodes in my neck, depression, memory loss, peripheral and demyelating neuropathy.  Since removing my implants, my ruptured implants, in 1994, all of my symptoms have slowly improved.  However, the last time I had my natural killer cells counted, they were low.  This is consistent with a finding of the Institute of Medicine report that states, "Consistent with animal toxicology studies, it appears that natural killer cells might be affected by exposure to silicone gel."

Inamed data has revealed the loss of natural killer activity in their immunotoxicity research.  How can this panel approve a device for breast-cancer survivors that compromises the immune system by lowering natural killer cells which are the cornerstone of the immune system, are responsible for controlling tumor-cell growth and are involved in control of infection?

In networking with other silicone-implanted women, I realize that many were being diagnosed with chronic inflammatory demyelating polyneuropathy and had memory problems or depression.  In 1996, I became a founding director of Chemically Associated Neurological Disorders which is dedicated to information and research.

German research by Flasseck et al., published 2003, found higher platinum levels in the fat tissue of breast-implanted patients when compared to control patients without implants.  The same researchers report the data of this work clearly show that the use of elemental silicon as an indicator of migration from breast implants to the surrounding tissue is not appropriate.

This becomes important because the Dow-funded Semple, et al., study on breast milk from implanted women used the inappropriate measurement of elemental silicon in an attempt to imply that breast feeding by implanted mothers is safe.  In an e-mail, I asked the German researchers, could you comment on whether you believe that elemental silicon is an appropriate measurement for breast milk from implanted women to determine safety.  Their response was, of course not.

How can this panel vote to approve devices for augmentation and revision of young women of childbearing age without a properly designed breast-milk study for safety of children born after implantation.

I have one final point.  Two years ago, McGhan Medical, now Inamed, filed a patent which makes the statement, "Silicone oil is not biocompatible with other human tissues.  Therefore, the use of silicone-based filler material has been discontinued in the industry."

How can this panel approve a device that harms human tissue, leaks silicone and heavy metals through the shell, ruptures at an underascertained rate, and the rupture may be silent.  There is no way to remove all of the toxic and hypersensitizing chemicals once they leak into the tissues.

Women with breast implants have not had informed consent.  I say, loud and clear, you are the scientists in charge who must stop this experimentation on women and our future generations and must say no to corporate greed.  I don't have a Ph.D. in suicide risk, but I can tell you, in networking with thousands of breast-implanted women, many do not have the money to have their ruptured and leaking implants removed.

Many have constant pain.  Many have declining health.  Many have lost their health insurance or have been denied coverage.  Many have children born after implantation with similar health problems and suffer guilt.  Many are now no longer able to take care of their families and do not want to be a burden.

Thank you.

MS. FERELLI:  Hello.  My name is Vanessa Rose Ferelli.  I have not been given any financial assistance from anyone related to this testimony.

I am 29-years old and I was implanted on August 7, 2000, as part of a study from Port Jefferson in New York.  I was misled about the truth of what the study really was about when I received the implants and I have been informed now that I wasn't even eligible to receive the implants at the time.

Over the year and a half that I had the silicone implants, I was hospitalized many times.  I was examined by a cardiologist, rheumatologist, my gynecologist and not once did anyone ever tell me that I could be sick or having some kind of a reaction to my implants.

I experienced abnormally heavy periods that lasted several months at a time, fatigue so bad that I slept 16 or more hours a day, numbness and pain in my arms and my hands.  I couldn't even hold a glass in the morning.  I also had severe breathing problems.  I lost my hair--facial sores and a lot more.

On December 1, I called my plastic surgeon from the hospital--I was in the heart hospital having heart problems--and asked them again, was there any possible way that this could be related to my implants.  And he told me absolutely not.

Finally, I met with a rheumatologist in New York that told me the truth about what was going on and he suggested I remove the implants as soon as possible.  He said I was having something similar to an allergic reaction to them.

On May 5, my implants were removed and they were sent to Inamed.  The right one was totally ruptured and the report showed granulomous inflammation from the specimen sent to the lab.  I didn't get a letter back from Inamed who received the implants right away until October 11.  I didn't feel that they were really concerned about what had happened or studying anything since now I found out I was part of a study.

Immediately after the implants were removed, I felt much better but I was still having breathing problems that were severe.  I ended up going to another plastic surgeon who removed all my natural breast tissue, the surrounding tissue that was contaminated and that also had silicone in it.

It has been about two-and-a-half years since this has gone on and I am much better now.  I haven't had any of those symptoms again.  The tachycardia and problems with my heart and the breathing problems are all pretty much gone.  I still have some minor symptoms but nothing like what was going on before and it cost me about $45,000 for the surgeries, the ordeal after the fact between the scans of my chest that never showed or revealed any kind of a rupture and I feel that I am lucky now to be okay, but I couldn't have imagined what would have happened if I didn't have the financial means to be able to take care of having them removed because a lot of things weren't covered under my insurance.

I just ask that you please make sure that the implants are safe before they are approved.  That's it.  Thank you.

MS. GILBERT:  Ms. Ferelli, my name is Alisa Gilbert.  I just have a question for you.  You said you were part of a study.  Did the study cover any of your expense when you had to go get the implants removed?

MS. FERELLI:  The only thing that they offered to me--I called them when I talked to the rheumatologist and realized why I was sick.  And I called them right away because I wanted a list of maybe chemicals that were now in my body, possibly.  I was nervous about that.

And they said that they would send me some information.  All they sent me was a packet and a claim form and they said I could get my money back for the implants which I think were, like, $1500.  That's it.  But as far as money--and my insurance company wouldn't cover--even if I wanted to go see if there was a rupture to get the--I think it is a C-scan or something when you go to scan.  Do you know what I am talking about?  I'm sorry,

MS. GILBERT:  An MRI?  Are you talking about an MRI?

MS. FERELLI:  Yeah, something like an MRI but there is a special one that supposedly they could really tell to see if there was a rupture and they had me going into all of these different types of scans, but nothing revealed any kind of a rupture.  But I had a lot of pain on my right side and burning.  And I kept complaining about the right side and the burning.

Then we didn't know until they removed it that the right one was totally ruptured.  Thank you.

PROF. DUBLER:  Dr. Whalen, question.  Could you come back to the podium for one minute.

MS. FERELLI:  Oh, I'm sorry.

PROF. DUBLER:  I am Nancy Dubler.

MS. FERELLI:  Hi.

PROF. DUBLER:  Before your first implant, in your discussions with the surgeon, what did he--he did not tell you it was a research protocol, involving a research protocol?

MS. FERELLI:  Actually, I went to two different plastic surgeons because I had saline implants first and they were ten-years old.  And I wanted to have them changed because I was told that, in ten years, you should.  Both of the doctors said to me that they just brought the silicone implants back on the market, that they did all these studies and tests and they have proven that they are 100 percent safe and that all the stories that you heard of people getting sick wasn't true.  So I believed the doctors.

PROF. DUBLER:  You never signed a document that was labeled "informed consent?"

MS. FERELLI:  What I got--

PROF. DUBLER:  What did you sign before your first surgery?

MS. FERELLI:  What I got from the--which I found out now, all of my records and stuff, what I signed was one general release that I read that talked about anesthesia and stuff and another one, one page, that just said--when they told me that they brought them back on the market, he said that I just have to promise to show up for several visits so that they can document and study my progress because they just brought them back on the market.  So I said, okay.

They never told me that you could get sick from them.  If I did, I would have known a year and a half before--when I had symptoms, I would have first realized that I could be having a reaction, but I went a year and a half of having these symptoms and nobody told me, not only the plastic surgeon that put them in.  I was at St. Francis Heart Hospital in New York.  I was at many different hospitals and I saw top doctors, cardiologists and nobody ever said that it could possibly be from that.

Now that I know what happened to me, I have gotten--

PROF. DUBLER:  Just a brief follow-up question.  So the only research element you were aware of was your obligation to return--

MS. FERELLI:  Right.

PROF. DUBLER:  --for examinations following the implant.

MS. FERELLI:  Absolutely.  That is why I called him from the hospital on December 1 to tell him that I couldn't make my appointment, because I had promised him I would be there.  On the phone, he told me that this could have nothing to do with that and I have given documentation, and he also wrote up a letter and sent that to the study that said that I called from the hospital, I asked them if any of this could have anything to do with the implants, and he said, absolutely not.

So he has even put in writing on two different occasions that absolutely no way could this have anything to do with the implants.  So I believed him.

PROF. DUBLER:  Thank you.

MS. FERELLI:  Thank you.

MS. WASSERMAN:  My name is Connie Wasserman.  I am a New York certified social worker.  I am here on behalf of my niece, Vanessa Rose Ferelli, who I will refer to as Rose.  I am not receiving any financial reimbursement of any kind for my testimony here today.

My niece has been experiencing ill health since the end of 2000.  She was experiencing heavy menstrual bleeding that left her body anemic and exhausted.  Her gynecologist placed her on the pill hoping to regulate her cycle.  She developed other symptoms such as hoarseness in her voice, paleness in her color, difficulty breathing and a chemical taste in her mouth.

She would have episodes of passing out and was becoming progressively worse.  I honestly did not know what to think.  Doctors were attributing all these symptoms to extreme stress from the loss of her brother.  I will never forget the terror in Rose's voice when she called me at 4:00 a.m. to tell me that she had figured out what her problem was.  She was having a reaction to her breast implants.

Was this all stress related?  Something in her voice convinced me that this was real and she was very sick.  I immediately contacted my physician for help when he told me that there was no evidence that linked breast implants with disease or illness.  This was the same statement that we had heard from each of the other doctors.

I indicated to him that I had heard this disclaimer from the other physicians, however, I wanted to know what he would do if this were his daughter.  He then referred us to a prominent plastic surgeon who suggested removal.

The doctor took X-rays and examined my niece.  The X-rays did not show any leakage.  The doctor blamed her chest pain and burning on the size of the implants and then recommended they be removed.  He explained the details of the surgery with us indicating that there was a strong possibility she would not feel any better after the implants were removed.

I counseled my niece by helping her to make a well-thought decision and not one in haste.  I was concerned that if the implants were removed and she continued to feel sick, she might become depressed because of the scarring that could occur.  She was adamant that, after being so sick for so long, she just wanted the implants removed and she did not care how she looked.

Unlike many young women, Rose could afford to pay the doctor to remove the implants.  Surgery proved what the X-rays did not show.  One of the implants had ruptured.  After the surgery, Rose began the slow road to recovery.  Today, she is experiencing better health and more energy as well as elevated mood in spite of everything she has gone through.

She is just grateful to be alive and feeling better.  When I asked my niece what made her decide to get implants in the beginning, she told me that she was self conscious about her body and the doctor told her it was completely safe.  My niece, along with many women who get breast implants, believe their doctors.

In order to remove the implants, doctors had to carve out the scar tissue that formed around them.  Many argue that it is important for some women who have poor body image and low self esteem.  I cried when I saw the scarring that she was left with after the implants were removed.

What about body image and self esteem now?  Please, don't let this happen to other women.  Please make sure breast implants are completely safe before they are approved.

Thank you.

DR. WOLFE:  I am Dr. Sidney Wolfe.  I do not get any money from Inamed, McGhan or any of the other companies and I drove out in my car today.

Both the FDA approval standard for medical devices, "reasonable assurance of safety," and the validity of informed consent, informed choice, are legitimate only if studies are long enough in duration and comprehensive enough to answer existing questions about safety.

Instead of reasonable assurance of safety, there is now, more accurately, reasonable assurance of ignorance and, therefore, reasonable concern about safety.  If the decision is to approve now, women will not get the information they need for truly informed consent before deciding whether to undergo silicone-gel breast-implant surgery.

It is ironic that the clear evidence of certain danger such as local and regional complications, often from rupture, and remaining questions about other risks that led to the more-than-eleven-year-old moratorium on silicone-gel breast implantation, except for post-mastectomy patients, were based on much longer exposures of those women than now exists for the implants being seriously considered.

Despite these deficiencies in the duration of the data being considered by the FDA, only the first three years of a planned ten-year study, there are many red-flag warning signals within the thousands of pages of documents put up on the Internet last week, more than enough to justify a decision to wait at least five more years to collect more information on the thousands of women who already have had implants rather than precipitously approving these unsafe devices for marketing to hundreds of thousands of women a year who will, once they are told by the government that devices have "reasonable assurance of safety," opt for the surgery to insert them and, thereby, participate in a massive experiment.

I think Dr. Krause gave you copies of my testimony because there are some charts which I just want to go through briefly.  The first chart shows that the complication rates for the three-year data in the 1990 Inamed study and the current core study are quite similar examining only the data on those implants in 1990 which Inamed is now seeking approval for.  For example, for reoperation, it was 19.9 percent in the three-year 1990 and it is 20.6 percent in the three-year core.

The next chart, though, shows the data from three years with that of five years in the 1990 study and shows the striking increase in the risk of various complications occurring just in the two years between Years 3 and 5.  For example, the risk of removal/replacement increased from 6.3 percent to 11.5 percent in two years, capsular contracture went from 9.5 to 14.9 percent.  The cumulative risk of breast pain which wasn't provided at three years in these data was already 25.8 percent and loss of nipple sensation was 17.2 percent at five years.

Since the five-year follow up in the 1990 study was not as great as it could have been, it is not unreasonable to hypothesize that, in the current core study, when--not if; when--the five-year data are collected prior to approval, hopefully, the risks will be as high or even higher than those found in the 1990 study.

If you flip, now, to the third page, there are just four charts here that show, within the 1990 study, the effect of time on the increased risk of various complications, and this is in the face of having lost a number of people to follow up by the fifth year in that 1990 study.  Breast pain goes from 14 percent at one year to 25.8 percent at five years, loss of nipple sensation from 11.8 percent to 17.2 percent, asymmetrical breasts increased from 5.9 percent in one year to 15.3 percent at five years and the effect of the implant duration on whether the implant is actually palpable or visible, sort of contrary to the idea of a cosmetic operation, goes from 12.5 percent to 23.3 percent at five years.

Moving on to the next two charts, these are derived from the unpublished data from the SEER study funded by Inamed and Mentor using the NCI, National Cancer Institute, surveillance data.  The most striking finding is that the Kaplan-Meier risk of implant removal--this is projecting out on what you know, what is likely to happen--was 21 percent at five years risk of removal in these women who had had post-mastectomy implants.

But it has risen to 50.5 percent by ten years, more than doubling for those women with breast cancer who had a silicone-gel breast implant.  In other words, by ten years, more than half of the women had their implants removed.  This compares to 40.7 percent of those women getting saline implants.  This is the only population-based study that we have on this topic with the advantages of better ascertainment to examine the problem of implant removal.  It is noteworthy that, of the 6,563 women with breast cancer in the geographic areas covered by the study, which are Iowa and Seattle, only 1, 159, or 17.7 percent of all those women had implants.  The rest chose, I believe wisely, in most cases if not all, not to.

There are in the study details for 1,012 of these; in other words, we have 87 percent ascertainment excellent follow up, much better than any of these studies.  What it shows is, as I mentioned, it rises from 21 percent at five years, 27.4 percent at seven and 50.5 percent at ten years.  I mentioned that rate was relatively, 20 percent, lower in the women getting saline implants.

In addition to the primary deficiency of only having three years of data in the ten-year core study, according to the FDA, "in the reconstruction study, only 59 percent of patients have reached the three-year follow-up visit.  Worse, the follow up in the adjunct reconstruction group was only 53 percent at one year and 27 percent follow up at three years.

Other problems noted by FDA reviewers were that, in addition to 92 protocol violations in the core study, many involving a lack of informed consent as Vanessa Ferelli just mentioned, the adjunct study had even more egregious violations including 211 patients in whom informed consent was not obtained before surgery, 86 patients implanted in a facility without ethical review-board approval, and 25 patients implanted by non-authorized investigators.

According to another FDA medical officer who I think will present data this afternoon, Dr. Dawisha was very concerned about the lack of ascertainment of asymptomatic rupture since only a small fraction of these core patients had MRIs to detect rupture.  This is particularly of concern because the FDA's own study, Dr. Brown's study, which she will talk about, based on 344 randomly chosen women who had had silicone-gel breast implants, found that 68 percent of the women had at least one ruptured implant using MRI.  The median age of implant rupture was 10.8 years; in other words, half of them had ruptures after 10.8 year.  21 percent had silicone gel outside the capsule in at least one breast.

Despite the concern about implant rupture and spread of silicone around the body, FDA stated that, "Inamed provided no discussion of the significance of implant rupture," which includes published cases of lymph nodes, abdomen, groin and so forth.

The final irony can be found in the section entitled Preclinical Shelf Life in one of the FDA documents.  The document states, "Based on all shelf-life data provided, Inamed supported a 2.5-year expiration date on their package label.  The tests included gel cohesion--obviously if it migrates to the body, the gel isn't terribly cohesive--shell ultimate break force, shell tensile set and shell elongation.  That is from the FDA document.

Even though so many women are led to believe these implants will last at least ten years if not a lifetime, this 2.5-year "expiration date" is a further insult.

The apparent unwillingness of the FDA to require silicone breast-implant manufacturers to provide longer-term data as a prerequisite to approval raises serious questions about whether the agency is enforcing the 1976 Medical Device Law to protect patients.

The major precipitating event leading to the passage of that law was the disaster concerning the Dalkon Shield intrauterine device.  Seventeen women died and thousands of women were rendered sterile because of the increased uterine infections caused by this device.

Requiring adequate preapproval testing for permanently implantable devices such as breast implants was a major principle underlying the passage of the 1976 law.  The agency appears poised to negate this principle as they consider the approval of breast implants.

This year marks Sybil Goldrich, who generously gave me some of her time, twentieth anniversary since her battle with breast cancer.  But that was only the beginning of her fight on behalf of women.  After undergoing a bilateral mastectomy in 1983, she received breast implants.  She, over the next couple of years, underwent seven operations, two mastectomies.

During this time, she experienced excessive pain, capsular contracture, migration of gel to her uterus, ovary and liver, extrusion necrosis and rupture.  Finally, she chose to have a tram-flap procedure.  In 1998, Sybil cofounded Command Trust Network.  I met her fifteen years ago when we originally got involved in silicone-gel breast implants, a national clearing house dedicated to providing thorough and unbiased information about breast implants.

Since then, she has worked tirelessly to ensure government and corporate accountability with regard to breast implants.  Her goal is to ensure a safe implant is available to all women.  She does not think that the short-term data submitted for FDA approval begins to approach such a goal.

Thank you.  With three seconds to go.

DR. WHALEN:  Are there any panel questions of Dr. Wolfe?  Seeing none, I would like to thank everyone who has testified this morning.  We are going to break for lunch and resume activities precisely at 1:20 p.m. when other society testimony will be heard.

[Whereupon, at 12:35 p.m., the proceedings were recessed to be resumed at 1:20 p.m.]

A F T E R N O O N   P R O C E E D I N G S

[1:30 p.m.]

DR. WHALEN:  We are approximately 40 to 45 minute behind.  I would urge everyone to try to keep within the time constraints as have been promulgated.  As soon as the first speaker can come to the podium or is allowed to get at the podium, please feel free to begin.

MS. NORSIGIAN:  Thank you.  My name is Judy Norsigian.  I am Executive Director of Our Bodies, Ourselves and coauthor of the landmark book about women's health and sexuality, also titled, Our Bodies, Ourselves.  Since the early 1970s, over 4 million copies in 18 languages have appeared around the world and we include a section on breast implants which will be updated in our upcoming eighth edition.  My organization has no conflict of interest with regard to this issue.

Our interest in breast implants goes back to the mid-1980s when Ester Rome, my coworker, began to work with hundreds of women, both locally and nationally, who believed that breast implants caused them serious problems.  Some of her interviews with women are reflected in the chapter that she wrote about breast implants in, Sacrificing Ourselves for Love, the book she coauthored just before her death from metastatic breast cancer.

We still continue to hear from women, some of whom were initially very happy with their implants, that many of them have developed debilitating conditions after five or ten or fifteen years.  Everyone knows that silicone-gel implants are more likely to break as they age, so it is essential that the FDA be concerned about what happens to women who have had their implants for at least five or ten or fifteen years.

Inamed's own data show that, even in the first two or three years, there are substantial problems especially for breast cancer patients.  I won't repeat them here, but I do want to underscore the fact that most of the 6 percent diagnosed with rupture had silent rupture and, therefore, leaking can continue for years without notice, as you have already heard today.

Inamed also reported an alarming increase in many symptoms associated with autoimmune diseases for all implant patients including joint pain, fatigue, hair loss and muscle pain.  If this is what is happening at two or three years using the best plastic surgeons that the company could find to participate in their study, what will happen to women across the country using less experienced doctors and with implants for ten or 20 or 30 years.

One of the more disturbing aspects of the ongoing public debate about the risks of silicone implants has been the largely absent voices of physicians who have examined women with breast implants, those doctors who have looked more closely at the joint aches, fatigue, dry mouth and eyes and the masses of silicone that have hardened in up to 70 percent of the women's chests over time.

To repeat Dr. Gary Solomon who has served as Associate Director of Rheumatology at New York University's Hospital for Joint Diseases Orthopedic Institute, "Any physician who has seen a dozen of these women knows there are illnesses caused by implants."

Unfortunately, as you have already heard, their symptoms and problems have not resulted in a clear clinical diagnosis.  This difficulty with nomenclature has made it easier to discredit the women who have been suffering as well as their physicians who are trying to help them.  Many experts agree that studies on silicone implants should include only women who have had implants in at least eight to ten years, since most leakage or rupture occurs after that period of time.

Studies by the FDA's Dr. Lori Brown and her colleagues look at such a targeted patient population and show an increase in fibromyalgia among women with leaking silicone-gel breast implants.  You will hear about that later.

Let us also recall that three years ago more than 190,000 problems from breast implants had been officially reported to the FDA including 123 deaths.  These complaints represented potentially more than 10 percent of the estimated 1.5 million women who then had breast implants.

We should also recall what happened in 1991 and '92 when the breast-implant industry and the American Society of Plastic and Reconstructive surgeons spent approximately $4 million in a PR campaign to try to downplay the potential risks of silicone breast implants and to block regulatory action by the FDA that might restrict access to these implants.

A well-funded publicity effort underscored the greed of trial lawyers involved in breast-implant litigation referred to the medical literature indicating problems as junk science and urged women to phone and write the FDA to protest any regulation of the use of silicone implants.  Rhetoric about giving women a choice was heavily employed and paid advertising spots featured women who wanted their right to choose breast implants.

In the entire media flurry, what was often overlooked was the fact that breast-implant safety had never been demonstrated in over 30 years of use despite the repeated pleas of women's health advocates and scientists at the FDA.  In a climate where cosmetic surgery and breast-implant procedures are heavily advertised and promoted, it is not surprising that augmentation is on the rise despite the mounting evidence of harm.

Many young girls are now given breast jobs as a high-school-graduation gift.  With such a backdrop to these deliberations, it is especially important that FDA exercise extreme caution in the matter of approving a product that could have such long-term negative impact on so many otherwise healthy women.

The FDA, which is charged to protect the public's health should deny approval of any breast implants until better data are available to resolve the ongoing debates about safety.  Choices are great, but we need safe choice.  What kind of choice is it to approve a product with very high complication rates and unknown health risks over time.

I also want to give to you a letter from almost a dozen breast-cancer advocacy organizations who expressed similar views.  They include the Mass Breast Cancer Coalition, other groups around the country.

I just want to add a small footnote that, as a lapsed member of NOW, and I have not been following everything that NOW does, I can tell you that the representation of NOW as espousing a view that women ought not to care about their appearance is just pure hogwash.

Thank you.

DR. BASH:  My name is Dr. Deborah Bash and I am a plastic surgeon at Mayo Clinic in Scottsdale, Arizona.  I am here today representing myself.  I have no financial interest in the manufacturer or any of their competitors.  My travel expenses are reimbursed by Mayo Clinic for my travel today.  I perform both reconstructive and cosmetic breast surgery with breast implants.  I am paid on a fixed salary which is independent of the type and amount of surgery that I perform.

I appreciate the panel providing me the opportunity to testify today.  This is an important issue and I believe I have a unique perspective to offer for consideration.  As a woman who has had a breast-augmentation surgery and a board-certified plastic surgeon, I can speak to both the clinical side and the patient side of these issues.

I decided to have a breast augmentation in 1985 when I was a third-year medical student.  At that time, of course, silicone-gel-filled implants were the most common device used for the procedure and that is what I had.  As a medical student interested in plastic surgeon, I was a particularly well-informed patient.

I knew what to expect and I was very comfortable with the decision that I made.  I believe now, as I did then, that silicone-gel-filled implants are a safe option for women considering breast augmentation.  I have not had any problems with my breast implants and, in fact, in 1999, I chose to have the original implants replaced.

I knew when I had my first procedure that the implants would not last a lifetime, just as any medical device will eventually need a replacement.  After 14 years, I chose to have my implants replaced with new silicone-gel implants rather than saline-filled implants because I had that choice.  And, as a patient and a board-certified plastic surgeon, I knew silicone-gel implants would give me a better cosmetic result.

Since my surgery in 1999, I continue to be a very satisfied patient.  My experience as both a surgeon and a patient puts me in a uniquely credible position to discuss breast reconstruction and breast augmentation with my patients.  I share with every one of my patients my history of breast implants and, in fact, my mother has also had a mastectomy and has silicone-gel implants at my recommendation.

While my experience allows me to relate to patients at the different level, I am concerned because they do not have the same access and choices that I had.  I believe it is important that patients understand the difference between silicone and saline implants.  As part of the informed-consent process I use with all of my patients, I ask patients to touch and feel both types of implants.

I explain the history of the implants and I tell my augmentation patients that, unfortunately, they do not have a choice for saline-gel implants.  In my experience, it is rare that a primary augmentation patient prefers a saline implant after holding and touching both devices.

It has also been my experience that, when I operate on women with aging silicone implants who desire replacement, that they chose to have a saline implant when they have a choice of silicone or saline.  I believe they make this decision based on positive personal experience and informed consent.

It has been my professional and personal experience that breast-augmentation patients should have the same options available that reconstructive patients have.

Thank you very much for your time and consideration.

MS. GILBERT:  Doctor, this is Alisa Gilbert.  Just out of curiosity, do you do any special protocol yourself in follow up or--how do you do your own follow up?  I mean, are you getting MRI, or you haven't had any problem whatsoever.

DR. BASH:  I have regular--my annual breast exams and my mammograms, just like any women would for routine breast care.  I am also on the Mentor silicone-gel adjunct study and I have done my follow up at my one-year, three-years, and then I have my five-years coming up.

MS. GILBERT:  Thank you.

MS. KEITHLEY JOHNSTON:  Good afternoon.  My name is Kathy Keithley Johnson and I am the Executive Director of Toxic Discovery, a national consumer-advocacy organization whose primary goals are education and emotional support for individuals who have sustained injuries from silicone and saline breast implants.

Toxic Discovery presently maintains a database of over 25,000 injured consumers with over 10,000 of those with breast implants.  My expenses of my staff to attend this hearing were paid by private donations from our members and from the injured consumers.  I also would like to submit to you today a collection of individual stories from the women called, We are the Evidence.

My personal journey has well prepared me for my present position as Director of Toxic Delivery.  I am a registered nurse.  I am a former surgical nurse and I am former Director of Hematology Oncology.  But, more important than my nursing background, I am yet another injured consumer.

I am also one of the nurses in the Harvard Women's Study.  I was implanted free of charge in Midland, Texas and was approached to become a spokesperson or poster child for the manufacturers.  I so elected not to take part.

During these hearings, the FDA will be determining once again whether or not to lift current restrictions on the use of silicone implants.  We ask you, the committee, what restrictions have the FDA truly enforced?  The breast-implant manufacturers and plastic surgeons of this nation would like to consumer to believe that silicone implants have been restricted over the years.  We say there has never been a true restriction or protection.

I ask you, who is minding the breast-implant store?  The experts here today are not you, this learned committee member, nor are the experts the breast-implant manufacturers.  The experts here today are not even the plaintiff attorneys or the defense teams.  The experts here today and remain the injured who have suffered in the name of corporate greed.

Mark McClellan, the FDA Commissioner, wrote a white paper a few years ago entitled, Protecting the Public Health; FDA Pursues an Aggressive Enforcement Strategy.  In Mr. McClellan's paper, and I quote, "The agency is committed to pursuing Food, Drug and Cosmetic Act violations."  Toxic Discovery believes that violations are being committed daily via the route of the silicone breast prosthesis.

In view of the supporting evidence and facts, silicone breast implants have never been approved safe.  lifting restrictions would appear purely motivated by special-interest groups and organizations of plastic surgeons and of the breast-implant manufacturers.  Lifting the restriction will allow the unchecked violations to continue.

It was Abraham Lincoln from my sister state of  Illinois that said, "Sin by silence when they should protest make cowards of men."  The breast-implant women of this world should no longer suffer in silence.  The women and children of this continued medical experimentation will remain the evidence with or without the FDA's so-called approval.

No committee including this one can change those facts.  In Missouri, we were fortunate in January 2000 to pass a law concerning informed consent directly related to breast implants.  The only other state that has such a law is the one I am visiting today, the Great State of Maryland.

I stand before you as injured consumer but one that knows the truth of product failure and corporate lies.  I hear you speak of research and yet you turn away from the truth.  We, the injured consumers, continue to be your research.  We, the injured consumers, are your evidence and we will remain just that, evidence that not medical device is totally safe but especially the new PMA application that Inamed would lead you to believe that is new and improved.

Informed consent should being with informed individuals.  Freedom of choice versus informed consent implies that women are aware of the consequences of their actions.  But the women requesting breast implants are not provided the truth and continue to be the victims of social pressure and medical fraud.

The right to choose is meaningless without the right to know.  The poor unsuspecting cancer patient has received a double insult to injury.  They have, in a large majority of cases, survived cancer to only learn in the years to follow that the very device that they believed would make them whole again would later cause further mutilation.

Let us also not forget that the largest contributors to the majority of cancer organizations are the breast-implant manufacturers and corporations.  As a former director of oncology, I know this very up-front and personal.  Do you really want to hand over a legacy of a new generation of a breast implant to the--

DR. WHALEN:  Can you conclude, ma'am?

MS. KEITHLEY JOHNSTON: --to the consumers who were so irresponsible to thousands?  May I say that I gave up five minutes for Allen Keithley this morning who was unable to attend.

DR. WHALEN:  Ma'am, you have five minutes and you are 52 seconds beyond that now.  If you could conclude, please.

MS. KEITHLEY JOHNSTON:  Women historically have been the cash cows of the medical profession.  Women continue to be spoon-fed from the cradle to the grave how they should look by a media-funded, manufacturer-driven, silicone cartel.  It is beyond time that we, the consumers, demand a safe breast implant.

We, the breast-implant women, demand from the FDA that you do not approve Inamed's application and we ask you to defer approval until only a safe implant can be offered.  The injured masses call on you, the committee, to save the next generation and the second generation of women that breast-fed their children.  Do not allow the silicone holocaust to continue.

MS. GILBERT:  Ms. Keithley, I have a question for you.  Alisa Gilbert.

MS. KEITHLEY JOHNSTON:  Yes, ma'am.

MS. GILBERT:  How many of the you said 25,000 patients or consumers that you--

MS. KEITHLEY JOHNSTON:  We have 25,000 members.

MS. GILBERT:  How many of those are breast-cancer survivors?

MS. KEITHLEY JOHNSTON:  Out of that statistic, I would say only probably 25 percent since I am an oncology nurse, or was a former oncology nurse.  Most of that is known and we do open our doors.  That is our first love is the cancer patient.

MS. GILBERT:  So you provide your main--the support that you provide to the participants that you represent is specifically support or--

MS. KEITHLEY JOHNSTON:  We are the only organization that I know of in the United States for breast-implant survivors and cancer survivors that had breast implants that failed that you actually get to speak to one of our people.  All of our registered nurses that answer the calls all had breast implants, they all failed and they all had them removed.  I have never received one nickel from the breast-implant manufacturer in any litigation nor, also, do I take a check from Toxic Discovery.  I do this from a passion.

DR. PAUL WEISS:  Good afternoon.  My name is Dr. Paul Weiss and I am a practicing plastic surgeon in New York City.  I am here representing the National Endowment for Plastic Surgery and thank you for the opportunity to present today.  I have no financial affiliations with the manufacturer, its product or its direct competitors.

As part of my practice of plastic surgery, I perform breast-implant surgery both for reconstructive and aesthetic indications.  I, therefore, derive a portion of my income from this type of surgery.  The American Society of Plastic Surgeons is paying my travel expenses.

The National Endowment for Plastic Surgery was established in 1994 to provide an ongoing secure funding source for specialty-based research projects.  The Endowment was established by and is maintained by the Plastic Surgery Education Foundation, the research, education, domestic and international service arm of the American Society of Plastic Surgeons.  The Endowment represents an unprecedented specialtywide cooperative effort and is funded solely through voluntary contributions of plastic surgeons and grateful patients.

The National Endowment is very important to organized plastic surgery and was created to support initiatives that are relevant to the immediate day-to-day clinical practice of plastic surgery.  In 1991 and 1992, the manufacturers faced the FDA's call for the PMAs for silicone-gel-filled breast implants.  The impact of this request created a crisis for implant recipients.  From that experience came the realization that, as a specialty, we needed to be much more proactive in our research efforts.

With a commitment to our patients, sound science and quality research data, organized plastic surgery established the Endowment as a perpetual source of research funding for the specialty.  Its mission includes supporting research directed toward defining the efficacy and safety of established plastic-surgical procedures, investigating innovative technology and plastic surgery, sponsoring research scholarships and fellowships and charitable programs in plastic surgery.

The review process used by the endowment for awarding research grants is rigorous.  The process and the application form are modeled after those of the National Institutes of Health.  The grant awards are selected by an expert panel of clinical and scientific researchers, many of whom have served on NIH review panels or are recipients of NIH grants.

While the Endowment is in the early years of awarding research grants, I would like to share with you information on the implant research funded to date.  In 2001 and 2002, the National Endowment provided funding to establish and conduct research at the Center for Implant Retrieval and Analysis at Washington University in St. Louis, Missouri, and Bioaesthetic Plastic Surgery Center, also in St. Louis.

The emphasis of this study was to determine the modes, causes and rates of implant failure.  The study determined and demonstrated the types of scientific analyses needed for each type of implantable device used by plastic and reconstructive surgeons.  The analyses were macroscopic, microscopic, mechanical, chemical and physical.

A significant outcome of this research is that a standardized protocol has been developed which gives surgeons a clear approach for management of explanted breast implants.

Another research project funded by the Endowment is a two-year study that measured the analysis of physical benefits of breast-reduction surgery.  Plastic surgeons and our patients have faced numerous challenges in recent years.  As a specialty, we are at the forefront of both reconstructive plastic surgery and aesthetic plastic surgery.  Perhaps, most importantly, we are more strongly committed than ever to facilitating and conduction good research and being more responsive as advocates for our patients.

In conclusion, the studies funded through the National Endowment for Plastic Surgery will provide information that will help determine best practices and deliver safe and effective care to plastic-surgery patients.  We are confident that the FDA will review the scientific data on silicone-gel-filled breast implants and make its decisions based on the best interests of patients.

Plastic surgeons want safe products for their patients and fully support the FDA's ongoing regulatory process for breast implants.  Plastic surgeons want good quality research data to drive the clinical decision-making.

Thank you.

DR. McGRATH:  May I ask you a question, Dr. Weiss?

DR. WEISS:  Yes.

DR. McGRATH:  You said that you founded the Endowment in 1994.  Ms. Kim Gandy, this morning, from NOW, made the comment or the statement that--I believe she said large sums of money have been spent by plastic surgeons to convince patients that implants can be used safely.

Could you tell us how you have used Endowment moneys, plastic-surgery moneys, for that purpose, please.

DR. WEISS:  All of the Endowment funds are used to fund research projects and nothing more.  There is no advertising or media money spent by the Endowment.

DR. MANNO:  Barbara Manno.  Doctor, have these studies been published in peer-reviewed journals, any of the data?

DR. WEISS:  The one study on the effects of Betadine in breast implants has been published but mostly the work that has been done regarding breast implants has been to establish protocols for their study rather than reaching conclusions at this time.

DR. CASAS:  Good morning.  My name is Dr. Laurie Casas.  Before I begin, I would like to let the panel know that the relevant statistics for my testimony can be found stapled to the copies of my testimony.

I am Dr. Laurie Casas.  I have no financial ties to any implant manufacturer.  I am neither a witness nor a party in an impending law suit.  I am Communications Commissioner for the American Society for Aesthetic Plastic Surgery, also called ASAPS.  My travel expenses to this hearing were paid for by ASAPS.

I am an associate professor of surgery at Northwestern University Feinberg School of Medicine.  I have published extensively on breast-surgery topics.  I am currently participating in an IRB-approved multi-site prospective outcome study on patient satisfaction following cosmetic procedures.  In addition, I am a board-certified plastic surgeon in clinical practice for more than 13 years.  Therefore, a portion of my income is derived from breast-implant surgery.

Plastic surgeons continue to fully support the FDA's ongoing regulatory process regarding breast implants.  The FDA, the medical literature, plastic surgeons and their patients have raised valid concerns about these devices and, for the first time, we have data to respond to these specific concerns.

The results of two surveys of more than 5,000 women with or considering breast implants will be published in the Aesthetic Surgery Journal, ASAPS' peer-reviewed clinical publication.  In these two retrospective outcome studies, these 5,000 women were anonymously surveyed in two separate questionnaires.

The first study surveyed 2,129 women with implants of which 13 percent had silicone gel and 1,625 who were considering implants.  They were asked questions related to their demographics, compliance with follow-up appointments, body image, motivation, reoperation and quality of life.

The second study surveyed 1,359 women with breast implants of which 13 percent had silicone gel.  They were asked questions regarding why women may not return for follow-up visit, the adequacy of informed consent, their ability to get health insurance, and overall satisfaction with breast augmentation.

The research was funded by the nonprofit Aesthetic Surgery Education and Research Foundation, ASERF.  Each survey was posted on the independent website implantinfo.com over a period of months.  The on-line surveys were implemented by professional research firms with protection in place to ensure respondents anonymity and to prevent multiple survey submissions.  Neither ASAPS, ASERF nor implantinfo.com had access to the data before it was analyzed.  The margin of error is plus-or-minus 3 percent at a 95 percent confidence interval.

Both surveys showed that women's motivation for having their breasts enlarged primarily relates to their sense of self.  In the first survey, the top reasons for choosing breast augmentation for 79 percent or more of the respondents were; to look better in clothes, to feel better about themselves, to feel more confident and to feel less self-conscious.  Pleasing others was one of their lowest priorities.

Results of both surveys showed that women's expectations were met by the procedure.  92 percent said they were happy about their decision to get breast implants.  89 percent said the augmentation completely or mostly met their expectations and 94 percent said they would recommend breast augmentation to their friends or family.  And if they choose, based on what they know today, would women with implants in this survey choose to get them again?  84 percent said they would very likely do so and 82 percent reported an improvement in self-confidence.

The outcomes research also shows that women considering augmentation were not impulsive about their decision.  77 percent considered augmentation for at least a year before undergoing surgery and more than half thought about it for three or more years.

The demographics from this data was very interesting.  It showed that the median age of the respondents was 34.  75 percent are either married or in long-term relationships.  84 percent had at least some college education and 45 percent at least had a Bachelor's degree.  And the respondents came from every state as well as Puerto Rico and the Virgin Islands

The 2,800 members and candidates of the American Society for Aesthetic Plastic Surgery fully support the FDA's ongoing review of implant safety and efficacy and rely upon the agency's evaluation of the current science regarding silicone products.

MS. GILBERT:  Excuse me.  I have a question.  Alisa Gilbert.  So you had 5,000 women in the study.

DR. CASAS:  Yes.

MS. GILBERT:  Where were they on this scale?  A year after their surgery?  I know that you said that some of them were considering surgery and some of them weren't.  Do you have any kind of--

DR. CASAS:  We have women from the '80s to 2002 that were augmented.

MS. GILBERT:  Okay.  How do these numbers reflect that?  I mean, so 91 percent wanted to look better in clothes or 92 percent said they were happy about the decision to get breast implants, but I guess what I want to know is how many of them were their first year out, their second year out, from implant surgery.

DR. CASAS:  We don't have the data broken up that way.  We have it--Dr. Leroy Young will complete the data for you and then, if he can speak to that, he will.

MS. GILBERT:  Okay.  Thank you.

MS. BROWN:  This is Debbie Brown.  I had a follow-up question.  Of the questionnaires that you sent out, what was your percent respondents?

DR. CASAS:  It wasn't a send-out questionnaire.  It was an on-line survey.

MS. BROWN:  And so did everybody respond or what was your percent respondents?

DR. CASAS:  Well, you mean the people that started the survey, did they finish the survey?

MS. BROWN:  Yes.

DR. CASAS:  Because once you sign on, then you complete the survey.  I don't have that information.

DR. YOUNG:  Approximately 85 percent of people that started the survey finished the complete survey.

MS. BROWN:  Thank you.

DR. YOUNG:  I am Dr. Leroy Young.  I have no financial ties to any implant manufacturer and am neither a witness nor a party in a pending law suit.  I am testifying for the Aesthetic Surgery Education and Research Foundation, also known as ASERF.  My travel expenses were paid by ASERF.  The Foundation funded the research that I will be telling you about concerning two surveys of more than 5,000 women with or considering breast implants.

I am a board-certified plastic surgeon and a portion of my income is derived from breast-implant surgery.  In the past, I have received research funding from McGhan, now Inamed, and have served as a consult to breast-implant manufacturers.  I am Chairman of the National and Co-Chair of the International Breast Implant Registries.  I am currently Co-Chair of the Breast Surgery Committee of the American Society for Aesthetic Plastic Surgery also known as ASAPS.  Results of my research will be published in ASAPS' peer-reviewed clinical publication.

Plastic surgeons have shared the FDA's concerns about certain issues raised in previous breast-implant hearings which is why I undertook the two surveys I will be speaking about today.  One of the most important issues raised previously was whether breast-augmentation patients have been adequately informed about risk associated with breast implants.

In the second survey, of 1,350 breast-augmentation patients, more than 90 percent of women remembered being provided with important information such as what to expect after surgery, specific complications, breast-implant deflation or rupture, and that breast implants do not last forever.  88 percent reported being actively involved in the decision about the size of their implants.

We also explored the issue of reoperation.  In the first survey of more than 2,000 women with implants, only 13 percent reported a revision and nearly all of them chose implant replacement, not permanent removal.  Most of the reasons for revisional surgery were cosmetic in nature such as the desire for larger implants.

Many women experienced breast pain associated with their menstrual cycles, nursing or other reasons.  Surprisingly, a larger percentage of women without implants reported breast pain than women with implants.  We also learned that 93 percent of breast-augmentation patients had all of their children before undergoing to procedure so nursing was not an issue.  Of the small percentage who had children and nursed them after augmentation, there was essentially no difference in the rate of reported problems.

There have been concerns about whether women with breast implants are being denied health insurance.  The survey results indicate that more than 90 percent of women with implants have insurance.  Of those who applied for insurance after receiving breast implants, 91 percent were able to obtain insurance.  A small number of women were charged at a premium and a number of others had a two-year exclusion on their policy for breast-related problems.  The margin of error of the surveys I am talking about is a plus-or-minus 3 percent at a 95 percent confidence interval.

In the first survey, we have data covering the past ten years of patients' experience with breast-augmentation.  There were more than 2,000 patients with saline and more than 250 with gel implants.  This is an expected distribution given that silicone was not on the general market.

We have some preliminary data on comparisons between the devices regarding patient outcomes.  For some findings, there are no statistically significant differences: fulfillment of expectations, whether patients would recommend the procedure, breast pain severity and frequency, asymmetry, skin wrinkling, hematoma and infection.

Other data indicate statistically significant differences between saline and gel regarding patient outcomes.  Examples are natural feel of the breast, capsular contracture and reoperation.   We plan to make this data available to the FDA as soon as analysis is complete.

Plastic surgeons are concerned that a small percentage of women say that breast augmentation does not meet their goals.  We are gratified that these results demonstrate an overwhelming number of women report satisfaction with the procedure and we are constantly striving to make women's experience with breast implants meet their expectations.

We fully support the FDA's ongoing scientific review of implant data and I thank you for the opportunity to present.

DR. BURKE:  I am Dr. Martha Burke.  I am Chair of the National Council of Women's Organizations.  We are the nation's oldest and largest coalition of women's groups representing 170 organizations and 10 million women collectively.  I have been paid by no one to appear today.  I am here representing the American women from all socioeconomic, ethnic and income groups that my groups represent.

We work on a wide variety of issues and, until now, we have never testified before the FDA.  But, after hearing the stories of women who have been harmed by breast implants, we are here to show our concern for them and for women across the country who might be considering the implants.

Some argue that breast implants have already been studied more than any medical device.  If that is true, perhaps there is a good reason.  More than $2 billion in legal settlements might influence the need for research and the incentive for companies to fund such research.  Unfortunately, even $2 billion cannot guarantee good-quality research.

I am here to let you know as panelists that American women who depend on the FDA to safeguard their health and safety are watching.  We are not here to pressure you to vote a certain way but we are here to urge you to carefully scrutinize the safety of this product and answer these two key questions.

If a product with high complication rates is approved, will the company that makes that product have an incentive to improve it?  And if a company provides inadequate long-term safety data and the product is improved anyway, will the company have an incentive to improve their research?

I would like to tell you about one of the many women who wanted to testify here today but could not come, one being Sharyn Noakes who is 55 years of age and she is too ill to be here.  She had two sets of silicone-gel implants.  The first was removed after only eight months due to extreme capsular contracture.  When the surgeon opened her up, the implants were both ruptured.  They were replaced with new implants that she kept for 24 years.  And this is what she had to say.

"I had the implants removed when the pain and deformity became too much for me.  The surgery that should have taken 90 minutes took over three-and-one-half hours.  The implants had ruptured.  The silicone and scar tissue had calcified and embedded into my chest.  I was rushed to the hospital next door."  This was apparently even outpatient surgery.

"I did not have cancer but I had to undergo surgery similar to a double mastectomy because of the silicone.  I am forced to live deformed and I am very ill with several autoimmune disease.  I believe there are many women like myself who will suffer when ruptured implants stay in their bodies.

"We are not like an automobile having our oil changed.  Surgery is expensive and many women won't have enough money to have the implants removed.  I am now on 13 different medications.  Fortunately, my husband is retired military and I am able to get my medication free.  Otherwise, it would cost me $1,500 a month to stay alive.

"Would you like to see what your wife, daughter, mother or girlfriend could look like after having the implants for 24 years?  Please look at my picture."

Sharyn's story reminds me that the Center for Medicare and Medicaid Services is suing several breast-implant manufacturers to recover millions of dollars that Medicare had paid out in healthcare for older women harmed by breast implants.  It doesn't make sense for one government agency to sue for healthcare costs while a different agency makes the same product even more available.

I hope some of you will ask Medicare how they made this decision to sue.  Do they know something that this panel needs to know?  Women should have many options in their lives including safe medical products.  Unfortunately, Inamed did not study the health risk or the cosmetic risk involved in this product.  As Sharyn Noakes' testimony shows, the cosmetic and health effects can be devastating.

In closing, I remind you, if the FDA improves silicone-gel implants, they can't require informed consent because the FDA cannot enforce that.  The FDA can't restrict the gel implants to cancer patients or to women over 18 because the FDA cannot enforce that.

The FDA can only make meaningful restrictions on access if this panel decides that the company has not proven that the implants are safe.  As you listen to testimony today, vote for approval only if you are confident that these implants are safe for long-term use.

Thank you.

DR. YOUNG:  Good afternoon.  I am Dr. Laurie Young, not to be confused with Dr. Leroy Young who preceded me.

I should first clarify that I am not paid by anyone who is involved in any kind of litigation or involved in the promotion of anything related to silicone gel breast implants.

I am the executive director of the Older Women's League.  The voice of midlife and older women is the only national grass-roots membership organization focused solely on issues affecting midlife and older women.

As you know, the risk of breast cancer increases with age and since our mission is to improve the lives of women over the age of 40, we are very concerned about the safety of silicone gel breast implants.

We know that women with breast cancer want to have choices.  Fortunately, however, these women have more options than ever before.  Most will not need a mastectomy and if they choose to have one, saline breast implants and other reconstruction surgeries offer alternatives to silicone gel breast implants, so the standard for approval should be are these that are being considered by a panel today safe for women to use.

We know that many women with breast cancer who were happy with their implants at first, have problems that arise and are entirely consistent with the high rate of complications that Inamed will be reporting to you today, such as 46 percent need additional surgery within two to three years, 25 percent have had their implants removed, 6 percent had a ruptured implant, 6 percent had moderate to severe breast pain, and 6 percent had necrosis which means that the skin or breast tissue around them died.

Everyone knows that silicone gel implants are more likely to break as they age.  When silicone gel implants break, the silicone can leak into the lymph nodes and other organs and unfortunately, Inamed did not study the health risks involved.

If there are so many complications in the first two or three years, what will happen to women with implants for 10 years or more?  Most women in their 40s or 50s, we know now will live another 25 to 40 years.

I would like to tell you about one of the many women who wanted to testify here today, but were unable to come.  Susan Nicewonger lives in St. Louis.  She received implants more than 20 years ago for augmentation.  This is what she says:

"Over time, my breasts looked and felt as if I had two croquet balls for breasts.  I will always remember how humiliated I felt when my gynecologist performed my annual breast exam.  Her facial expression revealed disgust at having to examine my breasts.  I actually had to request that procedure from her each year.  Clearly, she wanted to avoid it.

"Sleeping on my abdomen was not an option, too painful.  Soreness was a way of life.  Anything that required air movement, hence, breast movement, was painful. It was when I began to notice a sound similar to scrunching cellophane when my right breast moved that I knew something was terribly wrong.

"A trip to the doctor and a subsequent MRI confirmed that I had a ruptured silicone implant and surgery would be necessary.  My implants were removed this past May. Because of the leakage, breast tissue had to be removed from the right breast which has left me disfigured and unable to wear most bathing suits or any neckline that would reveal the sunken half of my right breast.

"I am, however, extremely grateful that those foreign, hard, painful objects have been removed from my body.  I can't be certain of the path that migrating silicone has taken inside my body or how it will eventually manifest, but it is on my mind every single day of my life."

I want to echo the comments expressed by Dr. Burke who said that the CMS, Medicaid and Medicare are now exploring legal methods to recover funds for the surgeries and medical treatments that women have had to experience as a result of these implants.

Again, I ask you today to be concerned and serious, and not to make judgments based on inadequate or studies that are not complete.  Women deserve complete information, if we are to be truly informed, it can't be on half-baked information.

Thank you.

MS. LOZIER:  (Read by Nancy Nealand.)  Good afternoon.  My name is Nancy Nealand and I am here reading for Christine Lozier from Hackensack, New Jersey.

"My name is Christine Lozier and I am from New Jersey.  I have no conflicts of interest.  I got my first set of silicone gel breast implants in the 1980s, but my breasts hardened soon afterwards and I had to have them replaced two years later.

"Cosmetically, the results of the second implantation were very good, but then, five years later, I developed autoimmune symptoms and other health problems I had never had before, like asthma, Raynaud's syndrome, and fibromyalgia.

"After living with my debilitating health problems for over a decade, I decided three years ago to have an MRI. It detected that the left implant had ruptured and was leaking.  When I went to a doctor to get both implants removed, it was discovered that my right implant was also leaking.  As a result of the leak, I had to have a partial mastectomy in my right breast to remove silicone, not cancer.

"I bled severely and had drains for two weeks.  Four months later, I underwent reconstructive surgery and had saline implants put in.  The right one deflated two months afterwards.  I had it replaced and two months later developed an infection and a rash on my left side.

"Because of the rash, I had to have both saline implants removed and replaced with new ones.  It was during this surgery that my doctors discovered silicone that had escaped my left implant during my very first explantation and was still in my body.

"Since my first explantation, I have undergone a total of six surgeries, and each time I have needed to have drains put in because of severe bleeding and infection.  It has been one year since my last surgery, and I still have pain in my left side from the silicone under my skin that cannot be removed.

"Each one of my reconstructions was paid for out of pocket and overall I have spent $60,000 trying to get my body back to normal.  Each time I have needed to be out of work for two weeks at a time and even though the silicone has caused severe health problems for me, I am stigmatized by my co-workers because I had implants and people have no sympathy for my condition.

"None of my doctors ever warned me of the possible effect silicone implants could have on my body and now that I am sick, they deny that silicone had anything to do with my illnesses.  I feel betrayed and lied to and I am telling my story because I don't want future women to have to go through what I have been through.

"Please, don't allow silicone gel breast implants to go back on the market unless they are proven to never leak.  If you require real proof of safety, I have no doubt that implant makers will develop a safer implant.  If you don't, they will continue to sell the kinds of implants that have harmed so many of us.

"Thank you."

DR. KASPER:  My name is Dr. Anne Kasper and I came here at my own expense.  I have no conflicts of interest.

I am a partner in Finding My Way, a support service for breast cancer patients that is in Bethesda, Maryland.  I am also the senior editor and an author of the book Breast Cancer: Society, Shapes, and Epidemic.

Breast cancer patients have more choices than ever before and while it is important to have choices, and tell this to all the women in our practice, when a woman is diagnosed with breast cancer, the choices are often very difficult to make.

The good news today is that three out of four women newly diagnosed with breast cancer can choose to have a lumpectomy with radiation instead of a mastectomy, and if a woman needs to wants a mastectomy, she can choose saline breast implants, silicone gel implants, or autologous tissue transfer procedures.

How important is it that silicone gel breast implants be widely available to cancer patients?  Under the current restrictions, the vast majority of women who want silicone gel implants for reconstruction are able to get them, but breast cancer patients are cautious, they want to live long lives and they have faced a life-threatening disease.

They also, however, are concerned about cosmesis and how they will look after cancer treatment.  Many will choose lumpectomy which allows them to keep their breasts, preserving their appearance and the sensory feelings of the breasts, but what women tell us--and this is the take-home message from my talking with you today--is that they only want silicone gel implants if they can be assured that they are safe.

Women are often encouraged to choose implants and sometimes a rosy picture is painted about the benefits of implant reconstruction.  In our practice, women are often disappointed to have reconstructed breasts that have none of the sensory, sexual, and maternal capacities of a normal breast and that often have complications.

Many reconstruction patients who choose silicone gel implants are happy with their implants at first, but unfortunately, many then experience the kinds of complications that you have heard today and that Inamed has found, which include 46 percent will need additional surgery, 25 percent will have their implants removed, 6 percent will have a ruptured implant, 6 percent will have moderate to severe breast pain, and 6 will suffer from necrosis.

In the Inamed core study of breast cancer patients, only 6 percent of the implants broke during the first two or three years, but over time we know that that percentage increases dramatically, and it is, of course, much more serious than when a saline implant breaks.

As you know, when silicone gel implants break, silicone can leak into the lungs, the lymph nodes, and other organs.

To determine the safety of silicone gel implants, Inamed should have studied the health of women whose implants broke especially since cancer patients may have compromised immune systems, unfortunately, they did not.

Another shortcoming of the Inamed research is that it only included African-American breast cancer patients in the core study and only five Asian-American women.  Obviously, this is not sufficient data on which FDA would make safety decisions for all American women.

Both of those groups have approximately the same rates of cancer as white women, and African-American women are especially susceptible to autoimmune disease.  All of these groups of women must be studied adequately to determine if implants are safe for them.

The bottom line, though, is that the complication rates for breast cancer patients are very high in just the first two or three years.  What is ahead for these women in five years, 10 years, or even 20 years?  Most breast cancer patients are living longer lives today and many choose mastectomy as a way to put the cancer behind them.

With implant reoperation rates of 46 percent in the first two or three years alone, and higher rupture rates in the years to come, many women will not consider implants a satisfactory and safe choice at all.

It is unusual to ask for more than two years of safety data for a new product, but silicone gel implants are not new.  Inamed is asking for approval for an implant that has not changed much, if at all, in more than a dozen years.

If FDA had strictly enforced restrictions on breast implants established in 1992, we would now have 10- year data rather than two or three years.

If FDA approves silicone gel implants based on the data that has been provided, it sends the message they are not as concerned as women have a right to expect about the health and safety of breast cancer patients and especially women of color.

If FDA allows silicone gel implants to remain on the market, but only as part of clinical trials, and this time requires a third party to monitor the studies, then perhaps breast cancer patients will finally get the information they need to make a meaningful choice.

Thank you.

DR. ANDERSON:  I have a question.  You have made the statement that cancer patients have these options and that silicone implants are widely available.  I work in the Seattle community and actually I don't know of the availability of silicone implants.

It would be an unusual plastic surgeon that would offer it as part of reconstruction, so I am curious what the basis of your statement that the silicone is easily available.

DR. KASPER:  Well, the ones that we see here in the Washington, D.C. area have been able to get the silicone gel implants when they wanted them.  I am basing my testimony on our practice.

DR. ANDERSON:  Thank you.

DR. WILLEY:  Members of the General and Plastic Surgery Devices Panel, I am Shawna Willey.  I am a medical doctor and a Fellow of the American College of Surgeons.  I practice in Washington, D.C., general surgery with an emphasis on breast surgery.  I am an Assistant Professor at Georgetown University.

I am proud to represent the 64,000 fellows at the American College of Surgeons and thank you for the opportunity to offer a general surgeon's perspective on the issue of silicone gel-filled breast prosthesis for patients who are considering breast reconstruction or augmentation procedures.

I have no financial relationship with any medical device manufacturers.

Today, I would like to focus my comments on breast reconstruction and augmentation as procedures that improve the quality of life for our patients.  Reconstruction of a breast that has been removed due to cancer is one of the most valuable surgical procedures available today.  New medical techniques and devices have made it possible for surgeons to create a breast that can come close in form and appearance to matching a natural breast.

Frequently, reconstruction is possible immediately following breast removal or mastectomy, so the patient wakes up with a breast already in place having been spared the experience of seeing herself with no breast at all.

The availability of breast reconstruction gives the woman an option over which she has some control at a very emotional time.  Both breast reconstruction and augmentation provide enormous benefits to women including an improved self-image and a sense of normalcy, but as is true with any operation, a patient must think carefully about her expectations and discuss them with her surgeon.

The American College of Surgeons has a long-standing commitment to improving the quality of surgical patient care and to enhancing patient safety.  Therefore, we are eager to hear more about the data collected by the manufacturer on the safety of silicone gel-filled breast implants and we welcome the FDA's review of these safety data.

If these devices prove to be unsafe, we, of course, would not support their use in patients.  Patient is always, and will always be, the primary concern of the College, but we also feel strongly that decisions that effectively limit patient options must be made on the basis of solid scientific evidence.

We are concerned that procedures with benefits that are life-enhancing rather than life-saving are occasionally viewed by society as somehow less important. Breast implants should be regarded as any other medical device that is being considered by this panel.  Further, distinction should not be made between the use for reconstruction or augmentation purposes.  If the device is safe in one group, it should be safe in the other.

Following more than 10 years of clinical studies and trials, if conclusive data are not found showing that these devices pose a significant risk to patients, we believe the choice to use them should be made by consenting patients.

The College also agrees that it is vital for any woman considering a breast implant procedure to be well informed of all possible risks associated with the procedure, as well as with the device being used.

Women must be informed, for example, that breast implants do not last forever, that complications can occur, and that additional operations may be necessary. Consequently, we support this panel's continuing efforts to ensure that breast implant patients have access to comprehensive and comprehensible material that can be used by patients to make informed decisions about these procedures.

To conclude, let me reiterate that all surgeons are obliged to be scrupulous in ensuring their patients are provided with the information they need to give truly informed consent for any operation.  We are certainly hopeful that the information collected by the manufacturer of these devices will contribute significantly to these efforts and help resolve the continuing uncertainty about the safety of these devices.

Thank you again for the opportunity to speak today.

MS. VOLPE:  Thank you for allowing me to present this statement.  I am Margaret Volpe, a breast cancer survivor with a silicone breast implant.  I have no financial ties to manufacturers or health care providers and am not being reimbursed for my appearance here.

I am a volunteer representing Y-ME National Breast Cancer Organization.  Y-ME has received a minuscule amount of funding from manufacturers, but none in the past 18 months.  Y-ME is committed to providing support and accurate information to empower individuals touched by breast cancer so they can select the most appropriate options in conjunction with their health care provider.

W-ME's comments concern silicone implants for breast reconstruction or other medical need.  We have no opinion on silicone implants for cosmetic purposes.  We are not endorsing a product or manufacturer, but feel strongly that silicone breast implants should be more widely available as an option for women requiring reconstruction.

When I was diagnosed with breast cancer in 1995, I faced fears and anxiety common to those diagnosed with a life-threatening illness.  It was important to have reconstruction, to feel whole again, to not be constantly reminded of my breast cancer and to get on with my life.

I have had no problems or complications with my implant since my reconstructive surgery in February 1996.  Options for reconstruction are limited.  There is tram flap reconstruction, major abdominal surgery with a painful recovery period, leaving us without our abdominal muscles. Women who are too thin do not have this option.

If a woman has a tram reconstruction on one breast, she is unable to have a second tram at a later date. The latissimus dorsi reconstruction usually requires an implant.  Saline implants are available, but for many they are unsatisfactory.  A silicone implant gives a more natural look and feel.

Y-ME emphasizes the need for a range of treatment options, so that each woman can choose what best fulfills her needs.

The availability of silicone implants is important to women with breast cancer.  It was difficult to get my implant because of FDA restrictions.  I signed a lengthy and detailed informed consent.  Even though a breast cancer patient theoretically can obtain silicone implants through a clinical trial, it is difficult for many women to enroll in a trial if the surgeon is unaware of or unable to participate in a trial.

Please adhere to the science when evaluating silicone breast implants.  The Institute of Medicine's exhaustive and definitive review of existing research found no evidence that silicone breast implants cause cancer or disease and no convincing evidence that silicone produces a immunologic response.

The report states such diseases or conditions are no more common in women with breast implants than in women without them.  U.S. Court's National Science Panel and several European Government Scientific Panels issued similar findings.

NIH recently reported to Congress that a large meta-analysis published in the New England Journal of Medicine concluded there was no evidence of an association of silicone gel-filled breast implants with connective tissue diseases or other autoimmune or rheumatic conditions.

Y-ME urges women to fully understand the risks and benefits of any medical or surgical choice.  Women should understand that no medical device lasts forever. Shunts, pacemakers, and artificial knees and joints have an expected life span and possible complications including capsular contracture.

Women must be aware of such side effects as capsular contractions, potential rupture, and the need for replacement.  Doctors must provide accurate information about the risks and benefits of silicone implants because informed consent is key.

As stated in the letter from several cancer and nonprofit health organizations, there should be long-term follow-up and data collected.  Privacy-protected patient registries and physician tracking are vital along with timely communication of significant findings to consumers, patients and providers.

In an effort to resume their life after a devastating disease, breast cancer survivors must be assured the right to select appropriate and effective medical therapies or devices.

Y-ME urges you to consider the plight of these women, but in the end, to act based on the science alone.

Thank you very much.

DR. WHALEN:  Thank you and on behalf of the panel, I would like to thank everyone who testified today for either their societies, organizations, or a individuals for taking time out of their schedules to talk with us.  There will be further public testimony as we initiate our activities tomorrow morning.

Before I go on to Dr. Witten, just briefly, there will be a slight change in schedule due to our lagging behind somewhat.  Dr. Witten will be giving a presentation on Regulatory Overview of Silicone Gel-Filled Breast Implants and we will then take the afternoon break following which the sponsor will make their presentation.

Dr. Witten is the director of the Division of General and Restorative Devices.

Dr. Witten.

Regulatory Overview of Silicone Gel-Filled

Breast Implants

DR. WITTEN:  Thank you for the introduction.

Good afternoon.  I would like to welcome everyone to this meeting of the General and Plastic Surgery Devices Panel.  I am Dr. Celia Witten, Division Director of the Division of General, Restorative, and Neurological Devices in the Office of Device Evaluation at FDA.

I will talk about why we convened this meeting, what we are asking you, our distinguished panel, to do and provide the regulatory background for silicone gel-filled breast implants.

We have convened this panel meeting to obtain your recommendations on whether there is a reasonable assurance of safety and effectiveness for the Inamed silicone gel-filled breast implant.  We ask that you base your recommendations on the valid scientific evidence contained in the premarket approval or PMA application for this device and on your scientific and clinical experience.

The FDA will listen to your deliberations and recommendations and all the evidence presented here today. After the panel meeting, FDA will continue to review the information contained in this PMA and we will consider the recommendations of this distinguished panel.

Ultimately, FDA will make a decision on the approvability of this medical device.

Now, I will move on to talk about the regulatory background for these devices.  At the time of the medical device amendments in 1976, silicone gel-filled breast implants were already being marketed.  FDA classified these implants as Class III in 1988, which means that they would require premarket approval in order to be sold.

However, like other devices that were grandfathered under the 1976 Medical Device law, they were allowed to remain on the market until FDA issued a rule calling for safety and effectiveness information.  This occurred in April 1991.

At that point, manufacturers had 90 days to submit PMAs with data demonstrating the safety and effectiveness of their implants.  Several PMAs for silicone gel-filled breast implants were submitted.

In November 1991, an advisory panel meeting was held to discuss the silicone gel-filled breast implant PMAs.  A voluntary moratorium was issued after that panel meeting, requesting that it cease marketing of silicone gel-filled breast implants while FDA reviewed the safety and effectiveness information that was submitted.

In February 1992, a second advisory panel meeting was held to discuss that new information.  In April 1992, FDA announced its decision that no sponsor had submitted adequate safety and effectiveness data to support approval of their silicone gel-filled breast implant.

Accordingly, silicone gel-filled breast implants became investigational devices at this point, however, FDA determined that there was a public health need for breast implants for reconstruction and revision patients.

To address this public health need, FDA approved adjunct studies designed to allow for continued access of silicone gel-filled breast implants for reconstruction and revision patients.

Inamed, in 1998, and Mentor in 1992, submitted protocols for an adjunct study and received approval.  These studies currently remain active.

At the current time, FDA has not approved any PMAs for silicone gel-filled breast implants.  They remain investigational devices, which means that today, a patient must be enrolled in an adjunct study or an IDE study in order to receive one of these implants.

I should add that in 1999, FDA issued a final rule calling for saline-filled breast implant PMAs, and in May 2000, two of these saline-filled breast implant PMAs were approved.

FDA has provided guidance to manufacturers both in written form and direct discussions to assist them in developing the data needed to support a PMA.  The FDA's guidance document provides manufacturers with information in several important areas.

The Chemistry Section describes how to analyze the device, the Mechanical Section recommends bench testing, and the Toxicology Section includes a description of the types of biocompatibility information.  All this preclinical testing is designed for evaluating a long-term implant.

The Clinical Section details information regarding study design, data presentation, and information on presentation of literature review.

So much for a brief summary of where we are with silicone gel-filled breast implants.  Now for the task at hand.

Inamed has submitted a PMA to the FDA for its silicone gel-filled breast implant, and that PMA will be the focus of our discussion over the next few days.  This is the first PMA submitted for a silicone gel-filled breast implant since 1992.

We are asking this panel to evaluate the information in Inamed's PMA and advise us as to whether there are sufficient information to provide a reasonable assurance of safety and effectiveness.

Each panel member has been provided a list of questions to consider as he or she reviews the data.  We have heard this morning a number of comments from the public and there will be an additional open public comment session tomorrow.

FDA appreciates the effort our stakeholders have gone to, to provide us with input, and I want to assure everyone here that we will take these comments seriously as whether or not to approve this product.

What follows after the break will be presentations from Inamed and FDA on this application.  Following that, we will solicit the advice of the panel on what action we should take with respect to the application before us.

The FDA very much appreciates your giving of your time and expertise to accomplish this important task.

I will now turn it back over to you, Dr. Whalen.

DR. WHALEN:  Thank you, Dr. Witten.

Are there any panel member questions for Dr. Witten in regard to her presentation?

DR. ANDERSON:  I have a question.  Does FDA offer guidance as to what defines safety and efficacy?

DR. WITTEN:  There is a definition of safety and effectiveness which is in our law, which we will read during the voting process.  It is something--I can't tell you the exact wording--but it is something like it is safe if there is a reasonable assurance that the benefits outweigh the risks in the target population for its intended use.

DR. WHALEN:  It is footnoted in your booklet.  I can't tell you the exact page right now, but each of those definitions are in two continuous pages in the information that was provided us.

DR. WITTEN:  Actually, it is cited in the panel question, because the panel questions that you have, have the definition in there, because we wanted you to keep it in mind as you answered our questions about safety.

DR. WHALEN:  No other questions?  Seeing none, we will take a 10-minute break.  I would ask the sponsor to be ready to start at 10 minutes to 3:00 their presentation.

[Break.]

Applicant Presentation

Inamed Corporation

McGhan Silicone-Filled Breast Implants

Introduction

DR. EHMSEN:  Members of the Advisory Panel, FDA staff, interested guests, I am Ron Ehmsen, Senior Vice President in charge of Clinical and Regulatory Affairs for Inamed Corporation.  As such, I do have a financial interest in the company.

My colleagues and I are very happy to be here today to have the opportunity to present a summary of the clinical and preclinical data that we believe supports our premarket approval application for McGhan silicone gel-filled breast implants.

[Slide.]

We are also joined by a number of consultants and experts in their various fields:  Professor Donald Berry from the University of Texas and M.D. Anderson Cancer Center; Dr. James Brenner, who is our central MRI reviewer, he is a radiologist in private practice in Santa Monica, California; Professor Michael Brook, Professor of Chemistry and McMaster University; Dr. James Lamb, who is a toxicologist with the firm of BDL Associates in Virginia; Joseph McLaughlin, Dr. McLaughlin is President of the International Epidemiology Institute here in Maryland; Professor Kathleen Rodgers from the University of Southern California; Professor David Sarwer from the University of Pennsylvania; Dr. Howard Smith from Cleveland Clinic; and Dr. Scott Spear, who is Professor and Chairman of the Division of Plastic Surgery at Georgetown University Medical Center in Washington, D.C.  Dr. Spear also serves as our principal medical consultant on the clinical studies that will be reported here today.

For clarification purposes, Inamed Aesthetics, which is a business entity of Inamed Corporation, was formerly known as McGhan Medical Corporation, hence, the reference to the trademark McGhan in reference to our breast implants.

Our PMA was submitted in December of 2002 for breast augmentation in women over 18 years of age, for breast reconstruction, and for breast implant revision, which includes removal and replacement of previous breast implants.

[Slide.]

This PMA covers seven device styles as shown in this slide, four smooth round, two textured round, and one textured shaped implant.  The photographs on the righthand side of the slide show the smooth round, the upper photo, and the textured shaped in the lower photo.

[Slide.]

Today, we will discuss the preclinical and clinical findings that were submitted in support of our PMA, as well as review the peer-reviewed research that we believe adds a great deal of credibility.

[Slide.]

First, our key findings.

[Slide.]

Our data demonstrates that McGhan silicone-filled breast implants are safe and effective.  Eighty-nine percent of the patients in the revision cohort, 93 percent in the augmentation cohort, and 96 percent of the patients are satisfied at two years follow-up.

Lower rates were found for most complications in Inamed's Silicone-Filled Breast Implants compared to Inamed's saline-filled breast implants which were approved by FDA in 2000.

[Slide.]

The peer-reviewed research demonstrates that there is no increased risk of breast cancer, no increased risk of rheumatic or autoimmune disease, and no negative effect on children of women with breast implants.

Tom Powell, our Director of Technologies, will now discuss the results of our preclinical studies.  He will be followed by JoAnn Kuhne, our Senior Director of Regulatory and Clinical Affairs, who will discuss the results of our clinical studies.

Let's begin with the preclinical.  Mr. Powell.

Preclinical Studies Review

MR. POWELL:  Thank you, Dr. Ehmsen.  Good afternoon.  I am Tom Powell, Director Technology at Inamed, and I have a financial interest in the company.

I am presenting the preclinical data from evaluations performed by Inamed.

[Slide.]

The key preclinical findings are:  Silicone materials employed in Inamed's gel-filled breast implant demonstrate biocompatibility and support the Institute of Medicine's report where the committee concludes that a review of the toxicology studies of silicones and other substances known to be in breast implants does not provide a basis for health concerns, and that silicone materials meet the physical property specifications for breast implants.

[Slide.]

Inamed's preclinical studies included toxicological evaluations, chemical evaluations, and the mechanical evaluations, meeting the requirements contained in the guidance for saline, silicone gel, and alternative breast implants, guidance for industry and FDA.

For the purposes of this presentation, the findings are summarized.  Complete documented results have been submitted to the agency.

[Slide.]

Toxicity evaluations demonstrated material biocompatibility as described in the FDA's guidance document for breast implants included pharmacokinetics studies, in vivo and in vitro standardized toxicity testing, and studies investigating special toxicological concerns of immunotoxicity, reproductive, and developmental toxicity, mutagenicity, chronic toxicity, and carcinogenicity.

[Slide.]

Inamed's pharmacokinetics evaluation was a scientific study tracking the movement of silicone gel implanted without an elastomer shell subcutaneously in rats. As shown in this gel study, the silicone gel is encapsulated and remains localized at the site of implantation with minimal distant movement.

This result agrees with the IOM report which had similar findings where studies showed that deposit of gel, whether free or in implants, remained almost entirely where injected or implanted.

Additionally, the submitted review of silicone distribution studies from literature further supported the IOM conclusion that, "There do not appear to be long-term systemic toxic effects from silicone gel implants or from unsuspected compounds in the gels or elastomers detected by animal experiments.

[Slide.]

A battery of standard toxicological testing was performed on silicone elastomers and silicone gel.  This battery investigated a variety of biocompatibility parameters including cytotoxicity where the silicones of breast implants were found non-cytotoxic, hemocompatibility demonstrating that silicones are non-hemolytic, pyrogenicity showing the silicones are non-pyrogenic, acute systemic toxicity demonstrating the implant materials are non-toxic, intracutaneous irritation finding that silicones are non-irritant, intramuscular implantation where silicones of breast implants are found non-toxic, and dermal sensitization where the silicones were found to be non-sensitizing.  All testing demonstrated the biocompatibilities of these breast implants silicone materials.

[Slide.]

Additional toxicological studies were undertaken to evaluate areas described as special considerations in the breast implants guidance document.

The immunotoxicity studies evaluating the test animals' immune response to silicone elastomers and silicone gel were performed using National Toxicological Program methodologies.

These assays investigated various immunologic parameters including antibody response to antigen, T cell, T cell subset, and B cell enumeration, mixed leukocyte response, and natural killer cell assay.

In all testing, no adverse immunologic effects of silicone materials was found.

[Slide.]

The reproductive and developmental toxicity assessment included fertility/teratology and reproductive toxicity studies.  The results from these studies sponsored by Inamed demonstrate that silicone materials used in breast implants are neither reproductive nor developmental toxins as there was no effect on parental, fetal, or offspring parameters.

Additionally, the submitted review of silicone reproductive and developmental toxicity studies from literature further supported these results.

[Slide.]

A series of standardized mutagenicity assays designed to investigate the potential for genetic damage was performed on silicone elastomers and silicone gel.  The studies used both bacterial and mammalian cell cultures and included Ames assay, forward mutation assay, mouse lymphoma assay, unscheduled DNA synthesis assay, cell transformation assay, and in vitro cytogenic/chromosomal aberration assay.

In all testing, no evidence of mutagenic activity was found.

[Slide.]

Subchronic toxicity studies demonstrated that breast implant silicones again were non-toxic.  Studies investigating potential for long-term organ or tissue damage and possible cancer development from exposure to silicone elastomers and silicone gel of breast implants were performed using the two-year chronic carcinogenicity tests in rats.

The carcinogenicity testing demonstrated that incidents of carcinomas were similar to controls except for the expected occurrence of solid-state tumors.  Rodent models for this testing are particularly prone to develop solid-state tumors at the site of implantation of foreign bodies.  This well-known rodent response to implanted materials is not a physiologic reaction observed in humans.

In this two-year chronic testing, the recognized gold standard to evaluate long-term toxic effects, no evidence of systemic toxicity or cancer other than rodent solid-state was observed in association with silicone materials.

[Slide.]

Overall, the extent of amount of toxicity testing performed has demonstrated that the silicone elastomers and silicone gel used in Inamed's breast implants are biocompatible and do not compromise the body's physiologic system.

[Slide.]

Looking at preclinical data, the preclinical chemical data characterizing silicone materials in breast implants support the findings of biocompatibility.  The chemical evaluations focused on the silicone materials extent of crosslinking, extractable compounds, and heavy metals content characterization.

[Slide.]

The physical strength and elasticity of the silicone implant shell and patch materials are a result of the extent of crosslinking achieved during the vulcanization process.  Physical property testing of elongation and tensile strength demonstrates the consistency across elastomer lots and that materials conformed to pre-established specifications

Validation testing ensured that the process that is used by Inamed to produce implant shells are adequate to achieve a crosslink density that assures for every lot the strength and elasticity to meet the American Society of Testing and Materials Standard F703, Specification for Breast Implants.

Every lot and every batch of silicone gel is tested to ensure the crosslink density conforms to predetermined specification using penetrometer testing prior to being released for use in the manufacturing of breast implants.

MS. GILBERT:  Excuse me.  Can I ask a question here real quick?

MR. POWELL:  Certainly.

MS. GILBERT:  How do you measure the strength of an implant?  Can you just give us an example of how you tell, I mean do you run a car over it, how do you measure it?

MR. POWELL:  There are number of physical tests that are run on the materials.  Some are standard physical property testing of elongation and tensile strength run on a testing piece of equipment known as a tensile tester Enstron.  They have various brand names.

We actually will talk about, in the Mechanical Section, additional testing that is run on the devices themselves.

MS. GILBERT:  Thank you.

[Slide.]

MR. POWELL:  Finished, sterilized devices were analyzed using techniques to detect extractable components including solvent extraction followed by gas chromatography using a mass spectrometer, ionization detector, and analysis by gel permeation chromatography.

Levels of extractable silicone materials were determined and results of the toxicological testing of the implant materials demonstrated their biocompatibility and non-toxicity and support the IOM report where the committee concluded that a review of the toxicology studies of silicones known to be used in breast implants does not provide a basis for concern at expected exposures.

[Slide.]

Heavy metal analyses were performed with separate chemical analysis on the shell, patch, and gel.  Metal concentrations were found comparable to approved saline-filled breast implants.  Specific analysis with respect to catalyst platinum and tin were performed.

Material toxicity testing demonstrated the biocompatibility  of gel-filled breast implant materials and supported the IOM Committee conclusion that evidence is lacking for an association between platinum in silicone breast implants and local or systemic health effects in women who have these implants.

[Slide.]

Addressing the preclinical mechanical data, it demonstrates that the materials meet physical property specifications for breast implants.

The mechanical preclinical evaluations included fatigue testing, gel bleed testing, and gel cohesiveness evaluations.

[Slide.]

Fatigue testing assesses the number of cycles at specific applied loads that a device can endure without the rupture of the shell.  Smaller sizes of smooth and textured devices, representative of Inamed's gel-filled implants were compressively test.

The testing was performed on individual worse case devices, those with a minimum acceptable thickness using load control equipment.  For this testing, implants were tested at applied load levels of 20, 30, 40, and 55 pounds until either device failure was observed or 6 1/2 million cycles was reached.

In addition to cyclic testing, a test assessing the force to rupture from a single compression of an implant was performed.  The minimum average force for this single compressive failure was 1,245 pounds force.  This force is significantly greater than what would be normally experienced in vivo, such as the maximum 55 pounds force exerted by mammography equipment.

Fatigue values from the applied load test cycles and the single compressive tests were used to generate applied force per number of cycle graphs and a threshold force, the force below which an implant can undergo an unlimited number of cycles without rupture failure was determined.

The minimum threshold value was determined to be 30 pounds force which calculates to a safety factor of 8.1 compared to the expected maximum in vivo cyclic load of 3.7 pounds force.  This means that implants can withstand more than 8 times the maximum expected cyclic load it would experience implanted in the body.

PROF. DUBLER:  Excuse me, I have a question.

MR. POWELL:  Yes.

PROF. DUBLER:  If that is the case--and perhaps you will deal with this later on in your presentation--and if the rupture rate is as reported, then, does the company have theories about what would undergird the rupture rate, if it's not a matter of force?  You seem to be comfortable in all of the data that you are producing about the strength of the elastomer.

MR. POWELL:  The fatigue test is an in vitro test.  There is a number of factors that are involved in the specific design of the fatigue test that may not be really predictive of an in vivo model.

PROF. DUBLER:  And is that true for other of the tests that you are now presenting the findings from?

MR. POWELL:  Yes, the gel bleed testing would be another in vitro assessment that really has no direct clinical correlation.

[Slide.]

Talking about the gel bleed testing, gel bleed testing assess the diffusion rates of silicone gel components through the implant shell and is a PMA requirement.

The study design was consistent with ASTM F703 method and involved the use of a silicone disk technique and results have no direct clinical correlation.

The calculated normalized diffusion rate value did not exceed 0.0019 g/cm2/week.

DR. WHALEN:  Mr. Powell, this is Dr. Whalen with a question on that.  When you assert that there is no clinical correlation to that, how would you then counter someone looking at that--and I am engineering challenged, I will admit--who wanted a sound byte to say, well, we have shown that it leaks, but it leaks slowly?

To me, that would have enormous clinical import, but you say that doesn't have clinical implications?

MR. POWELL:  The silicone materials are really not readily soluble in an aqueous solution.  This test methodology represents the worst case test method that can be assessed with some type of reliability in terms of a result.

DR. WHALEN:  Well, to follow up, I would assume that that standard is incorporated into the requirements for a device like this because of in-vitro concerns about leakage.  If not, why else would it be done?

MR. POWELL:  This was a test method involved in the ASTM Specification F703 to get an assessment technique, so that different designs of gel implants could be assessed for different rates of gel bleed diffusion, more to be used as a comparative for different designs.

[Slide.]

Gel cohesiveness testing assesses the cohesion characteristics of silicone gel.  Finished device gel testing was performed as per ASTM F703 methodology for gel cohesion.

Cohesion result values were well below the accepted 4.5 cm maximum value identified in the ASTM standard for gel-filled breast implants.  The results ranged from zero to 1.1 cm.

DR. CONANT:  A quick question.  In terms of the cohesive testing, does silicone tenacity change at different temperatures?  I assume these are done at room temperature versus body temperature?

MR. POWELL:  The testing is done at room temperature, that is correct.

DR. CONANT:  But how does the--

MR. POWELL:  The crosslinking of the silicone is very thermally stable, so once it is crosslinked, it is not going to change with temperature.

DR. MANNO:  Excuse me, I have a question also.

This gel bleed that you mentioned, was that the method that was presented in the material using the filter?

MR. POWELL:  The earlier drafts of the gel bleed was using filter paper and now we use silicone disk as the absorbent media.

DR. MANNO:  You are using a silicone disk to check for silicone gel leakage or bleed?

MR. POWELL:  We used an extracted silicone elastomer disk that is more absorbent to the silicone materials coming through the silicone implant shell than filter paper would be, so it's the worst case technology.

DR. MANNO:  How are you getting from to get to your amount of bleed from there?

MR. POWELL:  Over an eight-week course of study, weekly measurements are taken of the weight gain of the silicone disk that the implants are placed on and then replaced in exactly the same location.

DR. MANNO:  Okay.

DR. MILLER:  Dr. Whalen, could I ask a question?  I think that you said something which caught a lot of our attention, that I just think we need to clarify.

You said that the tests you are using have no clinical correlation or something like that, and I think that from the engineering point of view, what you are trying to say is that there are no comparable events in the body that are similar to the contrived test you are doing in a materials engineering sort of way, but the results you are getting have clinical implications because they are telling you something about the nature of the materials in the implants.  Is that accurate?

MR. POWELL:  I would say that would be accurate if you wanted to compare between two devices, if you had similar experimental data between two different designs, then, you could make somewhat of an assessment in terms of a clinical outcome, but until you have a correlation between both the physical mechanical testing and the clinical outcome, it may not be directly related.

DR. MILLER:  There is a limitation--I mean I just want to make sure I understand you--when you do the tests, you get results.  The results tell you something about the materials.  How you move from those results to what is going to happen in a clinical situation is subject to some limitations.

MR. POWELL:  It would be correlated with the clinical studies, yes, sir.

PROF. DUBLER:  Dr. Whalen, I don't understand.  I don't understand what you would like us to do with the data that you have just presented in terms of the sorts of deliberations that we are required to have.

Can you tell us how you would like us to view the data that you have just presented?

MR. POWELL:  Let me refer to my colleague, Dr. Brook, maybe could better answer that question.

DR. BROOK:  My name is Michael Brook.  I am a Professor of Chemistry at McMaster University in Hamilton, Canada, and I am a paid consultant to Inamed.

I can understand the sense of confusion on the part of the committee and perhaps I can try and clarify the situation.

When one is trying to do a materials test, you want reliable and reproducible tests that allows you to understand just how easily material is going to bleed, is going to leak out through this implant.

We all know silicones are typically used for waterproofing materials.  If you want to put shoe polish on that contains silicones, so that the water beads up, and if one was to try and design perhaps more clinically relevant tests, for example, into water, the problem is that the measurements are almost impossible to do.  You don't get reliable data because so little goes into water or saline or other artificial materials that you can't get reliable data.

So, the ASTM test is designed kind of like Kleenex in water.  We all know that the water wicks up into the Kleenex.  If you choose silicone, which is the material in which silicones will dissolve, you kind of get the worst case scenario.  You learn how bad it could possible be under the worst possible case.

You are essentially allowing it to go into a sponge that has exactly the same kind of properties.  So, this allows you to develop a test in which you have reliable data, measurable data, quantifiable data, so you have good certainty, good numbers, and it counts as a worst case scenario.  The problem is that you can't do that same kind of test into water, which is what you might like to do in order to make the clinical correlation.  It is just so little that goes in, you can't use it as a reliable test.

DR. CHOTI:  Just a question.  Do these preclinical tests, any differences between the smooth and textured, or the double lumen and single lumen?

MR. POWELL:  All these tests were done on both smooth and textured devices.

DR. CHOTI:  Are there differences in bleed in mechanical properties?

MR. POWELL:  Yes, there are differences in what I tried to describe here are the worst case scenarios.

DR. CHOTI:  If so, one useful correlate would be whether there are differences when you talk about the clinical data between leak, rupture, and so forth, between smooth and textured, for example.

Does the textured have less of the stronger mechanical properties, for example, than the smooth?

MR. POWELL:  Actually, it has weaker, some of the break force values, however, in the fatigue test it was stronger, so it is not strictly a direct relationship with the testing that has been done.

DR. CHOTI:  And is the patch a weaker area or a stronger area?

MR. POWELL:  The patch is definitely the strongest area on the implant, and it is the same for both.

DR. BOULWARE:  Dr. Whalen, I would like to ask a question.

Am I interpreting your gel bleed data correctly, 0.0012 gram/square centimeter/week, so if I extrapolate that out for 52 weeks, about 0.1 gram/square centimeter, and assuming that it's about 6 inches across, two-sided, I just did it quickly here, about 150 centimeters, that is 15 grams per year bleed?

MR. POWELL:  In the smallest case, the calculated value in a year would be 14 grams of gel.  We have never had any implant come back from a returned device analysis that has ever had that much.

DR. BOULWARE:  One more question.  The permeability to water for the envelope, is it permeable at all to water?

MR. POWELL:  It is permeable to water vapor, but permeable to water, I don't believe that is measurable.  Dr. Brook?

DR. BROOK:  No.

DR. MANNO:  On these numbers you are giving as normalized diffusion rates, I have noticed that you have given a single number, but should you not have associated something like a standard deviation in there someplace or a standard error, or a range of responses?  This is normally how it is done in the biological community rather than the straight chemical community.

Secondly, I would go along with the earlier question concerning having tested this, it's fine to test it at 25, but if you really want to see how it is going to work and you don't want to have to stand there and make a reference, that there is no clinical relevance here with this or correlation, then, it should have been done, I would think, at 37 degrees, which would be body temperature.

MR. POWELL:  Thank you.  Again, this is following the ASTM methodology, which is fairly precise in the way they want it calculated.

DR. MANNO:  I would suggest that when you are looking at guidelines in developing a device and you are looking at safety, when guidelines should be considered to be a minimum requirement for testing, and anything else, while it may be costly and you may not think much of it, it is certainly icing on the cake to give us more to work with safetywise.

MS. BROWN:  That leads to a question that I had. Were you following or what was your understanding of the guidance that you were supposed to be using for the testing, is it the breast implant guidance?

MR. POWELL:  Yes, it was the FDA guidance that we were following.

MS. BROWN:  Okay.

DR. OLDING:  Michael Olding.  I just want to back up for a second.  You said this is the worst case scenario, and backing up to the mechanical evaluations, one would be surprised, given that information, that there would ever be a ruptured implant, but those values are really on just implanted implants, if you can relate them clinically, is that correct?

They were obviously not in vivo, but they could have been removed after they had been in at a year or two and tested, so that it would have been more clinically relevant?

MR. POWELL:  Yes.  We have accelerated shelf life studies that compare implants right after sterilization to those that would be on the shelf for five years, and find that the physical properties are not at all compromised in those particular comparisons.

Then, in the initial retrieval analysis, looking at in vivo testing that we are doing, we find also that at this point, with the sampling that we have, that the physical shell strength requirements also have not diminished with the samples that we have tested in the retrieval program.

DR. OLDING:  How far out is that?

MR. POWELL:  The retrieval program has tested samples up to 11 years, but it right now is a small distribution.

DR. OLDING:  Thank you.

DR. LEITCH:  Could I use more clarification about the bleed?  It sounds like on your retrieval studies, you have verified that the bleed rate has not ever been to this extent, the 14 grams per year?

MR. POWELL:  Correct.

DR. LEITCH:  So, what is it?

MR. POWELL:  What we evaluate when we have a returned product that comes back, we will look at the weight of the implant, we will weigh it and we will see if it's within the specifications for that implant, which is within 5 grams of the specified volume.

To my knowledge, we have never had an implant--

DR. LEITCH:  So, you have never had one that has been more than 5 grams difference?

MR. POWELL:  No, ma'am.

DR. LEITCH:  Like how many would that be that you have looked at?

MR. POWELL:  In the report that we submitted, we had 339 returned implants back.

DR. CONANT:  Another quick question about gel bleed or the reverse, the solubility or the entry of things into the implant, how do you explain, with aging implants, the change of the internal contents, the opacifications, sometimes on imaging we may see water droplets or presumably serum within--suggesting that things are also getting in as things are getting out?

Have you noticed that in your retrieval program, they look different?  I think we saw images this morning also of it.

MR. POWELL:  There is a tendency, I believe--and, Dr. Brook, maybe you can help me with the potential for colorization of the implants?

DR. BROOK:  When one exposes any material to a complicated environment, everything is going to be soluble to different degrees, so, for example, alcohol in water are infinitely soluble in water, whereas, sand in water, sand doesn't dissolve very well in the water.

So, those are the kinds of things that we also see with implants.  Essentially, this outer layer is a membrane, and it's acting to prohibit some transport at some level, so that the rule of thumb is that things which are similar chemically, typically dissolve in one another.

So, for example, lipids might be anticipated to dissolve to a better degree in the silicone than would water, for example.  So, I don't know the answer specifically with Inamed, is it like that some lipids cross that membrane that go into the gel, yes, the answer is that there is going to be a higher solubility of lipids.

Similarly, I should come back to this question of water.  Absolutes are very dangerous.  Small amounts of water are likely to go into the inside of the implant.  The converse, the solubility of silicones in water is about 12 parts per million.  It will be sort of comparable the other way around - small amounts can transform either in the gaseous form or in the liquid form.

So, this is a membrane that is trying to keep the two parts of the body and the gel apart, but it doesn't work perfectly.  It is extremely difficult to design a barrier that is a true barrier for everything, so there will be some transport in both directions.

DR. MILLER:  Could I ask one question, too?  Do you have any theoretical calculation that could tell us, in the in vivo environment, what the bleed rate would be based upon the numbers that you get in your tests?

MR. POWELL:  No, I do not.

DR. MILLER:  Is it something that is possible to do, to make a theoretical calculation like that, or is that something that you just can't do?

DR. BROOK:  I don't think so.  I mean first of all, this is my personal opinion, essentially, everybody is slightly different, what is the exact local environment, could you come up with a dummy solution that would be a very good mimic.

DR. MILLER:  Well, a physiologic solution of some kind, just an idealized solution, could you just put the numbers in a chemical, you know, some kind of equation for analysis and calculate what the bleed rate would be based upon these worst case scenario bleed rates?

DR. BROOK:  My opinion is that it would be very difficult to try and do that.  Among other things, the environment around the implant is changing with time as you start to build up a capsule, and so at which point do you want to try and mimic the local environment.

I guess by doing these tests into silicone itself, you know what the upper limit is, if you did it into saline, you probably would be able to measure, if you could, what the lower limit is, and it will be somewhere in the middle in vivo.  But I don't know of any way to calculate that.

I guess since we are out 40 years and people have been worried about this for a long period of time, I am not the only one who has been unable to come up with a reliable test to get at that number in vivo.

[Slide.]

MR. POWELL:  So, just in summary, the preclinical evaluations performed in accordance to the FDA's guidance document demonstrates that the silicone gel-filled implant materials meet the physical property specifications for breast implants and the biocompatibility of silicone breast implants.

I will now turn the presentation over to Ms. JoAnn Kuhne for the clinical discussion.

DR. WHALEN:  During that transition, let me just say to the panel members without any regard to where we stand on our schedule, I would like to strongly encourage the greatest possible scrutiny over both this presentation and FDA's.  I don't want to restrict it in the least.

However, I think it might make sense from a flow point of view that we let Ms. Kuhne and then each of the FDA presenters, or whoever is representing Ms. Kuhne, do their entire presentation, make notations of what questions you want to ask, and at the conclusion, at their own respective conclusions, we will take breaks and then ask some questions.

Thanks.

Clinical Studies

MS. KUHNE:  Good afternoon.  I am JoAnn Kuhne, Senior Director of Regulatory and Clinical Affairs, and I do have a financial interest in the company.

[Slide.]

First, I would like to address the key clinical findings from our ongoing clinical trials.  From my presentation, you will note that Inamed's clinical data demonstrate safety and effectiveness of McGhan silicone gel- filled breast implants.

General patient satisfaction ranged from 89 percent in the revision cohort to 96 percent in the augmentation cohort with revision patient satisfaction falling within this range at 93 percent.

There were lower rates for most complications for Inamed's silicone-filled breast implants as compared to Inamed's FDA-approved saline-filled breast implants.

[Slide.]

Now, let me briefly introduce what I will be discussing for the next 40 minutes.  My remarks will encompass the scope of our clinical studies, the safety results obtained, the effectiveness results, comparison of the complication rates between the silicone gel-filled implants that are the subject of this PMA and our saline breast implants that have been FDA-approved for over three years.

In addition, some highlights from published peer-reviewed clinical research, and finally, the conclusions from our prospective clinical trials.

[Slide.]

Now, for a brief overview of the three main studies Inamed has conducted on silicone gel-filled breast implants, which I may also refer to throughout the presentation as gel implants.

All three studies were prospective, multi-center trials, but the most comprehensive clinical trial, and the one that will be the primary focus of my presentation here today, is the core study.  The core study provided the primary data for this PMA submission.

The study allows for evaluation of both safety and effectiveness of the gel implants.  Our results include complete two-year follow-up data, as well as partial three-year data for those patients who have already returned for their three-year visit.

The study includes women in three cohorts or groups with approximately 50 percent in the augmentation group, 25 percent in reconstruction, and 25 percent in the revision group.

The adjunct study is a large ongoing safety study that also includes both reconstruction and select revision patients with a current enrollment of over 25,000 patients.

The AR90 was a five-year study which included a small sample size of both augmentation and reconstruction patients.

[Slide.]

For a quick look at the scope of the three studies, we can see that the core study was initiated in 1999 and enrolled 940 patients.  Patients are currently being seen for their three- and four-year follow-up visits.

The core study protocol is designed to provide for 10 years of clinical follow-up.

The adjunct study began in 1997 and continues to enroll new patients.

The AR90 study began in 1990 and was completed after five years of follow-up.

[Slide.]

Our long-term post-approval studies for our core clinical study patients, should we receive PMA approval, consists of two phases.  The first phase involves continuing to collect patient data via direct physician evaluation through five years post-implantation.

This phase also includes obtaining MRIs through five years on the patients enrolled in the serial MRI portion of the study.

Phase II of our proposed post-approval plan involves continuing to follow up on these patients through 10 years via mailed surveys.  We have many measures in place to assure high compliance in mailed survey studies, such as premailings and follow-up reminder letters, as well as phone calls.

[Slide.]

Now, for more detail regarding the core study, which again is a prospective, multi-center clinical trial involving 940 patients enrolled at over 50 sites nationwide.

Patients enrolled into the core clinical study with three different indications - augmentation, reconstruction, and revision.

Augmentation patients are those patients who seek breast implantation, reconstruction, and revision.  Augmentation patients are those patients who seek breast implantation to enhance the size and/or shape of their breasts.

Reconstruction patients are women in whom one or both breasts are reconstructed after illness or injury.  The majority of reconstruction patients choose to have breast implants placed after a mastectomy to treat breast cancer.

Revision patients are patients who receive breast implants in a previous augmentation or reconstruction surgery and are having them replaced either due to cosmetic or complication reasons.

The following slide show results of the augmentation and reconstruction procedures.

[Slide.]

This slide shows the before and after results of a cosmetic augmentation with McGhan style 45 silicone-filled breast implants.  The patient was unsatisfied with the appearance of her breasts and opted for breast implants to enhance their shape and size.

[Slide.]

This slide shows the staged reconstruction results of a patient with McGhan style 10 silicone-filled breast implants.  The patient previously underwent a bilateral mastectomy and desired to regain her shape.

The picture in the top left shows this patient pre-mastectomy.  Moving directly to the right, we see the patient post-mastectomy after placement of tissue expanders.

The bottom left slide shows the patient after placement of breast implants, and immediately to the right, we see this patient approximately one year postoperatively.

Again, the revision indication refers to those patients who have had previous augmentation or reconstruction breast implant surgery and have their implants removed and replaced either due to choice or complications.

[Slide.]

Now, for a description of the patient population. We see that the median age for augmentation patients is 34 years with those patients in the reconstruction and revision groups being somewhat older.

Most of the women in this study are Caucasian and approximately half or more in each group are married.

[Slide.]

For augmentation patients, over 90 percent of the implants used were round with a smooth surface somewhat more common than textured.  In reconstruction patients, 65 percent of the devices used were the textured shaped implants.  Round implants were more common in revision patients with 71 percent choosing round devices fairly evenly split between smooth and textured surfaces.

[Slide.]

High compliance rates have been maintained throughout the study with 90 percent of augmentation patients returning for their two-year follow-up visit, 95 percent of reconstruction patients, and 87 percent of revision patients.

Compliance rates are similar for our partial three-year data.

[Slide.]

Turning to the safety results.

[Slide.]

The complications we collected include reoperation, implant removal also referred to as explant or explantation with or without implant placement, rupture, capsular contracture, infection, and a number of others which we will discuss in greater detail shortly.

The data are based on physical assessment of the patient through a physical exam.  The method used to analyze these complications at three years was a cumulative risk analysis based on the Kaplan-Meier product limit method.  This is the same method that was used for the analysis and presentation of our data for Inamed's FDA-approved saline-filled breast implants.

[Slide.]

Our protocol was designed to collect data on more than 30 different complications, all of which can be grouped into the following three categories:  general breast surgery complications, breast implant related, cosmetic complications, and breast implant related, non-cosmetic complications.

[Slide.]

General breast surgery complications are defined as complications which are related to breast surgery in general regardless of whether the surgery involves breast implants.  This category includes surgery-related complications, such as breast pain, swelling, infection, and tissue or skin necrosis.

In the general breast surgery category, breast pain and swelling each occurred 7 percent or less in all cohorts.  All other complications occurred at less than 5 percent.  As you can see, infection ranged from 0.2 percent to 2.3 percent in the three cohorts, and tissue, skin necrosis ranged from 0.2 percent in the augmentation group to 4.3 percent in the reconstruction group with the revision rate falling in between.

[Slide.]

Breast implant related cosmetic complications are defined as complications resulting from the breast implant surgery which are strictly related to the cosmetic appearance of the breast.

This category contains the following cosmetic complications:  Asymmetry, implant palpability, visibility, implant malposition, and wrinkling.

In the cosmetic category, asymmetry is 2.8 percent for augmentation, 15 percent for reconstruction, and 5 percent for revision.  Asymmetry could be expected to be higher in the reconstruction group because 37 percent of this group underwent only unilateral implantation.

All other complications are 5 percent or less in all groups with wrinkling ranging from 0.7 percent to 5 percent, and implant malposition ranging from 3 percent to 5 percent.  Implant palpability is 2 percent or less in all groups.

[Slide.]

Breast implant related non-cosmetic complications are defined as complications occurring as a result of the breast implant surgery which are not considered strictly cosmetic in nature.  This category includes the following non-cosmetic complications:  capsular contracture, implant rupture, capsule calcification, and implant extrusion.

For the non-cosmetic complications, capsular contracture is 8 percent for augmentation, 16 percent for reconstruction, and 9.8 percent for revision.  Rupture is 1 percent for augmentation, 6 percent for reconstruction, and 3.6 percent for revision.

Both capsule calcification and implant extrusion are less than 1 percent in all cohorts.

[Slide.]

Regarding other complications in the core clinical study, silent rupture, meaning patients who may have experienced a ruptured implant, but have no physical symptoms of a rupture, was studies via a subgroup of 351 patients who are undergoing postoperative serial magnetic resonance imaging or MRI.

These patients underwent routine MRI every two years.  Of the 597 implants, 317 patients that have undergone at least 1 MRI in the serial MRI portion of the study with 170 implants having undergone 2 MRIs, 15 implants have been classified as having a silent rupture.

The calculated Kaplan-Meier risk of silent rupture is 2.7 percent at three years.  The risk of 2.7 percent includes 2 out of the 15 implants that have been classified as indeterminate for implant rupture meaning it has not been confirmed that these two implants are actually ruptured.

It is important to note that because the patients are not experiencing any symptoms or adverse events related to the implant, some physicians do not recommend removing the implants that may have ruptured silently.

[Slide.]

Now, we will be presenting the reoperation findings from the core clinical study.  It is important to note that reoperation is not synonymous with implant removal or explantation.  Instead, reoperation means any surgery that occurs after implantation.

As you can see on this slide, the three-year risk of reoperation was 20.6 percent in the augmentation group,  45.9 percent in reconstruction, and 33.4 percent for revision.  Many different types of procedures are formed during reoperations as will be seen in the next slide.

[Slide.]

Presented here are some examples of the types of procedures performed during reoperation, such as aspiration, biopsy, capsule procedures, flap or pocket revision, mastopexy, and nipple procedure.

As shown next in this list, removal/replacement of an implant is just one of the many types of procedures performed during a reoperation.  Reposition implant and scar revision, wound repair are some of the last examples of possible procedures performed during a reoperation.

Of all these procedures listed for reoperations, the most common were capsule procedures.  A capsule procedure involves removal, suturing, or position manipulation of the capsule that forms around the implant.

[Slide.]

Within the augmentation group, 19 percent or 94 of the 494 patients underwent a reoperation.  Of the 94 patients who had a reoperation, 32 had an implant removed for a total of 60 explants.

[Slide.]

As seen in the pie chart on the left, of the 60 implants that were removed in the augmentation group, many were removed to treat capsular contracture; 18 were due to patient choice for a different size or style implant, and another 11 implants were removed for cosmetic complications, such as asymmetry or implant malposition.

The pie chart on the right shows that the majority of explants, 39 out of 60, were replaced with another silicone gel-filled implant, and only 9 or 15 percent were not replaced.  Twelve explants were replaced with other breast implants that are not part of this study, for example, with gel-filled implants that are part of another study or with saline-filled breast implants.

Another significant finding was that upon explantation of these devices, there were no findings of extracapsular gel.  These means that during the removal of these 60 breast implants, none of the physicians found that gel had migrated outside the capsule.  This is consistent with feedback we received from plastic surgeons in general who indicate that it is not common to find extracapsular gel when removing an implant.

[Slide.]

Forty-two percent or 92 of the 221 reconstruction patients had a reoperation.  Of the 92 patients who had a reoperation, 46 patients had an implant removed for a total of 56 explants.

[Slide.]

As shown in the pie chart on the left, of the 56 implants that were removed in the reconstruction group, 4 were due to patient choice for a different style or size implant, while 27 were removed to treat an unsatisfactory cosmetic result, such as asymmetry.

The pie chart on the right shows the majority, 45 explants were replaced with another gel implant, and only 9 percent or 5 explants were not replaced.

Furthermore, there were no findings of extracapsular gel during any of the explant procedures.  This means that during the removal of these 56 devices, none of the physicians found that gel had migrated outside the capsule.  This is also consistent with our augmentation group.  You will see in the next slides that our revision cohort did not experience any extracapsular gel either.

[Slide.]

Thirty-one percent or 70 out of the 225 patients in the revision group underwent a reoperation.  Of the 70 patients who had a reoperation, 29 had an implant removed for a total of 46 explants.

[Slide.]

Of the 46 explants that were removed in 29 patients in the revision group, many were due to patient choice for replacement with a different size or style implant or for an unsatisfactory cosmetic result, such as malposition, ptosis, or scarring.

The chart on the right shows that 29 implants or 63 percent were replaced with another gel implant.  Again, there were no findings of extracapsular gel at the time of explantation of the 46 devices.

To reiterate, in this large controlled clinical trial, we see that none of the 162 implants removed showed gel outside the capsule formation upon implant removal.  Plastic surgeons have shared with us that extracapsular gel is not a common event and is usually associated with a traumatic injury, such as a car accident or closed capsulotomy, which is a procedure that is not as common today as it used to be.

However, it is important to note that if the surgeon had observed extracapsular gel, it is expected that the gel would have remained at the site as demonstrated in our preclinical experiments where free gel was injected into experimental rats and stayed at the site of injection.

[Slide.]

Moving on to other evaluations of safety - breast cancer, connective tissue disease, autoimmune disease, lactation, and reproduction.

[Slide.]

In the augmentation group, there was one report of breast cancer post-implantation.  Reconstruction had 5 reports and revision had zero reports.  It is important to note that all the reconstruction patients with post-implant breast cancer also had pre-implant breast cancer with 2 of these patients experiencing recurrent breast cancer in the same breast.

[Slide.]

Of the 494 augmentation patients, there was one diagnosis of rheumatoid arthritis at 11 months postoperatively.  Of the 221 reconstruction patients, there was one diagnosis of systemic sclerosis/scleroderma at four months postoperatively.  Of the 225 revision patients, there was one diagnosis of fibromyalgia at 11 months postoperatively.

These connective tissue diseases occurred in less than one-half of 1 percent of all patients in the study.

[Slide.]

In the augmentation group, there were 5 reports of lactation problems primarily mastitis and inadequate milk production, and 8 reports of reproductive problems primarily spontaneous abortions.  One of these 8 patients had a planned abortion to treat a medical problem pre-implant and experienced a miscarriage post-implant.

Reconstruction had no lactation problems and 2 reported reproductive problems.  One patient experienced a miscarriage pre-implant and no menses post-implant.

Revision had 4 lactation problems, the majority of which were inadequate milk production postoperatively, and 6 reproductive problems mainly spontaneous abortions and infertility.  Two of these 6 patients experienced infertility both pre- and post-implant.

[Slide.]

For the effectiveness results.

[Slide.]

Now, I will review the various components of our effectiveness results.  Besides the endpoint related to increasing the breast mound, we examined three different measures:  patient motivation, patient satisfaction, and psychosocial changes.

Within our satisfaction measures, we captured general patient satisfaction at follow-up and patient satisfaction with the size and shape of the breast pre- and post-implantation.

Within our psychosocial measures, we captured information pre- and post-implantation to assess any change in women's expectations of their breast implant surgery, self-esteem, and quality of life.

[Slide.]

One of the areas we explored was patient motivation for implant surgery.  Therefore, we asked patients to rate their motivations prior to implantation.  As the side shows, the vast majority of women were motivated to improve the way they felt about themselves or make them feel better about their physical appearance.

It is interesting to note that motivations regarding partners and/or sex life were not a strong motivating factor for most of these patients.

[Slide.]

Patient satisfaction was extremely high with one year rates at 96 percent for augmentation, 94 percent for reconstruction, and 87 percent for revision.  Satisfaction was 96 percent, 93 percent, and 89 percent respectively two years after implantation.

[Slide.]

When measuring patient satisfaction with breast size and shape at baseline and at the two-year follow-up, a significant positive change was observed in all three cohorts or groups.

[Slide.]

In terms of patient's psychosocial changes, a questionnaire was developed by Dr. Julia Roland, the director of the National Cancer Institute's Office of Cancer Survivorship.  This questionnaire was used to observe the pre-implant versus two years post-implant changes in general psychosocial concepts.

This graph shows that self-image increased significantly in the augmentation and revision cohorts postoperatively while it remained constant in the reconstruction cohort.

On the improved social relations subscale, there was a positive increase in all three groups.

[Slide.]

For the Tennessee Self-Concept Scale, scores for all patients started high and despite any postoperative complications that may have occurred, remained high.  The scores show no change in the augmentation and reconstruction groups pre- versus two years postoperatively.

We observed a statistically significant decrease at two years post-implantation in the way revision patients scored on this measure.  This measure included aspects of how patients viewed their bodies and state of health, skills, their attitude about appearance, and sexuality.

[Slide.]

Similarly, the scores were high both preoperatively and two years postoperatively for the Rosenberg Self-Esteem Scale.  The scale provides for a maximum score of 40.

The scores obtained revealed a statistically significant decrease two years post-implantation in self-esteem for revision patients, whereas, no significant self-esteem changes were found in the other groups.

[Slide.]

Additional information was collected using the Body Esteem Scale.  This scale measures body image based on three subscales.  These subscales are sexual attractiveness, weight concern, and physical condition.

The bar charts show the changes observed in the total score in the sexual attractiveness subscale.  The bar charts on the left show that the total score does not change significantly in any of the groups, and the bar chart on the right shows significant increase in the sexual attractiveness scale in the augmentation cohort.

[Slide.]

Various SF-36 scales were measured.  In all cases, the quality of life scores remained well above those of the general U.S. female population both at baseline and at two years.  Although scores started high and there were no major changes overall, some of the individual measures decreased postoperatively.

This may be attributed to ceiling effects seen at baseline, for example, 65 percent of the patients in the augmentation cohort scored 100 preoperatively, which is the maximum score indicating best possible quality of life at baseline for the role limitations due to physical health problems measure.

Therefore, this 65 percent of the patients could not have had any increase postoperatively because they were already at maximum score preoperatively.

[Slide.]

In the next few slides, I will discuss the comparison of our current three-year core study complication rates to the three-year complication rates reported for Inamed's FDA-approved saline-filled breast implants which have been commercially marketed since May 2000.

The protocol for the saline study was very similar to the current core gel study with similar endpoints collected and similar follow-up intervals.  In addition, many of the investigational sites in the saline study are also enrolled in the core gel study.

You will note that the complication rates for the silicone gel-filled breast implants that are the subject of this current PMA we are discussing today are comparable and, in many cases, lower than those same complications for our already PMA-approved saline-filled breast implants.

[Slide.]

For the augmentation groups, all complications with risks greater than or equal to 5 percent in either the current gel-filled breast implants or Inamed's previously FDA-approved saline-filled breast implants are listed.

All the complications that showed gel-filled breast implants with a lower complication rate than those found in the saline study are noted in yellow.  Almost two-thirds of the risks shown associated with breast implant surgery are considerably lower for gel-filled breast implants versus saline-filled implants.

For instance, we see that wrinkling in our gel patients occurs at 0.7 percent at 3 years, which is less than one-tenth of the risk of wrinkling in our patients with FDA-approved saline-filled devices.

Furthermore, implant palpability, visibility showed less than 1 percent risk with our gel-filled devices, while the risk of this complication was 9 percent with our saline-filled devices.

It is also important to note that the complication of implant malposition presented in our saline augmentation group where the risk of 8 percent is more than twice as high as the same complication in our gel-filled devices with the risk of only 3 percent.

Gel rupture is one-fifth the risk of saline deflation with gel rupture occurring at a risk of only 1.2 percent in the augmentation group.

[Slide.]

Now, with a similar list of complication rates for the reconstruction group, we see that asymmetry at 15.3 percent with gel-filled implants is half the risk seen with saline-filled implants.

With the saline-filled implant, wrinkling and implant palpability each give a risk in the 20 percent range, while with our gel-filled devices, these two complications are dramatically lower.

Implant palpability in the gel-filled implants is only 0.5 percent.  The complications of loss of nipple sensation and skin paresthesia further the evidence supporting the gel-filled breast implant safety.

In the saline-filled implant study, we saw that these risks were 12 percent and 5.6 percent respectively, but the rates for these complications in the core gel study were zero.

[Slide.]

For the revision group, we see three complications that are displayed in yellow, indicating significant differences.  All three are significantly lower with our gel-filled implants versus our saline-filled FDA-approved implants.  Again, implant palpability shows a dramatically lower risk in the revision group.  This complication was reduced significantly across all groups - augmentation, reconstruction, and revision.

Asymmetry and wrinkling are also significantly decreased with the use of gel-filled breast implants versus saline-filled breast implants.  While the risk with saline-filled implants are over 15 percent for each of these complications, gel-filled implants provide one-third the risk at 5 percent each.

It is of interest that asymmetry and wrinkling were also significantly reduced in the augmentation and reconstruction groups showing that the gel-filled implants consistently reduced the concerns regarding asymmetry and wrinkling.

To summarize the comparison of complication rates of silicone gel-filled breast implants to Inamed's FDA-approved saline-filled breast implants, the complication safety profile for silicone gel-filled breast implants meets or exceeds the complication safety profile for saline-filled  breast implants.

DR. ANDERSON:  May I ask a very quick clarification question?  I didn't hear you say what the follow-up time for the saline group was.  We know for the gel, it is two years, but are these really comparable groups?

MS. KUHNE:  These are three years to three years. This is our three-year gel data, as well as our three-year saline data.

[Slide.]

Comparing our core study results to results reported in the medical literature, we find that with the exception of reoperation, all of the core study results fall at the low end of the rates found in other published studies with various lengths of follow-up.

[Slide.]

To summarize, the core study, conforming to FDA guidelines, demonstrates safety and efficacy of McGhan silicone gel-filled breast implants.  We see that 89 to 96 percent of all patients were satisfied at two years follow-up.  Lastly, we see lower rates for most complications for Inamed's gel-filled breast implants compared to Inamed's FDA-approved saline-filled breast implants.

[Slide.]

Examining the findings of peer-reviewed research, we have chosen to present findings from many independent reviews of published research by prestigious organizations rather than selectively citing individual published articles as you will see on the following slides.

Although these studies are not specific to Inamed's silicone gel-filled breast implants, they are representative of the body of knowledge that currently exists on silicone gel-filled breast implants relative to systemic diseases.

[Slide.]

As a result of concerns raised in the early 1990s regarding the safety of silicone gel-filled breast implants, many government-funded agencies critically reviewed the available scientific literature concerning implants and disease.

These agencies have been consistent in finding no evidence that silicone gel-filled breast implants are a significant health risk.  Without exception, they have found that silicone gel-filled breast implants do not increase the risk of breast cancer or autoimmune disease, nor do they affect the health of children born to mothers with breast implants.

[Slide.]

With regard to cancer, many independent reviews of the literature have been performed to examine the relationship of breast implants to cancer.  Specifically, an assessment of the toxic potential of silicone implants by the International Agency on Cancer found that there is no evidence that silicone implants cause breast carcinoma, and there is inadequate evidence that silicone implants cause other types of cancer.

In its review of the literature current to 1999, the Institute of Medicine concluded that there was no increase in primary or recurrent breast cancer in implanted women.

[Slide.]

Similarly, the National Institutes of Health 2003 stated that the most recent studies, which have been considerably larger than previous studies, and have been able to assess long-term risks, do not support the notion that implants are related to breast cancer risk.

The European Committee on Quality Assurance and Medical Devices in Plastic Surgery 2000 Consensus Declaration on Implants stated that medical studies have not demonstrated any association between silicone gel-filled breast implants and cancer or any other type of malignant disease.

[Slide.]

Still more independent agencies have presented information regarding cancer.  The American Medical Association Report 8 of the Council of Scientific Affairs showed no evidence of increased risk of cancer among recipient of silicone breast implant.

Like the United States, the Health Council of the Netherlands' review found there were no indications of an increased risk of cancer in women with breast implants, nor of a delay in the detection of breast cancer.

[Slide.]

With regard to rheumatic or autoimmune disease, many independent reviews of the literature have been performed to examine the relationship of breast implants to rheumatic or autoimmune disease.

In the United States, the Institute of Medicine found no evidence that silicone implants contribute to an increase in autoimmune disease.

The American Medical Association stated that no studies have found evidence of a relationship between silicone gel implants and connective tissue disease, and the Federal Judiciary National Science Panel found no association between breast implants and any of the individual connective tissue diseases combined or other autoimmune or rheumatic conditions.

[Slide.]

A meta-analysis performed by Janowsky et al., published in the New England Journal of Medicine in the year 2000 found no conclusive evidence to suggest an abnormal immune response to silicone from breast implants, nor an association with established connective tissue disease.

But most importantly, the American College of Rheumatology has weighed in on the controversy and found compelling evidence that silicone implants do not expose patients to additional risk for connective tissue or rheumatic disease.

[Slide.]

Other national agencies outside the United States that have investigated the relationship between silicone breast implants and autoimmune disease have come to similar conclusions.  For example, the Health Council of the Netherlands found that there is unlikely to be a connection between silicone breast implants and autoimmune disease.

A United Kingdom independent review group found no conclusive evidence for an abnormal immune response to silicone from breast implants and no epidemiologic link between silicone gel implants and any established connective tissue disease.

[Slide.]

The European Committee on Quality Assurance and Medical Devices in the Plastic Surgery 2000 Consensus Declaration on Implants concluded that medical studies have not demonstrated any association between silicone gel-filled breast implants and autoimmune or connective tissue diseases.

Finally, Australia's Therapeutic Devices Evaluation Committee in 1997 stated that several large studies have failed to establish a link between silicone breast implants and well-defined connective tissue diseases including scleroderma.

[Slide.]

The potential for silicone gel-filled breast implants to impact breast feeding or affect child health has also been considered by a number of government-funded agencies.

The Institute of Medicine Report found that the evidence for an association of maternal silicone breast implants and children's health effects is insufficient or flawed and that there was no data to support the suspected transmission of silicone to infants in breast milk of mothers with implants.

In the United Kingdom, an independent Review Group on Silicone Breast Implants found that published literature to date does not substantiate the claims that there are significant clinically apparent second-generation effects in children born to silicone breast implant patients.

[Slide.]

Among the medical associations that have considered the issue, the American Academy of Pediatrics stated that evidence currently does not justify classifying silicone implants as a contraindication to breastfeeding, and the American College of Obstetricians and Gynecologists feels that women who have had breast implants may still be able to breastfeed.

[Slide.]

As shown in the previous slides, many agencies in the scientific community have reviewed information regarding silicone gel-filled breast implants.  These respective agencies include the Institute of Medicine, the National Institutes of Health, the American Medical Association, and other associations within the United States.

In addition to those prestigious authorities, other agencies outside the United States have presented reviews including the European Committee on Quality Assurance, International Agency on Cancer Research, and the Health Council of the Netherlands.

Each of their reviews supports our key finding that there is no causal relationship between silicone gel-filled breast implants and breast cancer, children's health, or rheumatic or autoimmune disease.

[Slide.]

In summary, the vast body of published literature, as well as numerous scientific agencies, found no increased risk of breast cancer, no increased risk of rheumatic or autoimmune disease, and no negative effects of breast implants and breastfeeding on children's health.

[Slide.]

This concludes the clinical portion of our presentation.  In summary, we believe that Inamed's preclinical and clinical data support the safety and efficacy of McGhan silicone gel-filled breast implants.

We found that 89 to 96 percent of all the patients in the core clinical trial were satisfied with their breast implant results at two years post-implantation.

Furthermore, we see that the complication rates for Inamed's silicone gel-filled breast implants were in most cases lower than the complication rates for Inamed's FDA-approved saline-filled  breast implants.

Lastly, exhaustive studies conducted by numerous prestigious institutions have found no causal relationship between silicone gel-filled breast implants and systemic diseases, such as breast cancer, connective tissue disease, or other autoimmune diseases.

At this time, on behalf of Inamed, I would like to express our appreciation for the opportunity to present our scientific data in support of this PMA application seeking FDA approval of McGhan silicone gel-filled breast implants.

The efforts of everyone at FDA, or our clinical investigators, our technical experts, and many, many other support personnel are greatly appreciated.

Thank you.

DR. WHALEN:  Ms. Inamed, I will start off with the questions.

Much of the data you presented to us is looking at two- to three-year data and as you rightly pointed out, that is fulfilling FDA's requirements in terms of what you needed to look at.  Whether or not that will serve to judge safety and efficacy will obviously be the crux of what the panel grapples with over the next 24 hours or so, and it may be a continuous 24 hours the way we are going.

Be that as it may, why, in looking at a long-term study that you described earlier in your presentation, in view of what seems to be at least a plethora of anecdotes, and I would submit a reasonable amount of data of problems beyond the first decade, would you consider sufficient long-term data to be restricted to a 10-year point?

MS. KUHNE:  To be restricted to 10 years?  I am sorry.

DR. WHALEN:  Well, as I understood your long-term data design, you had two phases.  One is year 1 to 5 post-implant, and the second is year 6 to 10.  That, to me, means that 10 years, you have judged that you have long-term data.  That, to me, makes no sense.

MS. KUHNE:  Well, currently, that is the way the protocol is designed.

DR. WHALEN:  I guess that's my question, then, why would you design it that way when it doesn't seem to judge what long-term data should be for this particular device.

MS. KUHNE:  Well, we believe that the data shows that it is safe and effective in combination with all of the other data and literature that has been published on breast implants in the past decade.

Perhaps Dr. Spear, who is our clinical investigator, and Chief of Plastic Surgery at Georgetown University, could speak about his long-term experience with silicone gel implant.

DR. WHALEN:  If it is going to be presenting data other than what is in your application, then, we are not going to be entertaining it.  This is really more of a strategic design question in view of what is a overwhelming public concern about the long-term safety of this device.

If you want to editorially comment upon that, then, we will listen to it.

DR. SPEAR:  My name is Scott Spear.  I am a paid consultant to Inamed and the medical monitor of the clinical research.

Actually, I think that is a very interesting question.  I think when these studies were designed, a 10-year milepost seemed like a good place to start.  I don't think anybody would tell you the follow-up of patients should stop at 10 years.  I think that is something that the FDA will have to sort out with the manufacturer as to whether or not they want to continue to track patients longer.

I think realistically, in all likelihood between now and 10 years from now, we may see other devices come forward which will be new and improved over these devices and it becomes kind of an infinite issue, you know, how long do you track something before you come up with something that is even better still.

I don't think anybody would argue against tracking patients longer.  I think that that is a perfectly reasonable thing to bring up.

DR. WHALEN:  The second question is easier.  When you presented the comments of the plastic surgeons who removed these devices upon gel extrusion, I assume what you are relying upon is simple gross inspection of the surgeon at the time of the operation and nothing more rigorous.

MS. KUHNE:  Yes.

PROF. DUBLER:  I want to come back to the two- and three-year issue because it puzzles me.  I am assuming that the company had some long-term plan to bring these back on the market, and why limit your date to two to three years when it seems as if the five-year mark is really an interesting one, and when the continued rate of failure grows over those years, so why did you limit your data collection to two and three years?

DR. SPEAR:  I can speak to the fact that I don't think it has been limited to two or three years.  I think you are seeing a snapshot of the data at three years, which has been arranged with the FDA as a reasonable time to look at the data.

I think there is precedent in terms of other devices in terms of doing that, but I don't think is anybody's intention to stop collecting data.  I think it will be collected again at five years and the MRI studies are designed to go out until nine years.

PROF. DUBLER:  I understand that going forward, but these conversations began with the FDA over a decade ago, and they have been implanted over all this time in protocols.

Where are the data that give us the experience of the company with those implanted devices over the last decade?

DR. SPEAR:  Well, the saline studies are at five years, we have five-year data for that, and I don't think the process between the FDA and the manufacturers has been all that simple in terms of even putting together the studies.

I mean this study was not started until three years ago, McGhan or Inamed didn't start its adjunct study until 1998.  I think it is a very reasonable question to ask why didn't that happen in 1992.  You would have to ask the FDA why it took from 1992 to 1998 to get that study rolling.

I think there was a lot of issues in front of the FDA at that time, but we do have the saline data five years. You have these devices in Europe, in many countries, the whole time, and there is data available from that, and that is what is quoted in some of this epidemiological information.

Frankly, I don't think anybody would mind seeing longer term data, but I am not sure many of these things are going to change long term.  I think the rupture rate is one thing that you would want to know about and certainly people suspect it will increase like everything else with time, but many of the other things are not probably time-dependent, such things as rippling as paresthesia, and things like that.

DR. NEWBURGER:  I have a question which is of the explanted prostheses, has the capsule in the area from which they have been explanted been removed, have biopsies been done, has the histologic reaction been characterized, have they been analyzed for presence of silicone?

Then, I have, after you answer that, one other question, please.

DR. SPEAR:  I don't think that was part of the Inamed PMA presentation.  That has certainly been done experimentally in other research, and depending upon what device and what the circumstances were, you can sometimes find silicone in the capsule.

I think the literature is pretty consistent that if there is silicone in the capsule, it is usually confined to the capsule, and once you get further away, it very rapidly drops off to a normal baseline.

DR. NEWBURGER:  I am interested in what is the mechanism whereby some people would form very painful capsule reactions which would require reoperation and others are just able to live with it.

DR. SPEAR:  Well, there are different theories as to what causes the capsular contracture or capsule reactions.  What is good in terms of these devices is that the incidence that is reported in the study is much lower than with earlier generation devices before the last FDA hearings in 1991 and 1992.

Probably the most widely accepted theory is that there is a low level of bacterial contamination that surrounds some implants, that causes a capsular response, but I am not sure that is the only cause.  It is probably multifactorial.  It can be related to hematoma, it can be related to multiple surgeries, and sometimes you just can't figure out what it is, but in the data for this study, in the augmentation group, it runs at 8 percent at three years, which is significantly lower than what you see in historical literature.

DR. NEWBURGER:  And the second question I think Dr. Spear might be the best to address it, could you explain to those of us on the panel who do not routinely implant breast prostheses what would be a difference in implantation technique for silicone versus saline that might in some way account for some of the differences in the reactions that you see down the line?

DR. SPEAR:  Well, I don't think it is very much technique-dependent.  About the only thing I can imagine that is different in the technique is the size of the incision.  A saline implant is typically put in uninflated and then it is filled, so it is a smaller incision, which would explain perhaps one of the complications which is a scarring, which is more with silicone.

In terms of the other complications, I think it is more the behavior of the device than the methodology of putting it in.  Saline implants do behave somewhat differently than silicone implants.

DR. NEWBURGER:  Thank you.

DR. ANDERSON:  In the public sessions this morning, there was a lot of discussion about rupture.  I note that you at no point in your data analysis did you break out the few patients that had ruptured to talk about their experience to see if, in fact, they really were similar to the other group.

Why is the rupture data not shown to us here?

DR. SPEAR:  You mean in terms of the sort of history of those patients, what kind of situation they had?

DR. ANDERSON:  Your argument, the argument that is being presented is that this is a safe and effective device and that it is equivalent or better than saline, but the concern about rupture between saline and silicone is at least theoretically different because of the persistence of gel.

But you didn't show us any data about your rupture patients, what happened to them, and I think that that would be helpful if we knew more about those patients.

DR. SPEAR:  I think the data, at least some of that data is in the presentation.  I think the most significant piece of data is that there was no extracapsular silicone in any of the rupture patients.

I think there were some 26 implants that were ruptured as of three years, and of those 26, there was no extracapsular silicone.  From a surgical point of view, it means that when the implant is removed, it is relatively easy to remove at least all the gross silicone.

Beyond that, if you are asking for sort of endpoints in those patients, did they all have reimplantation, I would have to go back and look through the data, but I think the vast majority did, because very few patients in the study had implants removed and left out, so almost all the replacement/removals ultimately had replacements with another device.

So, there were very few patients who dropped out and had their implants left out, and there was no extracapsular silicone in any of the patients, so it would leave me to believe at least that nobody suffered any harm from the silicone that got outside of the implant, at least that we could see.

DR. LI:  I think it was asked once before, but could you review how you looked for extracapsular silicone?  Was it you just looked and see if it was there or not?

DR. SPEAR:  I suppose each surgeon has his own way of looking for extracapsular silicone, but usually, each surgeon was asked, when he dealt with a implant rupture, to answer the question was there extracapsular silicone.  It usually presents itself as a lump outside the capsule.

So, if you have extracapsular silicone, it's a palpability issue, you can feel it or see it because the tissues will look like they have silicone in them.

DR. LI:  So, it doesn't rule out, for instance, diffusion of low molecular weight materials or if it doesn't clump, in other words, there could be several grams of liquid silicone, if you will, or of different molecular weights around the extracapsular area and you would never see it, is that correct?

DR. SPEAR:  I am not sure I would agree with the "is that correct" question, but is it possible that there were small amounts of silicone which were not grossly detectable in the capsular soft tissues around an implant, I think that is possible.  It was not grossly detectable.  Whether there is microscopic or other methodologies for picking up very small amounts, I think that is a possibility.

DR. LI:  I guess I was asking this in conjunction I think with what Mr. Powell said, in your retrieval collection, you did not see any retrieved components that lost more than 5 grams of material.  Well, 5 grams is actually a lot of material.

In orthopedics, a half a gram would cause osteolysis in the same time period, so 5 grams is actually a whopping amount of material.  So, if you had 5 grams distributed kind of uniformly rather than in a lump, that possibly could be missed, is that correct?

MR. POWELL:  Let me just correct that.  What I mentioned was the fact that upon returned devices' evaluations, the weights are always within the specified weight for that particular implant.

I think in our preclinical study where we injected unsupported gel into the subcutaneous portion of the rat, where actually only I think 0.06 percent of the total mass of the silicone was not at that site of administration of the gel, showed that the gel really doesn't move around in the body.

DR. LI:  I don't want to jump ahead too much in the detail, but in your retrieval collection, I think you had something like 339 retrievals and 133 of them were evaluated as having rupture, so even in those retrieved devices that had rupture, there was no apparent weight loss?

MR. POWELL:  No, I didn't mean to infer that.  Where they could weigh it, and they were able to weigh the intact devices, there was no gel apparently missing from the devices.

DR. LI:  So, for those devices that were retrieved, and appeared to have an opening, substantial or significant amounts of silicone could have been lost?

MR. POWELL:  I am not aware of the complete data in that regard, but I don't believe--I would have to find the data out on that.

DR. LI:  It is somewhat hard to believe that you could have a significant rupture or opening and have no loss of silicone, because if that was the case, we wouldn't really worry about ruptures, right?

MR. POWELL:  Right.

DR. LEITCH:  Perhaps you have some imaging data about the patients who had rupture that, you know, you had MRI documentation of some of those ruptures which should tell you whether or not there was extracapsular extravasation, and I expect some of those patients had mammograms, as well.

DR. CROTTEAU:  Back to the question that was asked about the explanted devices.  My name is Patrick Crotteau, I am Vice President of Quality Systems for Inamed Corporation.  I do have a financial interest.

In our explant analysis, when a device has ruptured, we have found returned devices that are fully intact, partially intact, or just pieces of shells depending on what has occurred during the explant surgery, and based on our analysis, you can't make a determination where that silicone was lost.

What we did determine is on returned explanted intact devices, that there was no apparent loss of silicone based on our specifications for filling those devices during manufacturing.

DR. LI:  Just to clarify, so on those devices that came back, say, the ones you classified as being broken or having a very large opening, did the surgeon at that time report any extracapsular silicone?

DR. CROTTEAU:  We would have to look at the actual clinical investigation forms.  Typically, we don't get that information with the returned device, so we would have to go back to the clinical information to determine what those comments were from the physician.

DR. LI:  Well, you see what I am driving at, that perhaps the operative report that says whether or not you see extracapsular silicone may or may not be a good indication that the silicone actually leaked into that extracapsular area.

DR. SPEAR:  Maybe to wrap this up, at least in clinical practice, I mean there is an entity of extracapsular silicone that very experienced surgeons have seen, and it is an entity that happens.

All I can tell you is that in the study population by multiple surgeons, I guess there were 60-some sites, none of these 26 ruptured implants had any detectable extracapsular silicone.  That doesn't mean there couldn't be microscopic amounts of silicone, but then the question becomes what is the significance of a microscopic amount of silicone in that capsule.  That is where the epidemiological studies begin to come into play - I mean what is the significance of that, if any.

Regarding the MRI question, as I know it, the MRIs did not reveal any extracapsular silicone either, and there were some 300 MRIs done, and there was no extracapsular silicone including I think it was 18 or so implants that were diagnosed as ruptured on MRI.  So, the MRI didn't show any extracapsular silicone either.

DR. LI:  My normal background is in the area of orthopedics, but in orthopedics and joint replacements, the most biologically active particles that we are worried about are ones that cannot be seen by eye.  So, the fact that they are small doesn't necessarily mean they are benign.

DR. BOULWARE:  As a rheumatologist on the panel, I would like you to clarify the slide 83, the American College of Rheumatology.  In the spirit of disclosure, I am a member of that organization, currently sit as a board of director.

Could you tell the independent peer-reviewed research done by them, or was this just a simple statement in mid-nineties that said there isn't existing data at this point to draw a conclusion, because I have asked the ACR if they have a position statement which I needed to disclose, and they did not have an official position statement or a policy on it.

DR. SMITH:  Would you repeat that question, please.

DR. BOULWARE:  Slide 83 indicates that the American College of Rheumatology--and the title says Significant Results:  Independent Peer-Reviewed Research-- the ACR, implants do not cause additional risk for connective tissue or rheumatic disease.

Is that really an independent peer-reviewed research that was done or simply the statement that was issued in the mid-nineties that said there is insufficient data at that time to draw a conclusion, and really not a research study?

DR. McLAUGHLIN:  My name is Dr. Joseph McLaughlin from the International Epidemiology Institute.  It wasn't their own research, it was a pronouncement by the College.

DR. BOULWARE:  Actually, I would be very interested in that because I am not only finishing my third year on the board, but served three years on the committee, so I have attended six years of meetings and asked them specifically if we have done this or sent it to committee, and have not.  I was told we do not have that.

Now, that is just one of many things that you have cited, but I do want that to be clear in the record.

DR. SMITH:  Hi, I am Howard Smith.  I am a rheumatologist.  I am Director of Research at the Cleveland Clinic Health Systems.

The ACR, about four or five years ago, sent out a fax to all the members giving this position. It was an official policy of the ACR.

DR. BOULWARE:  Would you submit that for record then for us in terms of I think that is probably one of the faxes that are sent out to interpret--

DR. SMITH:  But it was a position statement.

DR. BOULWARE:  No, I don't think it is.  I think if you asked the ACR, they would say it is not a position statement, because they are very, very rigid about that.  I am a rheumatologist, like you, and we spend all our time doing nothing but criteria--

DR. WHALEN:  Obviously, it is a significant point, but the right wing is getting antsy, so I have to get to some questions over here.  If you could indeed produce that in writing for us, since we are indeed going to be here again tomorrow, that would be appreciated, or if you can't, please inform us of that.

DR. EHMSEN:  Mr. Chairman, just a couple of comments, please.

First, with regard to the extracapsular or intracapsular gel, I believe that the safety data that we presented demonstrates that regardless of where the gel is located, it does not create any health concerns.  So, that is one point that needs to be made.

Also, I would appreciate it if you would give Professor Brook an opportunity to clarify a point that he raised earlier in response to the question regarding the relationship between the preclinical and the clinical data, if you might just give him a moment, please.

DR. BROOK:  I think I misunderstood Dr. Miller's question, which was the relevance of preclinical data to clinical data.  There shouldn't be a disconnection between these two, of course, but preclinical means just that, and the idea is you put incredibly difficult constraints on these materials.

I mean in the case of the fatigue testing, you take two plates and you put the implant between them and six a half million times you apply at least 30 pounds of pressure to see whether or not there is going to be rupture, and all of them survived at least 30 pounds.

Of course, there is a relevance because you expose these materials to something they would never be expected to see in the clinical situations.  You go far beyond what you anticipate is going to be the real exposure.  So, in that sense, there has to be a relevance between the preclinical and the clinical data.  You are trying to make sure that a material that is going to go into the clinic, every expectation is it is going to pass all of the requirements once it gets into the clinic.

DR. CHOTI:  Three quick questions.  First of all, regarding that, if that is the case, then, why is there still a reasonably high rupture rate even at two and three years, if it is put through these rigorous tests, what are these women doing within three years that there is still rupture rate, and the rupture rate keeps going up?  I mean if it is so rigorously tested in vitro.

DR. BROOK:  I think I am probably the wrong person to answer that because I can talk about fatigue testing outside of the body.

DR. CONANT:  May I comment on that?  Again, I am a breast imager, but what I hear from women is that there is a shearing, there is a motion, their arm raised, that is retropectoral, there is a shearing of the implants that is not reproduced well in the laboratory by direct compression.

There is a twisting of the implants, possibly wrinkles, as they have described, weakening at wrinkles.  I would love to hear a comment on that, but it is a very different in vivo situation, the stresses that an implant undergoes.

DR. SPEAR:  I think there are two answers to that question.  One is the retrieval data on broken implants or ruptured implants, and one is just clinical experience.  Clinical experience is that some implants unfortunately are probably damaged in the process of insertion.  What percentage of this couple of percent of the ruptures that we see at three years, I don't know, but some are, and they will show up apparently as smooth, sharp openings on retrieval.

I think you are referring to the issue of folding, that the testing done in preclinicals is a pummeling test, it is not a tumbling test or a folding test, and it is possible that some of these devices will fail because of repetitive folding and creating a fold crease.

I think that is the reason why despite the fact that they do so well in these preclinical lab tests, a combination of the surgical trauma and then the real life trauma, which is not just pummeling, is I think the explanation about why you see failure anyway.

DR. CHOTI:  But if that is going up over time, it doesn't suggest that it's an acute process.  It suggests there is some weakening or some process over time, not just from the mechanism of putting it in.

The second question is just is there any change, any improvements in the product, in this product, compared to the one that was looked at in 1992, or is it the same product?

MS. KUHNE:  It is basically the same product as it was 10 years ago.  I mean there has been some minor manufacturing efficiencies, the gel crosslinked a little bit more, but basically, it is essentially the same product.

DR. CHOTI:  So, the problems that were addressed over a decade ago, there was no modification on the device at all from over a decade ago to this product?

DR. SPEAR:  Let me answer that.  This is not a preplanned answer to that question, but back in 1991, these devices were not turned down maybe because there was data they were not safe, the problem was that there wasn't enough data for these devices.  They were being looked at in the context of two or three generations of devices which had never gone through a PMA process because they were pre-1976 devices.

So, I don't think that these devices that were presented in '91, '92 FDA hearings were ever characterized as being unsafe.  It is just that the FDA said we need data for these devices.  They are an improvement over the historical devices, which a lot of the ill will associated with breast implants is from, going back into the seventies and eighties.

So, this is a different device than was used in the seventies and eighties, but it is not a device that is much different than presented in the PMA of 1991 and 1992.

Did you want to say something, Joe?

DR. McLAUGHLIN:  This is Dr. McLaughlin again from the International Epidemiology Institute.

We published the first incidence rupture study of breast implants ever done in the July issue of The Archives of Surgery, and we followed the women for two years, and the rupture rate for the type of implant that is now being used, the so-called third generation, is approximately 1 percent per year.  It parallels the rupture rate of the Inamed product.

Up until that study, in the literature when you read about rupture rates, they were all basically cross-sectional prevalence studies or explantation studies.  So, the rupture rates are not as high as people have been led to believe.

DR. CHOTI:  I just wanted to clarify that the product is really the same in a decade.  This isn't a fourth generation, it's the same.

The last thing is just I am a little bit troubled by this long term.  You know, everybody made an emphasis that women are informed that this does not last a lifetime, and yet there hasn't been really clear guidelines as to is there a time that you recommend--this may be beyond the clinical study--but are you recommending replacement at 10 years, at 12 years, and if so, based on what?

DR. McLAUGHLIN:  Again, in our incidence study, we found that at five years, 98 percent of the implants were still intact, and at 10 years, 85 percent were still intact.

So, then again, the amount of rupture is a lot lyre than people imagine.

DR. CHOTI:  Do, just to clarify guidelines, leave it in until it ruptures, replace it at a certain time point, what are you recommending as you get beyond 10 years?

DR. SPEAR:  It is interesting, this came up in 1991 and '92, when the FDA decided to limit the availability of the devices, and the question was what do you tell women who currently have implants.

The question, the answer really hasn't changed very much in my opinion.  I think the advice to women, whether they have an implant for reconstruction or for augmentation, is that they should be seen periodically by their surgeon every--I think one year is too short, five years is too long--every two to three years to be monitored for evidence of problems, and the most likely problem would be a silent rupture.

The good news is that the term "silent rupture" tells you that these are usually asymptomatic, in other words, they don't cause symptoms, they don't make people sick, they don't cause physical problems in the vast majority of patients.

But on the other hand, I think, as has been pointed out, we don't want to see silicone outside of the capsule, and I think if it's outside the implant, there is probably a greater likelihood it will get outside the capsule with trauma, that they should be replaced once there is a diagnosed rupture.

So, I think the advice would be not that every implant that is 10 years old should be removed, but that every patient with an implant should be seen periodically because at least most surgeon's clinical experience, if the patient doesn't have a capsular contracture, you can diagnose a ruptured implant just by physical examination. You don't need a $1,000 MRI to do it.

DR. CHOTI:  So, you are suggesting guidelines leave it in until it ruptures.  That is basically what you are saying?

DR. SPEAR:  Well, I can tell you from clinical practice if someone comes in to see me and they have concern, and they have had an implant in, let's say, 15 years or 20 years, and they say I am worried, I have had an implant for 15 or 20 years, what should I do, and I tell them if you are worried, the fact is that your risk of rupture is time-dependent, maybe you should go ahead and replace it and reduce your worry.

If they are not worried, frankly, and they have no clinical symptoms or signs of a problem, I would tell them just leave it in, be seen every two or three years.  You can usually diagnose a problem just on physical examination.

DR. WHALEN:  Brent, do you have a question?

DR. BLUMENSTEIN:  I have a number of questions about the methodology.  On slide 43, you mention that what is represented, I assume on the subsequent slides are Kaplan-Meier estimates of three-year event rates, and I just wanted to confirm that that is the case.

DR. BECKSTRAND:  That's true.

DR. BLUMENSTEIN:  Then, these estimates do not include censoring for prior complications as specified in the guidelines?

DR. BECKSTRAND:  That is true.

DR. BLUMENSTEIN:  Then, subsequently, on slide 72 and some other slides, you present comparison of the Kaplan-Meier estimates from the silicone-filled implants to the saline-filled implants.  Are the estimates for the saline-filled implants, Kaplan-Meier estimates without censoring?

DR. BECKSTRAND:  That is true, exactly the same analysis.

DR. BLUMENSTEIN:  You recomputed them, because a few years ago, when you were presenting the saline-filled estimates, you were censoring for prior events.

DR. WHALEN:  Excuse me, can you just identify yourself, so that we know who you are.

DR. BECKSTRAND:  Maggie Beckstrand, Inamed Corporation.  I have a financial interest.

DR. WHALEN:  And your capacity with them?

DR. BECKSTRAND:  I am sorry?

DR. WHALEN:  You are their statistician?

DR. BECKSTRAND:  Senior biostatistician.

DR. WHALEN:  Thank you.

DR. BECKSTRAND:  Actually, when we presented the saline data, that was a misinterpretation, and we were not censoring for previous complications.

DR. BLUMENSTEIN:  So, your estimates then are in compliance with the guidelines that the FDA put out in February.

DR. BECKSTRAND:  That is true.

DR. BLUMENSTEIN:  You present a lot of different complications and a lot of different Kaplan-Meier estimates of risk at three years.

Why are you choosing not to pull all this together in one composite risk of, say, something to bad thing, as a composite of women who experience something bad as opposed to women who do not?

DR. SPEAR:  I think much of how the data is presented was by the discussion with the FDA about how the data be presented.  I don't think anybody would have a problem saying what percent of patients had any complication as a total, but many of these complications are misleading in the sense that one person could have had four things, so the total number could be more than 100 percent if you did it that way, but if you wanted to look at how many patients had some problem or other, I think there would be no problem doing that.  It just wasn't the way the data was presented,  I think more or less by agreement.

DR. BLUMENSTEIN:  To me, it doesn't really show the single thing that I would think that most women would want to see, and that is what is my risk of having something bad happen.

DR. BECKSTRAND:  We did actually break it down into three different types of bad things, which would be the general breast surgery risks, breast-related cosmetic risks, and breast-related non-cosmetic risks.

We have those numbers.  For example, you will see the slide, it's general breast surgery complications, and lists all of the complications included in that Kaplan-Meier, three years' augmentation, 43 percent of the patients experienced a complication general breast surgery, and then we have another slide for the cosmetic complications, 9 percent, other than non-cosmetic complications 13.

So, we did not choose to lump it all into one big, any type of complications.

DR. BLUMENSTEIN:  So, you don't have a back-up slide with such an estimate on it?

DR. BECKSTRAND:  Where it brings together all three?  No, in advisement with Dr. Spear, we chose not to, because we thought it was very important to break out the general breast surgery complications that occur.

DR. BLUMENSTEIN:  I can see the sense in breaking it out, but it would also seem useful to have one composite bad thing, risk of a bad thing.

It really brings me to another issue.  I see that you are choosing to present the three-year Kaplan-Meier estimates.  I don't see any graphs of risk either as 1 minus Kaplan-Meier or whatever, however you choose to do it.

It seems to me that this becomes extremely important when you talk about things like reoperations because the timing of the reoperations would seem to be a big factor in looking at the data as a whole as opposed to looking at the cumulative risk of pre-operations at a single point in time, say, three years.

DR. BECKSTRAND:  We have got some graphs.

[Slide.]

This would be our risk at 1, 2, and 3 years Kaplan-Meier.  It is not the survival curve, but it's at those three time points, this is the risk.

DR. BLUMENSTEIN:  I personally would prefer to see a more conventional Kaplan-Meier presentation.

On slide 60, this has to do with reproductive effects and lactation effects, and I was just curious about what the denominator was for the computation of those risks and the appropriateness.  Apparently, you used the denominator of the whole cohort of patients studied, and I am curious as to what your rationale for that is.

DR. BECKSTRAND:  Right, we did use the denominator for those, but we also looked at it from the standpoint of how many patients, in terms of lactation, how many patients attempted to breastfeed, and that denominator was 32.  There were 5 lactation problems out of 32 patients who attempted to breastfeed, which was 15.6 percent risk, which was the same as--there is a 15 percent problem with lactation pre-implantation in the augmentation group, as well.

We originally looked at it with the whole cohort, and FDA asked us to look at it from the other perspective.

DR. BLUMENSTEIN:  Second to last question.  In the slides 79 through 89, this is just like fingernails on a chalkboard to me, but on the slides you say there is, well, like the one that we were discussing a few minutes ago with the American College of Radiology, it says, "Implants do not cause additional risk of connective tissue or rheumatic disease."

Other slides, you say, "There is no evidence of an immune"--like just above that you say--this is on slide 83--you say, "No evidence of an abnormal immune response."

The issue here is that these are studies that looked for something and failed to find it, and the conclusion of no cause or no relationship, and so forth, is not a correct conclusion.  If you look for a needle in a haystack and you fail to find it, you can't say that there is no needle in the haystack.  You have to say you failed to find it, and then you have to characterize how hard you looked, with a sensitive metal detector or just rummaging around.

So, the language that was used to describe this often did indicate that there was no evidence or something like that, but your slides actually state these as definitive conclusions.

The other question I had was do you folks market a silicone breast implant in Europe?

MS. KUHNE:  Yes, we have been marketing it there for over 25 years.

DR. BLUMENSTEIN:  Do you have data from studies in Europe?

MS. KUHNE:  Not data that would satisfy FDA's requirement for clinical endpoints.  There has been no controlled clinical trials in Europe on Inamed's breast implants.

DR. WHALEN:  Dr. McGrath.

DR. McGRATH:  I have a couple of sort of clinical questions.  One, I am a plastic surgeon and user, and I appreciate seeing the percentages on the silent ruptures and the rupture rate, but I want to know why they ruptured.

Now, we have danced around this question before and no one answered it, and what I want to hear from you is do you know why these implants break down over time?  What happens to them?

DR. SPEAR:  Dr. McGrath, I think the answer I gave you is going to be the answer, that some are ruptured because of trauma at the time of insertion, not all of them obviously, and over time some rupture or fail I think because of the fold flaw method.  I think the fold flaw is probably the most likely suspect over time.  It is the thing that they are most vulnerable to.

They are very strong in some ways, but they don't handle--

DR. McGRATH:  These are your assumptions, Dr. Spear.  Is there any data on this question?  I think we need to understand whether the polymer is thinning or the seal is coming off or what is happening to these things that they fail.

DR. CROTTEAU:  In our retrieval study, the primary purpose of the study--and this is a study that is conducted both for saline and silicone breast implants--we are trying to determine modes of failure for each type of device.  On both studies, they are ongoing studies of looking at returned devices, we do not see any changes in the shell physical characteristics when we can measure it.

As far as the elongation, the strength of the devices, on explanted devices, we do not see any change in thickness or thinning of the shell.  We do not see any weakness in our patch-adhered joint.

What we do see is in the case of the silicone devices, 18 percent of devices have manifestations of surgical damage, so when we get devices, we go back and we evaluate the devices both macroscopically and microscopically to determine if we can tell how the device failed.

We have studied typical surgical damage and have a set of standards that our technicians compare device damage to based on forceps, sutures, scalpels, and characterize those type of failures.

We also have an ongoing study that shows or tries to determine why we have both smooth-edged openings and sharp-edged openings in devices.  When we have a smooth-edged opening that is typically and has been demonstrated to be caused by a fold and actual abrasion of the devices, this is not a typical mode of failure for silicone devices, it is a more predominant mode of failure on saline devices.

For sharp-edged openings, beyond surgical damage, we have an ongoing study to determine why those are occurring, but we have not determined statistically any modes of failure that determine why we have a sharp-edged opening in the device.

This is part of an ongoing program, part of our conditions of approval for saline, and again part of an ongoing program on silicone gel.

DR. LI:  Could I follow up on that?  In looking over your retrieval data, I was a little confused over the smooth edge/sharp edge description because those aren't normally materials terms that describe rupture.

But I was a little confused over this issue about what you call smooth edge openings occurring at a fold.  Now, in the breast implants retrieved that I have got, many of them are quite supple and there are no permanent folds or deformations in the device.

So, are you saying that some of the retrieved devices come back with a permanent crease in it?

DR. CROTTEAU:  You can actually see physically a crease in the shell, and at the apex of the crease, you see a thinning and then a wearing through of the shell.  This is much more evident in saline devices, but we can determine, you actually see a crease in the shell.

DR. LI:  So, this is a permanent deformation of the shell?

DR. CROTTEAU:  Well, it is either a permanent deformation or is a shadow left in the shell.  Again, it is a manifestation where you can see where the device was creased and then again it shows that typically at the apex, that crease, thinning of the shell.

DR. LI:  For what you call your sharp edge openings that you said didn't appear to follow a crease line, is it possible that there was a fold while it was in the patient, that when you took it out, it was still soft enough to where the fold disappeared?

DR. CROTTEAU:  That's possible.  Again, we cannot determine what is causing the majority of the sharp edge openings at this point in time, but we do not see any changes in mechanical properties or physical properties of devices that are explanted.

DR. LI:  Did you do any microscopy?  Oftentimes you could see, for instance, if it's a scalpel mark, if the fracture is outside-in, or inside-out.

DR. CROTTEAU:  Yes, and we actually do photomicrographic analysis when we need to.  We have photomicrographs of typical surgical damage that the technicians use when they compare the devices.

DR. LI:  So, for the ruptured devices that you believe were surgically induced, those were all outside-in ruptures?

DR. CROTTEAU:  Well again, if it is in the case of a suture, you will see an entry and an exit, so you can't say they are all outside-in, you can depending on the type of damage.  It it's a forceps damage, you may see that it has been damaged actually on the inside of the device and then propagated to the outside.

DR. LI:  One more detailed question on the surgically-induced ruptures.  First of all, these shells are relatively thin, they are three-tenths to 0.45 millimeters thick.  So, you are saying that it is possible essentially if you nick the wall with a scalpel, that you have nicked it enough, so the surgeon wouldn't detect it as a leak, but you have only gone through 0.1 mm instead of 0.13 or 0.3 mm, so wouldn't you expect those ruptures to happen earlier?

So, is there a correlation between those types of ruptures and the length of implantation?

DR. CROTTEAU:  Again, when we receive the devices back, they have failed, so we can't determine whether it was a nick that then propagated into a failure or it was a cut through.  Now, typically, when you look at them, you can tell what the initial contact was, but you can't tell when that propagated to a point of failure.

DR. LI:  Just one last question for the moment.  Of the 339 retrieved devices that you reported on, were all these 339 from the three cohorts, or were they from the 79,000 implants that you have put in?

DR. CROTTEAU:  No, those 339 are from these three cohorts.

DR. LI:  Are exclusively from these three cohorts.

DR. CROTTEAU:  Again, we have ongoing analysis of other devices, but the data that was presented represents the clinical study.

DR. LI:  Because it's a little misleading because in some of your tables, you used as the numerator 79,000 devices and reported percentage, so that is a little misleading.

DR. CROTTEAU:  That is I think in our postmarket information, not in our explant or retrieval.

DR. WHALEN:  Dr. Conant, did you have a question?

DR. CONANT:  Again, I am a breast imager, but I have got a couple of questions about the patient satisfaction side of it.  I find it a little contradictory that when questioned about how satisfied they are with their implants, 89 to 96 percent of the women were very satisfied, but at the same time, at two years, the reoperation rate was so high, 20 percent, 45.9, 33 percent.

I am questioning how patients were queried on satisfaction, and if I take the lowest level of that group, at 89 percent, still more than that number had reoperations. I can't imagine being satisfied if you have to go back to the OR in the same period of time.

Could you clarify that?

MS. KUHNE:  For the general overall satisfaction rate, it was just a question asked at different points in time postoperatively how satisfied are you with your breast implants on a 5-point scale.

Indeed, as you indicate, I mean many of these patients had complications, but that didn't cause them to be unsatisfied with their breast implant surgery in general.

DR. CONANT:  So, they were questioned at multiple times during their three years.

MS. KUHNE:  Right.

DR. CONANT:  Were they pooled responses?

MS. KUHNE:  Right.  The satisfaction question that you are referring to was not a pre- and post-satisfaction, it was just post-implantation annually, and they are asked that question how satisfied are you with your breast implant.

DR. CONANT:  Did you look at this relative to whether they had been back to the operating room, whether that group was less satisfied, whether that impacted on their satisfaction?

MS. KUHNE:  We didn't look at each patient in terms of, you know, specific for unsatisfaction, but we did look, for instance, at instances for, you know, capsular contracture is a very common complication, but only like 18 percent of those patients indicated that they were not satisfied.

DR. CONANT:  My next question is since we only have two years I think for reoperation data, I am wondering if you all were able to extrapolate.  You showed when he asked, I am not a statistician, he asked for the Kaplan-Meier and showed three different curves slightly increasing over time, not a straight line, if you extrapolated to five and 10 years in terms of reoperation, because as you said, quite a few of these reoperations were capsular procedures which were in many cases I think repetitive procedures, whether you could comment on what the expected reoperation would be at later years like five and 10.

MS. KUHNE:  I don't think we can really extrapolate from the information.  I don't think there is a mathematical way to do that.  I don't know if Dr. Spear has a comment on what he would expect from the clinical standpoint.

DR. SPEAR:  Well, I think most of us have learned that things deteriorate with time, it doesn't matter what you are talking about, and I think as patients have implants longer and longer, the number of patients who have a reoperation is going to come up for whatever reason.

I am not sure it will all be for capsular contraction.  I think what that data shows me, interestingly, is that even though the numbers are there, that from the patient's perspective, the numbers of this and that don't change the fact that 90 percent plus, particularly of the augmentation and revision patients, are happy with their device despite the fact that 20 percent have a reoperation, despite the fact that 8 percent have a capsular contracture, they are still overall happy with their device.

A few patients have removed their device because they weren't happy, but when you look at the totality of the thousands of patients in the study, 90 percent are still happy despite the fact that we have listed very definitely everything that one could think of that someone could have a problem with, from bruising to itching to paresthesia, I mean a pretty definitive list of things that people could answer yes to, but with all those things in there, still 90 percent plus were satisfied.

DR. CONANT:  Three years.

DR. SPEAR:  Yes, and I expect that with time, it might change.  You know, things do change, but it didn't change much from one year to two years, it changed a very small amount.

DR. CONANT:  I have just a couple short comments and questions.

You compared the saline versus the silicone in terms of rupture, and the saline had a much higher instance of rupture, still not very high, I think around 5.6, something like that.

I think that is a very difficult group to compare, because we know that when saline implants rupture, they deflate, and one knows that, patients are aware of that,  versus dealing with this silent rupture.

I just think that that has to be said to the audience because it's a very different clinical situation.

DR. SPEAR:  Can I respond to that?

DR. CONANT:  Sure.

DR. SPEAR:  I think that is a very valid point, the fact that one is not silent and one is silent, but I think that is why the MRI study was included in the study, and the MRI study still gives you a lower number of silent ruptures than you would have of the saline ruptures.

DR. CONANT:  Which takes me to the next point, which is the MRI study.

DR. WHALEN:  Can I just interject before you get to that point?  I want to again echo that I strongly encourage any and all questions.  We are going to hopefully get to a point where the panel is going to be making comments, so please ask your questions at this point, reserve your comments until a point this evening when we get to that.

DR. CONANT:  Question.  How did you decide the sample size for the MRI study?  I notice the compliance dropped off considerably at year 2.  I am wondering how you recruited the patients, what that plan was?

MS. KUHNE:  The sample size was determined based on the specific precision that was required actually per the FDA guidance document.  It was about 34 percent of the overall patient population is what it ends up to be.

In terms of compliance, we are attempting to get these patients back in to get their MRIs.

PROF. DUBLER:  Dr. Whalen, I had a follow-up question.

DR. WHALEN:  We will get to you in just one second, Dr. Dubler.

DR. LIEBERMAN:  I wanted to follow up a little bit on the satisfaction data and ask how you can square again this very high satisfaction rate with the decrease in so many measures of well being.  You put up that they were above average even at the two- or three-year mark, but most of them had decreased and many of them had significantly decreased.

MS. KUHNE:  I think again a lot of that has to do with the fact that they started so high to begin with and there wasn't much room for them to go but down in some instances.

DR. LIEBERMAN:  It's a pretty consistent pattern.

DR. SPEAR:  Can I respond to that, too?  I think that is a very interesting question.  It has troubled I think some of us, as well as you, but I want to emphasize that the decision to have elective surgery is a courageous decision regardless of what it is, and I think that you get these people at a point when they are feeling very good about things in general.

The fact that even though there is a decline in some of these measures, and even though it is a statistically significant decline, I am not sure it's a clinically significant decline because it is a very small measure, too.  They are still very, very high on the scale.

But I think the point that they start at the top, there really isn't much room for them to move except down.  I mean if you get 100 on the exam, you can't get 101, so the fact is sort of it's just a peculiarity of the measure that they couldn't do better, they could only do the same or worse, because of the fact that they scored perfectly, many of them, at the initial assessment.

But I think the fact is even though it is a statistically significant change, I am at a loss to give it any clinical significance.

DR. LIEBERMAN:  We mostly have yes's or no's in significance.

The other question I had you talked about cancers and you talked about breast cancer, and you talked about rheumatologic disease.  There are some studies that have suggested increases in lung cancer and brain cancer specifically glioblastoma, and you didn't address that point, so I would like to hear what you think about that.

MS. KUHNE:  We didn't address that point because that wasn't an endpoint in the study, and we didn't specifically collect that information, although we did look at, in terms of the numbers of lung cancers and brain cancers that we had on some of the patients who had report of complication of cancer postoperatively, and if you give me one moment, I can find that information for you.

DR. LIEBERMAN:  I think your numbers, I mean given the size of your cohort, the numbers are probably pretty small, so I think part of it has to be based on the literature, as well as what is found in the study.

DR. SPEAR:  Let me just say that in terms of the literature, what is interesting is that the overall mortality of the patients who had breast implants or breast augmentation is lower than the general population.

There are some markers that are higher, there are some that are lower, and I think it is my opinion anyway that I would not give breast implants a protective effect for preventing certain cancers, and I don't think there is any evidence that there has a causative effect for causing certain cancers.

I think that there is some that it is better in and some that it is worse in.  It is probably much more a reflection of the underlying population who got the implants rather than the effects of the implants.

DR. WHALEN:  If you about to present information, this is information in the PMA, but not today presented, or information that was not in the PMA?

MS. KUHNE:  It's in the PMA.  It's we didn't collect those as specific endpoints, we didn't solicit that type of information.

DR. WHALEN:  Then, go ahead, please.

[Slide.]

MS. KUHNE:  These are patients in the reconstruction group of the core study, and there was one patient who had brain cancer.  In the revision group of the core study, I am sorry, there was no patients in the revision group.  In the reconstruction portion of the adjunct study, there were two patients who had lung cancer and two patients who had brain cancer, and, of course, these patients were all ones who first presented with breast cancer.

DR. WHALEN:  Dr. Dubler.

PROF. DUBLER:  Thank you so much.  I had a follow-up question to Dr. Li that I didn't want to lose.  We are very interested, obviously, in what makes these things rupture.  Since surgical technique has been noted by a number of people as a likely reason for some ruptures, what would be your solution to that problem?

DR. SPEAR:  Well, I think it's the reality that you can't control every surgeon.  I think the company has the intention of having an educational program to try to educate surgeons how to best use these devices if they are approved, but I mean as long as there is humans, there is going to be human variability and there will be some surgeons who do better than others, there will probably be some surgical error.

A lot of times it sounds like these errors are not things you can even detect in the operating room.  They may be subclinical trauma to the implant, not enough to actually rupture it putting it in, but enough to stress it or to create a point of propagation so that later on it would rupture there.

So, short of making a very large incision, which plastic surgeons are loathe to do, there is going to be some mechanical trauma with insertion by some surgeons.

PROF. DUBLER:  Do you have any data that would show that board-certified plastic surgeons have a lower rupture rate than others?  How do we bend our minds around this data?

DR. SPEAR:  I wish I had that data, but I don't have that data.

DR. MILLER:  I just have a couple of questions briefly.  It was my understanding that in the last five, six years, that significant changes had been made in the manufacture of the devices, but that is not true?  When was the last time that you made a major change in the manufacture of the devices in regard of the envelope and limiting the gel bleed and that sort of thing?

MS. KUHNE:  I can't answer right now when we made the last change, maybe Pat can answer that, but it is the same device that was used 10 years ago, we have not made any major changes since that time.

DR. CROTTEAU:  But to the manufacturing method, again, the design of the device hasn't changed, but we have made significant changes to the manufacturing method to enhance the reliability of the manufacturing method.

In devices that were made previously, the manufacturing methods tended to be uncontrolled, unvalidated manufacturing methods.  In the last few years, the manufacturing methods have all been validated to demonstrate that they can repeatedly produce the desired device.

So, those type of controls in the manufacturing method have been enhancements made to the manufacturing method surrounding the same specifications to the end device, so that the variability in the product coming out of our production line is much less variable than it was, say, 10 years ago.

DR. MILLER:  Thank you.  I had another question. Because you didn't see in the animal studies or didn't observe in the core study any presence of silicone gel outside of the capsule, are you suggesting that this never occurs or does not occur?

DR. SPEAR:  I am not suggesting that.  In fact, I think one of the things that has changed over the last decade is surgical practices, as well as the design of the devices.  There was a time surgical practices included something called closed capsulotomy, where if you are dealing with a patient with capsular contracture, the treatment was to squeeze the implant with as much force as you can humanly muster and fracture your thumbs and break up the capsule.

Undoubtedly, that was the cause of a great many patients who did wind up with extracapsular silicone because you could rupture the device, rupture the capsule, and send silicone into the soft tissues all in one moment.

I think it is pretty clearly known by board-certified plastic surgeons that this is not a good idea to do, it is discouraged by the manufacturers in the package insert, and I think a change in culture is part of the reason why I think extracapsular silicone is a much rarer event today than it was 15 years ago because the devices are made to higher specifications and I think because surgical practices have changed.

It can still happen, it can still be microscopic amounts.  I think if someone is in an automobile accident and hits the steering wheel, they could break a device and send silicone into the soft tissues, but those are very unusual events, they are not common events.

DR. WHALEN:  So, the capsulotomies now are only done during mammography, I guess, right?

DR. SPEAR:  Well, if you heard the pressures, you know, I am not the radiologist in the group, but pressure exerted with a mammogram is substantially less than with a closed capsulotomy.  I think it's 30 pounds per square inch or less on average.  Closed capsulotomies were much more forceful.  So, I think it is unusual to have an implant rupture with a mammogram also.

DR. McLAUGHLIN:  I would like to get back to a question from Dr. Boulware.  I have in my possession now the position statement.

DR. WHALEN:  Good.  That is something that we have asked that copies be provided to us.  We don't further comment at this particular juncture.

Dr. Olding.

DR. OLDING:  I still have a problem with detection of rupture.  To me, if you accept the possibility or the probability, whatever, that, in fact, distant disease is not a problem here, take that whether you believe it or not for a moment think that, so then it becomes local complications.

If it is a local complication, if it is going to rupture, my patients are going to want to know.  It is called silent rupture presumably because you can't detect it in any manner except perhaps for an MRI.

Why would I not advocate an MRI for my patient on a yearly basis, you pick the time interval?

DR. SPEAR:  I think the term silent rupture really refers to whether the patient is aware of the rupture.  I don't think it would be appropriate to say that it refers to the fact that it can't be diagnosed.  I think it just means that a patient is not aware, it does not have any symptoms.

When I see a patient in clinical practice who either thinks they have a rupture or who has had an MRI for whatever reason, then, comes to see me, I think the definitive test is actually the physical examination.

I have often, I mean I have removed implants when the MRI suspects rupture, and the implant is intact, and I have had patients who come to me with an MRI suspicion of a rupture and, in fact, on clinical examination, I am absolutely positive the implant is intact.

So, I think the term "silent" does not mean that a physician cannot detect a broken implant in the absence of capsular contracture.  If you have a capsular contracture, all bets are off because you can't really examine the implant.

But then the other issue is what is the danger of a silent rupture, and although I don't think it is desirable to have a broken implant in a patient without addressing it, the fact is that the risk of a broken implant that is not symptomatic is very low.

DR. OLDING:  Presumably, that is a step that precedes extracapsular rupture.  So, if you can predict in advance of an extracapsular rupture by some method, and I believe MRI is something like 96 or 98 percent specific and sensitive, somewhere near that, although there are going to be some that it misses, the vast majority it should pick up.

So, again, I have had patients who I have explanted that I did not think were ruptured, and were ruptured, and vice versa.  Would you not want to recommend an MRI?

DR. SPEAR:  I personally don't recommend MRIs to my patients as a routine.  I think it's a very expensive test for very little yield personally.  Now, there is nothing to stop a surgeon from recommending an MRI if the patient wants one or if they think it is appropriate.  I don't think it's the standard of practice at the moment to get MRIs because of the length of time a device has been in place.

It is going to be part of the ongoing study here. There is an MRI study in place for a total of nine years. The incidence you just gave me the numbers again, the silent rupture incidence at three years was 2.7 percent in this particular study.

I just don't think it's a necessary test because although there is a natural progression from intracapsular silicone to the greater risk of it being extracapsular, again, none of the patients in this study had extracapsular silicone at least that was grossly discernible.

I think the chances are for a non-traumatic rupture, in other words, one that doesn't happen from a closed capsulotomy, the chances of silicone getting outside the capsule are still very low.

DR. OLDING:  I would agree with that, but I still know that my patients are going to want to know if they have ruptured implants because that is going to be the decisionmaker for them or most of them.

MS. KUHNE:  I just wanted to clarify something.  We did calculate silent ruptures separately, but those silent ruptures are included in our overall rupture rate.

The rupture rates that are reported for the different cohorts be it the 1.2 percent for augmentation, and 6.3 percent for reconstruction, 3.6 percent for revision includes ruptures that are both asymptomatic and symptomatic, confirmed and nonconfirmed, so I just wanted to clarify that silent ruptures is included in that overall rupture rate, so patients would know.

DR. BRENNER:  My name is Jim Brenner.  I am a radiologist and a paid consultant to Inamed.

A couple of points.  First of all, perhaps not every plastic surgeon is as adept as Dr. Spear at examining patients, and perhaps not every radiologist is as adept at identifying true intracapsular ruptures from folds, so that the discussions Dr. Spear had may represent, in part, variability both in interpretation of the MRI scans, which are not necessarily a reproducible automated test, and the variability of the examiner.

In the first place, in this study, there was on-site reading by the radiologist and there was a secondary reading by a central reviewer to ensure as high an accuracy of identifying intracapsular rupture, indeed, implant failure.

I think of at least one occasion where I was called by the plastic surgeon that his or her radiologist did not believe there was a rupture, I had read a rupture, and he was kind enough to call me back and inform me that, indeed, there had been a rupture, so there was an attempt during the study to maximize the sensitivity of rupture, and not underestimate its presence.

In fact, as you will note from JoAnn's presentation, the two indeterminate readings were also included as probable ruptures, so that when an intracapsular rupture is identified, it's a notable clinical finding.

Now, how you conclude with respect to the possibility of silicone getting beyond the capsule is a matter of discussion for this panel.  A great deal of evidence has been submitted that the biocompatibility of that silicone with respect to the potential pathologic effects doesn't seem to be supported by the literature or to a limited extent in this study.

So, all I can report is what we found during the study.  I should note, because the difficulty in reconciling why implants rupture over time was raised by Dr. McGrath and others, is still not easily reconciled.

If, in fact, the theory, as Dr. Spear suggested, is that folds may be a weak point, I can tell you that in virtually, well, certainly over 90 percent of the implant studies I reviewed in this cohort, as well as those which I have reviewed over the last 10 years, I do see folds.  So,  folds are almost an ubiquitous finding when you look at implants intact, and whether or not they are the source of potential failure is extraordinarily difficult to identify.

Most of the implant failures that I observed in this study, and I have observed elsewhere, do not represent complete collapses of the envelope itself, but rather more focal collapses, so-called inverted loops or subcapsular signs, that may lend credence to Dr. Conant's observation that shearing forces may play a small part if there is a slight adhesion because of oil migration through the implant.

DR. WHALEN:  If I am reading the panel correctly, Dr. Manno's question may be the last for sponsor.  So, everybody knows where we are going, I plan on getting through FDA's presentation before we break for dinner.

I also have strong hopes that we can get through all panel discussions and FDA's questions before we adjourn for tonight or tomorrow morning, whenever that may be.  So, that is where we are headed for tonight.

Dr. Manno, the last question for sponsor.

DR. MANNO:  I guess because of the time, I will limit myself to one question.

On slide 34, it was the core post-approval proposal Phase II, where you say you had mailed surveys.  You said you have a mechanism for being sure that nobody is missed.  Could you tell me again how you are going to do that?

MS. KUHNE:  To ensure compliance.  I didn't mean to allude to the fact that we would have 100 percent compliance for those surveys.  What it meant was that we would follow up carefully with those patients, they are direct mailed surveys, at least that is what the proposal is, and we have done this before, direct mail surveys, and we send out reminders to the patients that the survey is coming up, we send out the survey, and then we send out other reminders and follow-up phone calls to encourage the patients to complete the questionnaires and return them to the company.

DR. MANNO:  Is there any difference in response or any trend that you might see developing when you have to go through, say, the phone calls at the end, because I think it may be documented, and I can't tell you where, at any rate, that a response can change not to your benefit when the people don't respond on the first, second, or third try, and you go to a personal contact, they get belligerent maybe.

I am just asking if there has been any notice in any of these that you have had.  I think that might make an impact.

MS. KUHNE:  Right.

DR. MANNO:  Along the same line with the mailed, it seems to me--well, that is a comment, I will hold it for later.

DR. WHALEN:  I would like to thank Inamed and I would ask Commander Allen to begin the FDA's presentation.

FDA Presentation

ODE and Office of Surveillance and Biometrics

Introduction and Mechanical Testing

CDR ALLEN:  I am Samie Allen, the lead reviewer for this PMA.

For your convenience, we have provided you a hardcopy of the slides.  Please refer to this handout because there are changes to the clinical slides from that posted last Friday.

[Slide.]

FDA's review team was comprised of may reviewers from across the Center for Devices and Radiological Health, as well the Center for Drug Evaluation and Research.  However, this slide shows those FDA reviewers who are presenting today.

[Slide.]

I will now present an overview of the mechanical testing, device description, retrieval study, and shelf life information.

[Slide.]

There are seven styles under PMA review for the indications of primary augmentation, primary reconstruction, and revision.  As a note, not all seven styles were included in each of the clinical styles to be discussed later by Dr. Dawisha.

The styles are available in different combinations of shapes, profiles, surfaces, and volumes.

All styles are single lumen but Style 153.  Style 153 is a double lumen device with inner and outer lumens filled with silicone gel.  The inner lumen is located at the lower pole of the implant and its function is to maintain the curved profile.

All styles are comprised of the same basic components a 3-layer shell, a 2-layer patch, silicone gel filler, ,and a silicone adhesive to seal the hole.  The material specifics will be provided by Dr. Arepalli as part of the chemistry overview.

[Slide.]

As per FDA's breast implant guidance, these are the 3 mechanical tests performed to support a silicone gel-filled breast implant.

[Slide.]

Gel cohesion, as a whole, is addressed by both gel cohesion testing as per ASTM F703, as well as penetrometer testing.  These tests assess the cohesive and cure characteristics of the gel.  All testing met the acceptance criteria.

[Slide.]

Gel bleed testing was performed a per ASTM F703.  The average weight gain for smooth Style 40 and textured Style 110 at 8 weeks are shown here.

ASTM F703 does not include an acceptance criterion for this testing.  The ASTM F703 test methodology was not established to mimic physiological conditions, but instead to accelerate the bleed diffusion process to compare various smooth implant designs.

[Slide.]

Fatigue testing was performed on smooth Style 40 and textured Style 110, both of which were considered representative of the styles under PMA review.  Final, sterilized devices with the thinnest shells were used.

Inamed performed the fatigue testing as per the recommended test method described in FDA's breast implant guidance.

The resulting endurance load levels at 6.5 million cycles runout were 55 pounds for Style 40 and 30 pounds for Style 110.  Based on an expected in-vivo load of 3.7 pounds, this equates to safety factors of 11 and 6, respectively.

Fatigue testing also included determination of the ultimate static load, or force to failure due to a single compression of the implant.  The results were far greater than the 55 pounds expected during mammography.

[Slide.]

The sponsor conducted retrieval studies to better understand the possible modes of gel-filled breast implant failure in vivo, which could lead to improvements in manufacturing, device design, surgical technique, and/or labeling.

The study included 339 implants retrieved in just over a 2-year period.  All 339 implants were those styles under PMA review.  Forty-five percent of these were from the core or adjunct study.  The remaining 55 percent were implanted prior to the 1992 call for PMAs, were returned unused due to an intraoperative observation, or for which a specific link to a study could not be made, such as an unknown serial number.

Physician observation, such as rupture or non-rupture, were made at the time of explantation.

Laboratory observations were made by Inamed to categorize each explant with a primary device characteristic, such a sharp-edge or smooth-edge openings.

Mechanical testing was performed, but this testing was inconclusive in assessing the mode of failure.

Analyses of sharp-edge openings were performed to characterize those created by surgical damage.

[Slide.]

This table summarizes the cause or mode of failure for each of the device characteristics based on all analyses performed.

Smooth-edge mode of failure is fold flaw.  Only 18 percent of the sharp-edge openings could be linked to surgical damage.  The mode of failure for the remaining 82 percent is unknown.

For broken devices, for example, those devices with the shell torn into pieces, possible causes include explantation surgical technique or propagation of a smooth-edge or sharp-edge opening.

For device surface observations, for example, scalloping at the radius or dimpling, one possible cause is excess stress caused by the implantation technique or capsular contracture.

For gel-related observations, for example, particles or bubbles in the gel, Inamed offers one possible cause for the bubbles as excess stress prior to implantation or during explantation that led to air permeating through the shell.

For functional devices, there was no opening or surface or gel observation.  Therefore, there is no mode of failure to assess.  However, these devices included those removed for complaints, such as capsular contracture.

Inamed considers these retrieval study findings to be inconclusive to determine any specific steps to take with regard to improvements in device manufacturing, device design, surgical technique, or labeling.

[Slide.]

With regard to shelf life, device and package testing were performed.

Inamed provided a combination of real-time and accelerated testing on their subject device, as well as on their approved saline-filled breast implants, to support a 2.5 year shelf life on their package label.

Inamed continues to collect data with a goal of supporting a 5-year shelf life.

[Slide.]

In conclusion, FDA considers the gel cohesion testing, gel bleed testing, fatigue testing, retrieval study, and shelf life information to be adequate.  However, the correlation between these tests and the long-term clinical performance of the implant is not well understood.

The panel should consider these results in your safety assessment of this breast implant.

Thank you.

Dr. Arepalli will now present FDA's review of the chemistry data.

Chemical Testing

DR. AREPALLI:  Thank you, Samie.

Good afternoon.  My name is Sam Arepalli.  I am the chemistry and materials reviewer for this PMA.

The next few minutes I will be giving a brief overview of the chemistry/material aspects of this device. Please bear with my, my presentation will contain some highly technical words.

[Slide.]

As Commander Allen stated earlier, the device components are the silicone shell, silicone patch, silicone gel filler, and silicone adhesive.

The shell is manufactured from two elastomers.  They are a mixture of high molecular weight vinyl containing siloxane polymer, a low molecular weight methylhydrogen containing siloxane crosslinker.

A platinum-silicone complex is used as catalyst and the elastomer is reinforced with a treated amorphous silica to enhance toughness.  The shell consists of inner and outer layers, which are also called base layers, and are sandwiched around a "barrier layer" designed to impede the diffusion of components of the gel through the shell.

All layers of the shell are produced using a copolymer of poly(dimethyldiphenyl)siloxane.  The barrier layer differs from the base layer in that it contains a higher percentage of diphenylsiloxane in the copolymer.

The patch is manufactured from two types of silicone elastomer.  The inner layer of the patch is a two-part platinum cure elastomer similar to the shell material, except the polymer is poly(dimethyl-methyltrifluoropropyl) siloxane copolymer.  The outer layer of the patch is a peroxide cure dimethyl silicone elastomer.

The gel is manufactured from a formulation that is approximately 15 percent by weight of a silicone polymer with chemically reactive molecules along the polymer chain that forms a 3-dimensional network.  The gel also contains approximately 85 percent of liquid silicone polymer which is also called silicone oil.

The RTV adhesive used to seal the fill hole in the patch is an oxime-cured silicone with an organotin catalyst.

[Slide.]

The chemical analyses were conducted in accordance with the FDA Breast Implant Guidance Document.  Both the shell and the gel were subjected to the analyses separately.

The first analysis carried out was degree of crosslinking.  The degree of crosslinking in the shell was measured by the Sal Fraction method and was found to be 3.4 crosslinked unit/molecule.  Physical characteristics like break force and elongation were measured for different lots to ensure uniformity of the crosslinking.

To ensure degree of crosslinking uniformity for the filler material, penetrometer studies were conducted on different lots of the gel and were found to be 3.5 to 7.5 mm which meets the sponsor's specs.

[Slide.]

Volatiles.  The device was tested for volatile compounds.  The volatile compounds were analyzed for the shell component and were found to contain 279 micrograms of 1,1,1-trichloroethane and 251 micrograms of isopropyl alcohol.  The gel was not tested for volatiles as it does not contain any volatile solvents.

[Slide.]

Extractables.  This slide talks about the analysis methods for the extractables.  The shell and the patch were separated from the gel filler.  In addition to the shell/patch and the gel of the finished device, virgin shell/patch that is not exposed to gel filler is also used for analyses of extractables.

The shell and the patch together, and the gel filler were separately subjected to exhaustive extraction with n-hexane at 35 degrees C.  The residue obtained was concentrated and was subjected to gravimetric analysis, gel permeable chromatography, infrared spectroscopic analysis. The results were consistent with those of typical silicone polymers.

[Slide.]

GC-MS Analysis.  The extracted residue was subjected to qualitative and quantitative analyses using gas liquid chromatography coupled with mass spectrometer.  The results were provided in FDA's summary panel memo.

The low molecular weight cyclicoligosiloxanes up to D10 are not detectable at the sensitivity level of the instrument used.  Higher cyclic and linear oligosiloxane concentrations of the subject device are comparable to those present in the FDA-approved saline-filled breast implant devices as far as ppm of the implant are concerned.

[Slide.]

Metal Analysis.  Analyses for heavy metals were carried out on shell, patch, and gel filler separately.  The metal analyses showed they are present in trace amounts as shown in Table 2 and Table 3 of the panel pack.

While saline-filled breast implants contain tin only, the gel-filled breast implants contain platinum, as well as tin, are used in the manufacture of the device.

This slide provides the tin and platinum concentrations in the shell, patch, and gel filler separately.

[Slide.]

Analysis for Silica.  Spectroscopic studies were conducted to show that there was no free silica present in the device.  X-ray diffraction studies indicated that the bound silica is present in the amorphous form, not in the crystalline form.

[Slide.]

Finally, in conclusion, the envelope shell and the gel materials were separated and chemical analyses were conducted in accordance with the FDA Breast Implant Guidance Document.

The shell and the gel were tested for degree of crosslinking and volatiles.  Each component was hexane-extracted in a soxhlet and the residue obtained was subjected to gravimetric, gel permeable chromatography, FTIR, and qualitative and quantitative analyses by GC/MS.

The shell and gel were analyzed for heavy metals.

The elastomer shell was analyzed for the presence of free silica and the bound silica was proved to be in amorphous form, not in the crystalline form.

Inamed has performed extensive chemical analyses of their breast implant device.  FDA notes no deficiencies in the information they have provided.  This is important information for toxicology testing that will be addressed next.  Also, FDA recommends that the panel consider these chemistry data in their deliberation on the long-term safety of the implant.

Thank you very much.

Dr. Berkowitz will present the FDA's review of the toxicology information.

Dr. Berkowitz.

Toxicology Overview

DR. BERKOWITZ:  I am David Berkowitz, the toxicology reviewer for this PMA.

The objective of the toxicology studies is to provide an assessment of the toxicological risks likely to be associated with the device.

The gel-filled prosthesis differs from the saline implant in two ways.  First, the potential exposure to the implant materials is about 40 times higher than the exposure to saline implants.  A woman exposed to two, 800 ml prostheses would be potentially exposed to 1,600 grams of silicone gel and about 40 grams of silicone elastomer.

The exposure to the saline-filled prosthesis is only to the elastomers in the two shells.  Secondly, the shell of this prosthesis contains a diphenyl-substituted polydimethylsiloxane which is not present in the approved saline implants.

[Slide.]

The gel-filled breast implant toxicological investigations can be considered under the six categories shown on the screen:  pharmacokinetics, biocompatibility testing, subchronic toxicity testing, which is implantation testing, reproductive and teratogenicity testing, immunotoxicology, genotoxicology, and carcinogenesis testing.  I will present a brief summary of the conclusions drawn from these studies.

[Slide.]

Pharmacokinetics.  By weight, the overwhelming component of the elastomer and the silicone gel are the long chain polydimethyl siloxanes as we have heard.  Inamed tested the fate of these compounds in rats by using subcutaneous implants of a gel made from radiolabeled polydimethyl siloxanes incorporated into the same mixture used to make the Inamed gel.

The radioactivity in the blood peaked at 21 days after implantation.  After 30 days, only 0.06 percent of the radioactivity was released from the gel, even though the gel was not encapsulated in a shell.  So, these high molecular-weight materials are stable and don't migrate to an appreciable extent in this time frame.

The important toxicology issues are likely to be related to the lower molecular weight cyclic siloxanes, mainly D4 and D5.  The pharmacokinetics of these compounds have been studied by the Meeks group at Dow Corning and by Hoan-My Do Luu and Joe Hutter at FDA.

These materials diffuse out of the implants, but at slow rates.  The toxicology studies were designed to detect any potential toxicological effects, including those of the cyclic siloxanes or any other low molecular weight compounds that emerge from the prosthesis.

[Slide.]

Biocompatibility testing is listed on the slide:   cytotoxicity, irritation and sensitization, acute systemic toxicity, implantation testing, subchronic toxicity, hemolysis, and pyrogenicity.  The results of these tests were all negative.

[Slide.]

Because many of the breast implant recipients are likely to be of childbearing age, the reproductive safety of the gel implants is an important consideration.  In addition, the shell contains a diphenyl-substituted siloxane.

This was of some concern because compounds such as cis-2,6-diphenylhexamethylcyclotetrasiloxane have estrogenic activity.  This compound is a D4 in which two phenyl groups replace two methyl groups.  Although diphenyl-substituted cyclic siloxanes were not detected in the chemical analyses, we considered the reproductive toxicity testing to be another check on this potential hazard.

The shell was extensively tested in rat experiments, in a 1-generation study, and in a 2-generation study, which included histological observations of the ovaries and uterus of the F1 generation.

In the 1-generation study, the numbers of anomalies in the sham and treated control groups were not significantly different.  There were 418 pups produced by the sham group and 448 by the treated group.  There was no evidence of reproductive or teratogenic effects of the shell in the study.

The shell was also tested in a 2-generation study. Both the F1, as well as the F0 animals, used for mating were also implanted with pulverized shell material 5 weeks prior to mating, so two successive generations were exposed to the shell materials.

As shown on the slide, the mating indices and the fertility indices were the same in both the F0 and F1 generations.  There were no significant differences between the sham and treated groups.  There were no significant teratogenic effects.  The gel was tested in a similar protocol and the results were also negative.

Comprehensive reproductive and teratogenicity testing revealed no reproductive or teratogenic effects.

[Slide.]

Immunotoxicology Testing.  The immunotoxicology testing was conducted in mice.  The reporting included spleen and thymus weights, thymus histopathology, hematology, including differential counts and natural killer cell activity.  The measurement of splenic antibody forming cells and the mixed leucocyte reaction were used to assess complex immune functions requiring more than one component of the immune system.

Some results of the testing of the gel are shown on the screen.  The sham control mice received 1 ml of saline and the test mice were implanted with 1, 2, or 3 ml of gel.  As shown on the slide, the results in the presence of cyclophosphamide, the positive control, were all significantly different from the sham control.  The only significant finding in the highest gel exposure, that is the 3 ml implant, was the higher spleen weight.  This was shown to be an artifact in subsequent testing.

We conclude that there was no evidence of significant immunotoxicity.  Similar testing was performed on the shell and the results were also negative.

[Slide.]

Genotoxicity was tested using bacterial mutagenesis, mammalian cell forward mutation assay, chromosome aberration assay, and a mammalian cell transformation assay.

The results of the in vitro genotoxicity testing were all negative.

[Slide.]

For carcinogenicity testing, approximately 120 rats per group were implanted with Inamed prosthesis materials and a low density polyethylene control, and the animals were followed for two years.

The animals were examined for tumors and for other toxicological effects.  Hematology, blood chemistry, gross pathology, and histological pathology were all reported.  The major pathology seen was foreign body or solid state carcinogenesis.  This is regularly seen in rodents.

The results with the gel were that the treated animals had a longer time to tumor and a longer survival time than the polyethylene control animals.  There were twice as many tumors in the polyethylene control than in the gel animals.  In the implantation study mentioned previously, fibrous tissue developed between the particles of material put in, but the gel was only surrounded by fibrous tissue on the external surface.  This may be why fewer foreign body tumors were seen at the gel implantation sites.

The phenyl/phenyl shell and low density polyethylene control were pulverized and delivered in gelatin capsules.  In this testing, the overall survival of the shell-treated animals was shorter than the survival of the polyethylene control animals, which was shorter than the survival of the sham control animals.

The difference between the shell and the polyethylene control and sham were not significant when the implantation site tumors were removed from the analysis.  The implant site tumors were of mesenchymal origin, as is expected in foreign body tumors and rodents.  The data do not demonstrate a carcinogenic effect with any of the implant materials.

Studies currently in progress at Dow Corning indicate that D5, which may be present at very low levels, if at all, in the gel, causes uterine adenocarcinoma in rats.  A risk assessment indicates that the risk is negligible, but we will continue to follow this until the study is completed.

The application contains extensive and comprehensive toxicology testing on breast implants.  My review of these data indicate that the toxicological effects were not significantly different than the effects of the control materials.

Dr. Sahar Dawisha will now present the results of the clinical investigations.

Clinical Data Overview

DR. DAWISHA:  I was going to say good afternoon, but I guess good evening is in order here.  I am Dr. Dawisha.

[Slide.]

I reviewed the prospective clinical data contained in the PMA.  I am going to be providing an overview of the data.  This data constitutes the majority of the clinical safety and effectiveness information.

I also will be pointing out important clinical issues that I would like you to address as stated in the panel questions.

[Slide.]

This slide shows the prospective studies conducted by the sponsor which are included in the PMA, all of which are open label, multi-center, and which collected local complications.

I should point out that these studies did not address long term and general health effects thought to be related to breast implants, such as the risk of cancer, connective tissue disease, reproductive/teratogenic effects, and later effects on offspring.  For these issues, the sponsor is utilizing the published literature and animal data.

At the conclusion of FDA's presentation, you will be asked to address the adequacy of this information in the assessment of safety of the sponsor's product.

[Slide.]

Of the three studies, the core study constitutes the majority of the clinical safety and effectiveness prospective data, and my presentation will focus on this study.

Patients with all three indications were enrolled, and the study is intended for 10 years of follow-up.

It is the only study which included prospective screening for asymptomatic rupture via MRI in a subset of patients at years 1, 3, 5, 7, and 9 after implantation.

It is the only study that also collected quality of life information and a variety of connective tissue disease signs and symptoms.

The quality of life information was collected to supplement the patient satisfaction information, and the CTD signs and symptoms were collected in an effort to determine whether a referral to a rheumatologist was needed.

[Slide.]

The adjunct study was intended to make the implants available for patients with public health needs of primary breast reconstruction or revision of an existing implant due to medical or surgical reasons.

Local complications are to be prospectively collected at years 1, 3, and 5 post-op, and the 1 and the 3-year data are available from this study.

[Slide.]

The 1990 study originated in response to FDA's pending call for gel-filled breast PMAs.  Follow-up was yearly for local complications up to five years after implantation.

Because the majority of the patients enrolled in this study were augmentation, I will be referring only to the augmentation cohort and only to the 4 of 11 styles for which the sponsor is seeking approval.

[Slide.]

Now that I have provided a brief overview of the three prospective studies conducted by the sponsor, I will discuss the results of each study individually beginning with the core study.

[Slide.]

Before I begin with each indication separately for the core study, I would like to point out the median ages of the three cohorts in the core study:  34 years for augmentation, 50 years for reconstruction, and 44 years for revision.

These ages are consistent with those reported by plastic surgeons for U.S. women seeking breast implants, and as you can see, the median age for the augmentation patients is considerably younger, and is within childbearing years. This needs to be considered in the determination of safety for each indication as stated in the panel questions.

[Slide.]

I will now discuss the results for the core augmentation group.

[Slide.]

This slide summarizes the patient disposition for the core augmentation cohort.

There were 495 patients implanted, with one patient excluded due to an age protocol violation, and data are reported for 494 patients.

At two years, 90 percent of the expected patients he a follow-up visit.  At the time of database closure, 83 percent of the augmentation patients were available for a three-year visit and follow-up at three years was obtained for 81 percent of these patients.

There was no information available for the cause of death of the one death reported in this cohort.

For the 13 patients who were discontinued through three years due to removal without replacement of their study implants, most had changed to saline-filled implants or to gel-filled implants from other manufacturers.

Ten of these patients reported at least one complication, which is reported in the Kaplan-Meier risk rates on the next slide.

[Slide.]

This slide summarizes the cumulative Kaplan-Meier risk rates of first occurrence of selected complications through three years, on a by-patient basis, with the 95 percent confidence intervals shown in parentheses.

The complication with the highest rate reported was all severity levels of swelling, at 23.3 percent, which is not shown here because the time to resolution was within a week.  Reoperation, at 20.6 percent, was reported at the second highest rate, and capsular contracture, Baker Grade III or IV is the next highest rate, at 8.3 percent.

Note that the breast pain and nipple sensation change shown here only includes reports of moderate or greater severity, and scarring includes all severity levels.

I will next discuss reoperation and implant removal, followed by rupture.

[Slide.]

With respect to reoperations in the augmentation group, there were 248 additional procedures performed in 19 percent of the patients.

The two most commonly performed procedures were capsule related, which is due to capsular contracture, and implant removal with replacement, which I will discuss next.

[Slide.]

This slide summarizes the primary reasons for implant removal.  If more than one reason was reported, the following hierarchy was used to determine a primary reason: rupture, infection, capsular contracture, extrusion, necrosis, healing related, pain, unsatisfactory cosmetic outcome, and lastly, patient request.

There were a total of 60 implants removed through three years in 32 patients; 51 of these were replaced and 9 implants were removed without replacement.

Of the 60 implants removed, the majority, or 70 percent, were removed to treat a complication or to correct an unsatisfactory cosmetic outcome.

Thirty percent of the implants were removed or replaced due to patient choice.

Of these 18 implant removals for patient choice, 13 were for a size or shape change and 5 were due to patient concern, which the sponsor refers to in their application as media anxiety.

The specific complications which necessitated implant removal in these patients are shown with the white background.  Of the 60 implants removed, the most common reason was capsular contracture, 45 percent of the primary reasons for implant removal.

[Slide.]

The next two slides summarize the implant rupture information for the core augmentation cohort through three years.

For simplicity, this information is based on the date of database closure in order to be consistent with the Kaplan-Meier values reported in the application, and does not include rupture status which has been confirmed since that date.

166 patients, or 331 implants, were enrolled in the serial MRI screening cohort.  At the first screening of approximately one year after implantation, 139 of the expected patients, or 87 percent, underwent MRI screening for asymptomatic rupture.

At the second serial MRI screening at approximately three years after implantation, only 83 of the expected patients, or 64 percent, underwent MRI screening for asymptomatic rupture.

There were a total of 145 patients, or 289 implants, who had at least one serial MRI screening, which constitutes about 29 percent of the total core augmentation implants.

Of these 289 implants screened with MRI at least once, there were 3 implants reported with asymptomatic or silent rupture.

The sponsor reports a silent rupture rate of 1.2 percent through three years on a by-implant basis based on this MRI screening.

[Slide.]

The previous slide summarized asymptomatic ruptures.  This slide discusses the symptomatic ruptures and the overall rupture rate reported by the sponsor.

For those patients who were not included in the MRI screening subset--that is the "No MRI" screening group--there were 2 implant ruptures which were symptomatic, one due to pain in axillary nodules, and one due to capsular contracture.

Note that there were two additional potential symptomatic ruptures which the sponsor excluded from the rupture rate, which were shown in figure 1 and discussed in my review memorandum: one symptomatic rupture determined intact with mammography but not with MRI or explant, and one explant reported intact by the physician, but found to have a sharp-edge opening on microscopic evaluation.

For the overall by-implant rupture rate of 0.6 percent, the sponsor combined the 3 silent ruptured or asymptomatic ruptured implants from the previous slide with the 2 symptomatic ruptures shown here using the entire cohort for the denominator, for a total of 5 ruptures in 5 patients.

The overall rupture rate excludes potential silent ruptures in the "No MRI" group, which is about 71 percent of the core total implants.

Most of the ruptures identified in this cohort were asymptomatic.

[Slide.]

In addition to local complications associated with the implants, the sponsor also collected other safety information, such as self-reported reproductive problems, lactation problems, breast disease, breast malignancy, abnormal mammogram reports, connective tissue disease diagnoses, and connective tissue disease signs and symptoms.

Through three years, there was no increase in reports of reproductive and lactation problems.  There were 32 new patient reports of breast disease and 12 reports of abnormal mammograms, the results of which were mostly benign.

There was one new report of a breast malignancy, which occurred 27 months after breast implantation in a 39-year-old women who was diagnosed with invasive intraductal adenocarcinoma in situ.

Note that there is no comparison group of patients with similar characteristics followed for the same duration, so conclusions cannot be made from these data.

There was also one patient, a 46-year-old, reported as having rheumatoid arthritis 11 months after breast implantation.  Review of the diagnosing physician's notes indicates no mention of joint pain or swelling, a negative rheumatoid factor, and treatment with a non-steroidal anti-inflammatory agent, indicating that the patient does not likely have RA.

[Slide.]

Recall that the sponsor collected a variety of signs and symptoms, some of which could be related to connective tissue disease, in order to assist in referring patients to a rheumatologist if needed.

This table summarizes selected reports of before and through two years after breast implantation.

The starred entries contain general categories which encompass a variety of signs and symptoms, and the unstarred entries show specific signs and symptoms of interest.

As you can see, all signs and symptoms increased in frequency over time for the core augmentation group.

Without a comparison group of women with similar demographic characteristics and without implants followed for the same duration, it cannot be concluded that this increase is attributable to having breast implants.

[Slide.]

With respect to effectiveness, while the majority of patients who completed two years of follow-up and who responded to satisfaction questionnaires reported being satisfied with their implants, there were small declines in mean satisfaction over time for both general satisfaction questions and for satisfaction compared to pre-operative expectation of satisfaction.

With respect to general quality of life measures, such as the SF-36 and the MOS-20, there were small declines in mean values of all scales at two years compared to pre-op, however, the two-year values were still higher than normative values compared to age and gender-matched normal controls.

With respect to specific quality of live measures, mean values of some measures improved over time, while others declined, as shown on this slide.

[Slide.]

Now, a summary of the core reconstruction patients.

[Slide.]

Enrollment for the core reconstruction group began later than for core augmentation, such that only about 58 percent of the patients were eligible for a three-year visit at the time of database closure.

At both two and three years following breast implant reconstruction, over 90 percent of the expected patients had a follow-up visit.

The 7 patient deaths through three years were due to breast cancer recurrence or were accidental.

Of the 16 patients who discontinued due to removal of their implants, most had replacement with saline-filled implants or with gel-filled implants from other manufacturers.

All of the discontinued patients due to implant removal had at least one, and in most cases several complications, which is captured in the Kaplan-Meier risk rates shown on the next slide.

[Slide.]

The by-patient cumulative Kaplan-Meier risk rates of first occurrence are shown here through 3 years.  The complication with the highest risk rate is reoperation, at 45.9 percent, followed by implant removal or replacement, at 25.3 percent.

[Slide.]

With respect to reoperations, there were 242 additional surgical procedures performed in 42 percent of the patients.  The two most commonly performed reoperation procedures were capsule related, which is due to capsular contracture, and implant removal with replacement, which I will discuss on the next slide.

Note that scar revisions, wound repair, as well as pocket revisions were also common in this population.

[Slide.]

The primary reason for implant removal is shown here using the hierarchy previously discussed.  Of the 56 implants removed through 3 years, 51 implants were replaced and 5 implants were not replaced.

Most implants were removed to treat a complication or to correct an unsatisfactory cosmetic outcome--that's 92.9 percent--as opposed to patient choice for size or shape change, which was 7.1 percent.

Of the specific complications which necessitated implant removal, cosmetic problems, such as asymmetry, malposition, contour deformity, and wrinkling were the most common, at 46.4 percent, followed by capsular contracture, at 21.4 percent.

[Slide.]

With respect to ruptures for the core reconstruction group, 108 patients, or 184 implants, were enrolled in the serial MRI screening, and there is data for 170 implants at around 1 year after implantation.  This constitutes about 47 percent of the total core reconstruction implants.

Of these 170 implants screened at least once, there were 8 implants reported as ruptured.  The by-implant silent rupture rate reported by the sponsor is 4.7 percent based on these 8 ruptures.

Although this is reported as through 3 years, only 2 implants in 2 patients had reached their second MRI screening at 3 years, therefore, almost all of the data is based on the first screening at 1 year.

[Slide.]

With respect to symptomatic and overall ruptures, the sponsor reported 5 symptomatic ruptures in the patients who did not undergo MRI screening.  Three of these were actually asymptomatic and discovered through MRI, but were not included in the silent rupture rate.  One was due to asymmetry and one was due to flattening of the implant.

For the overall by-implant rupture rate of 4.2 percent, the sponsor combined the 8 silent ruptured implants, shown on the previous slide, with the 5 ruptures shown here, for a total of 13 ruptures in 12 patients.

This overall rate excludes potential asymptomatic ruptures in the "No MRI Screening" group, which is about 53 percent of the core reconstruction implants.  Most of the ruptures in this cohort were asymptomatic and identified through MRI.

[Slide.]

With respect to other complications through 3 years, there were no increased reports of reproductive or lactation problems.  With respect to breast disease, there were 5 new reports of breast malignancy, all of which were either recurrence or metastases of the original tumor.

There was 1 new report of a connective tissue disease, a 42-year-old patient was reported as having scleroderma four months following breast implantation.  The physician notes, however, associated with this diagnosis indicate that the patient has undifferentiated connective tissue disease.

Without a comparable group of patients, conclusions regarding the association of these reports with the implants is not possible.

[Slide.]

This table summarizes the results of selected self-reported connective tissue disease signs and symptoms, again with the starred entries representing general categories encompassing a variety of signs and symptoms.

As you can see, reports all increased from before to through 2 years after implantation.

Without a comparison group of patients with similar demographic and co-morbidities and without implants followed for the same duration, it cannot be concluded that this increase is due to the breast implants.

[Slide.]

With respect to effectiveness, most patients completing 2 years reported being satisfied, however, there were declines in overall patient satisfaction and in satisfaction compared to pre-operative expectation.

Mean values of general quality of life measures improved for this cohort, and the mean values of specific quality of life measures, some of which improved and some worsened over time.

[Slide.]

Now, for a summary of the core revision patients.

[Slide.]

At the time of data cutoff, about 91 percent of the core revision patients were eligible for a 3-year visit.

The follow-up rate is comparable to that of the core augmentation group.

Of the 10 patients having removal of their implant, most changed to saline implants or to implants from other manufacturers.  Five of these patients reported at least one complication, which is included in the Kaplan-Meier rates on the next slide.

[Slide.]

Like the core augmentation and core reconstruction groups, reoperation is the most frequently reported complication, at 33.4 percent, which is followed by implant removal or replacement, at 13.4 percent, and capsular contracture, Baker Grade 3 or 4, which is 9.8 percent.

[Slide.]

With respect to reoperation, there were 190 additional surgical procedures performed in 31 percent of the patients through 3 years.

Like the core augmentation and core reconstruction patients, the two most common reoperations were capsule related followed by implant removal with replacement.

[Slide.]

The primary reason for implant removal is shown here using the hierarchy discussed previously.

Of the 46 implants removed through 3 years, about two-thirds were removed to treat a complication, including unsatisfactory cosmetic results, and about a third were removed due to patient choice.

Of the specific complications necessitating implant removal, cosmetic problems, such as asymmetry, malposition, or ptosis constituted the most common primary reason, at 32.6 percent, followed by capsular contracture, at 15.2 percent, and followed by implant rupture, at 13 percent.

[Slide.]

With respect to MRI screening, there were 77 patients, or 148 implants, who were enrolled in the serial MRI screening cohort.

There is data for 138 implants with at least one MRI, which is 94 percent of the expected implants, and which constitutes about 38 percent of the total core revision implants.

Of these 138 implants screened at least once, there were 4 implants reported as ruptured.

The by-implant silent rupture rate of 2.9 percent reported by the sponsor through 3 years is based on these 4 ruptures.

Again, although this is reported as "through 3  years," most of the data is based on the first screening at 1 year, with only 2 implants in 1 patient having reached their second screening at 3 years.

[Slide.]

This slide summarizes the symptomatic ruptures and the overall rupture rate.

For those patients not included in the MRI screening subset, there were 4 implant ruptures which were symptomatic, 2 implants that were reported with nodules or bulges noted in the same patient, 1 due to a motor vehicle accident, and 1 due to softening of the implant.

For the overall by-implant rupture rate of 2.2 percent, the sponsor combined the 4 asymptomatic ruptures shown in the previous slide with the 4 symptomatic ruptures shown here, and used the entire cohort for the denominator, for a total of 8 implant ruptures in 7 patients.

This overall rupture rate excludes potential asymptomatic ruptures in the "No MRI" group, which constitutes about 68 percent of the core revision implants.

Half of the ruptures in this cohort were asymptomatic and identified only through MRI screening.

[Slide.]

Through 3 years, there were no increased reports of reproductive or lactation problems.  There were 13 new reports of breast disease, all of which were benign.

There was 1 new report of a connective tissue disease, a 50-year-old patient was diagnosed with fibromyalgia 11 months after implantation.  Review of the diagnosing physician's notes confirms this diagnosis.

[Slide.]

This table summarizes selected self-reported CTD signs and symptoms results.

All entries in this table represent general categories encompassing a variety of specific signs and symptoms.  As you can see, all reports through 2 years compared to before were increased.

[Slide.]

With respect to effectiveness, of the patients who completed 2 years of follow-up, most patients reported being satisfied with their implants, however, mean satisfaction levels declined over time for both general satisfaction and for satisfaction compared to pre-operative expectation.

Mean values of general quality of life measures, such as the SF-36 and the MOS-20, declined over time, as well.  However, the 2-year values were still higher than those reported for normative age and gender matched populations.

Of the specific quality of life measures, only the Body Esteem and Sexual Attractiveness subscale improved over time, while all other worsened.

[Slide.]

I would like to point out the limitations in comparing the data from the sponsor's gel-filled implants to their saline-filled implants.

First, the rupture rates cannot be compared because rupture of a saline-filled implant is immediately evident, while, as you can see from the data, most gel-filled implant ruptures are asymptomatic and not immediately evident.

The saline-filled breast implant data is a historical control group, with inherent differences in baseline demographics, in surgical characteristics, and in surgical practices, which is not addressed in the sponsor's comparison.

Finally, the sponsor did not provide confidence intervals for their comparison.

For the complications of reoperation, removal, and capsular contracture, which are less subject to interpretation than implant wrinkling, palpability, or breast pain, for example, the confidence intervals are overlapping for the saline-filled implants compared to the gel-filled implants, suggesting no significant differences.

[Slide.]

The adjunct study and 1990 studies are summarized in this slide.

The follow-up rate is poor in the adjunct study.

The complication rates reported in these studies are comparable to those in the core study.

However, the implant rupture rates cannot be compared because asymptomatic screening for silent rupture was not performed in these studies.

[Slide.]

In summary, reoperation, followed by implant removal or replacement, is the most frequently reported complication.

Capsular contracture reoperation procedures constituted the majority of these reoperations.

For implant removal, most implants were removed to treat a complication, including unsatisfactory cosmetic outcomes.

The frequencies of signs and symptoms related to connective tissue disease increased over time.

With respect to effectiveness, while patient satisfaction is high, it declines over time.

General quality of life measures improved only for the reconstruction cohort.

Of the specific quality of life measures, only the Body Esteem Scale, and specifically only the Sexual Attractiveness subscale consistently improved.

[Slide.]

Implant rupture, in particular asymptomatic rupture, is a key clinical issue regarding focal complications.

The overall rupture rate reported by the sponsor is under-ascertained because MRI screening for asymptomatic rupture was performed in only a subset of the patients and only in the core study, and most implant ruptures were asymptomatic.

As you can see, of the total 26 implants reported as ruptured in the core study, 15 implant ruptures, or 57.7 percent, were asymptomatic.

These asymptomatic ruptures are based on a third of the total core study implants enrolled and essentially includes only the first MRI screening at 1 year.

Because of the critical nature of implant rupture in the determination of safety, you will be asked to address the adequacy of the rupture rate information in the determination of safety of this PMA.

A final and important note is that almost all of the asymptomatic ruptures detected via MRI were intracapsular, that is, within the fibrous capsule.

[Slide.]

There is currently no method of determining whether and/or when an intracapsular rupture will become extracapsular, however, it is known that extracapsular silicone is associated with local and potentially systemic effects (to be discussed) and is difficult to remove, prolonging explant surgery time and complexity.

Because there are no current standardized guidelines for the frequency (or for the method) of screening for asymptomatic rupture, you will be asked to address screening methods and frequency for screening for asymptomatic rupture if you recommend approval.

You will also be asked to provide a recommendation for the necessity of removal of asymptomatic ruptured implants.

Thank you.  This concludes my presentation.

Dr. Irony will discuss the statistics.

Statistical Review

DR. IRONY:  Good evening.  I am Telba Irony, the statistical reviewer for the Inamed silicone-filled breast implant study.

In this presentation, I will comment on the statistical techniques that were used in the design and analyses of the study presented by the sponsor.  I will also point out to possible biases that should be taken into account when the results of the study are analyzed.

[Slide.]

The safety and effectiveness of the Inamed breast implants were evaluated through four clinical studies.  My statistical comments refer only to the core study, a prospective, multi-center, 10-year study initiated in 1999. It is the primary study to support this pre-market approval submission.

The reported follow-up times were 4 weeks, 6 months, 1, 2, and 3 years.

The results through 2 years post-implant are complete and the data reporting through 3 years has not been completed yet.

Safety data will continue to be collected at scheduled follow-up times through 10 years.

[Slide.]

The Core Study.  The core study was composed of three cohorts:

The augmentation group started with 494 patients and 83 percent of those patients had traversed the 3-year visit.  Excluding the patients that did not reach 3 years after implantation, the patients that had their implants removed, and the dead, 398 patients were expected at 3 years.  322 provided 3-year results, translating into 19 percent lost to follow-up.

The reconstruction group had initially 221 patients and 58 percent had traversed the 3-year visit.  By excluding the patients that did not reach the 3-year visit, the patients that had the implants removed, and the deaths, 116 were expected at 3 years.  Since 105 presented 3-year results, 9 percent were lost to follow-up.

Finally, the revision group started with 225 patients and 91 percent traversed the 3-year visit.  192 patients were expected at 3 years.  Since 160 patients had 3-year data, 17 percent were lost to follow-up.

[Slide.]

It is important to point out the descriptive nature of the study.

There were no claims, targets, or control groups.  No hypothesis tests were performed.  As a result, the safety and effectiveness endpoints for this product are summarized by the use of point estimates and confidence intervals.

Consequently, the adequacy of the sample size should be evaluated based on the width of the confidence intervals, that is, the precision, for the rates of the adverse events.

The acceptability of the resulting rates and their precision must be evaluated from a clinical perspective, keeping in mind the limitations of the study and the possible biases that could have influenced the results.

The resulting rates should help the decisionmakers to form an opinion about the safety and effectiveness of the implants and weigh their risks and benefits.

[Slide.]

Safety was addressed by estimating rates of complications and rates of secondary surgeries.

The following statistical techniques were employed:  Kaplan-Meier analyses, "prevalence", and "incidence."  The quotes mean that in this case, prevalence and incidence have a slightly different meaning than the usual prevalence and incidence.

I will address these three techniques and discuss their merits and flaws.

[Slide.]

Kaplan-Meier Analyses.  For each adverse event, including implant rupture, reoperation, implant replacement, and removal, a Kaplan-Meier analysis was conducted on the time to first occurrence of that event.

The result of such analysis, at each considered time point, is the estimated probability that a patient will experience that adverse event from the time of implant up to the considered time point.

The advantage of this technique is that it allows women who were not followed for the entire duration of the study to contribute information to the analysis for the time they were in the study.

However, it is very important to recall that this technique is based on the assumption that the censoring mechanism is independent of the occurrence of the adverse event.  In other words, it assumes that the reason for the patient not returning for the follow-up visit is not associated with the fact that she has or has not experienced the complication.

Panel members should evaluate how reasonable this assumption is in our context.  If patients are lost to follow-up because they experienced adverse events, then the actual rates should be larger than the estimated rates.

[Slide.]

Despite the fact that the follow-up times were in weeks, months, and years, the Kaplan-Meier rates were computed by considering days as units of time.

This makes the estimated rates much more accurate, alleviating the problem of interval censoring, since there is not a large time interval between the occurrence of the event and the recording of that event.  This procedure assumes that reporting the occurrence of the event is accurate to the day, which may not be possible for all events.

Correlations among adverse events were not investigated.  If adverse events are positively correlated, fewer patients will be affected although the ones affected will tend to experience more adverse events.

Whenever Kaplan-Meier curves are constructed for several adverse events, that may or may not be correlated, the problem of competing risks arises.  If a patient that experiences the first complication is removed from the pool of patients, that patient is not a candidate to experience another complication.

In this case, this did not happen because the patient that experienced an event was returned to the pool of candidates to experience another event.  The only exception was when the patient experienced implant removal with or without replacement.

[Slide.]

In this slide, the Kaplan-Meier rates for reoperation are displayed to exemplify the results of the procedure.

The intervals in the table are 95 percent confidence intervals for the chance that a patient will experience reoperation by 4 weeks, 6 months, 1, 2, and 3 years for each cohort, augmentation, reconstruction, and revision.

Since we are dealing with an estimation procedure, it is very important to take into account the whole confidence interval, and not only the midpoint.

So, how should we read this table?  For instance, read the contents of the last cell on the first row as follows:  For an augmentation patient, the 95 percent confidence interval for the chance of reoperation by 3 years goes from 16.8 percent to 24.4 percent.  The mean point is 20.6 percent.

The 3-year data collection is not complete yet.  A correction for the incomplete data was performed for the reconstruction cohort--see the red star on that table--because for that cohort, a large fraction of patients did not reach the 3-year window.  Note that for that case, the interval is wider.  For the other cohorts, the correction was not necessary.

[Slide.]

In addition to the Kaplan-Meier rates, the sponsor estimated the prevalence and the incidence of adverse events at each follow-up time point.

Here, the prevalence is the percentage of patients seen at a given follow-up visit who are experiencing a specific adverse event.  The incidence is the percentage of patients seen at a given follow-up visit who are experiencing a specific adverse event, not experienced at earlier visits.

What distinguishes the prevalence and incidence in this case from the usual prevalence and incidence is that the percentages are conditioned on the fact that the patients return to the follow-up visit.

These measures are extremely sensitive to biases generated by losses to follow-up.  Perhaps the patient did not come back to the follow-up visit because she has had a complication.  On the other hand, she may not have returned because she was doing fine and saw no reason to come back to the follow-up visit.

[Slide.]

As additional safety information, data on connective tissue disease signs and symptoms were collected and analyzed.

Patients completed an activity and lifestyle questionnaire prior to receiving their breast implants and 2 years after implantation.

The frequencies of patients reporting 8 categories of signs and symptoms of connective tissue disease before implantation was compared to the frequencies 2 years after implantation.

For all 3 cohorts, all 8 frequencies increased after implantation.  The only exception was the urinary symptom in the reconstruction group.  Its frequency remained the same before and after implantation.

It is very unlikely that all these increases happened by chance.

[Slide.]

In order to assess the statistical significance of the increases in the frequencies, the sponsor performed a Bonferroni correction that was too conservative for this case, especially if there are correlations among the signs and symptoms, which cannot be ruled out.

As a consequence, it was difficult to detect statistical significance with the study sample size.

[Slide.]

Despite the conservative statistical analysis, the increase in the frequency of some signs and symptoms was still statistically significant.

However, because there was no control group to be compared to the breast implant patients, the clinical interpretation of these results is problematic.

These were the main statistical issues in this submission and at this point I conclude my presentation.

Dr. Lori Brown will now present FDA's review of the surveillance data and literature.

Thank you for your attention.

Medical Device Surveillance and Literature Review

DR. BROWN:  Good evening.  I am Lori Brown and I will review medical device surveillance and published literature on breast implants.

Because silicone gel breast implants have been available for nearly 40 years and there is extensive medical literature on them, Inamed was asked to review and summarize the literature on this product as part of the PMA.

[Slide.]

Device manufacturers are required to report adverse events to CDRH.  This includes deaths, serious injury, or malfunction (that is a failure of the device to meet its performance specifications or otherwise perform as intended.

We also receive voluntary reports from medical providers, consumers, and others through the MedWatch program.

[Slide.]

There are three basic surveillance databases used at CDRH.

Currently, MDR and MedWatch reports are entered into MAUDE database.

Some reports from manufacturers are reported through the Alternative Summary Reporting system.  This system is for reporting events that are well known and numerous, such as rupture or capsular contracture for breast implants.

Instead of sending us 100 individual reports that say there was a breast implant rupture, the manufacturer sends one quarterly report saying there were 100 breast implant ruptures reported to them.

The third database, Device Experience Network was used prior to the development of MAUDE so has historical data.

I will show you overall reports for silicone gel breast implants but then restrict my remarks to reports from MAUDE.

[Slide.]

Just some precautions about surveillance systems in general.  Surveillance is not the equivalent of a clinical study.  Because of under-reporting, rates cannot be calculated.  Also, the number of individuals at risk may not be known.  The number of devices ever sold cannot be used as the denominator.

[Slide.]

The FDA does not check for the accuracy and completeness of reports and these reports do not establish a causal link between the device and the reported event.

[Slide.]

These reports are useful for providing a signal to regulators of potentially important adverse events and for providing an important signal to manufacturers to guide them in improving the design and manufacture of their product.

[Slide.]

This slide shows reports for silicone gel breast implants overall and for Inamed silicone gel implants.  A total of 134,000 reports were sent to the FDA between 1984 and mid-2003.  Shown in the righthand column are adverse event reports for Inamed, previously McGhan Medical.  By mid-2003, there were 14,414 reports for Inamed silicone gel breast implants.

[Slide.]

Medical device reports in the MAUDE database are coded with device problems and patient problems.  This slide shows the three most commonly reported device problems for Inamed implants.  Each report can have multiple adverse events coded.

The most common problem reported was explantation, a close second was implant rupture, and the third most commonly reported problem was migration.  Some of these reports were for implant migration and others were for gel migration.  Data from the DEN database and summary reporting databases are similar in that implant rupture and explantation are the most frequently reported adverse events.

[Slide.]

This slide shows the most commonly reported patient problems for Inamed breast implants.  Reports from the other two CDRH databases are generally consistent with this.

[Slide.]

In addition to reviewing MAUDE reports overall, we perform analyses to examine issues of interest in greater depth.  These next two slides show the results of such analyses.

The first MAUDE analysis is on breast implant rupture during mammography.  We performed this analysis in response to a patient complaint in November 2002.  After identifying 17 reported cases of breast implant rupture during mammography from the literature, we searched the MAUDE database for reports of such incidents.

This analysis includes silicone gel and saline breast implants and is not specific to Inamed.  Between 1992 and 2002, FDA received 33 reports describing breast implant rupture during mammography.

In these 33 reports, 19 were for silicone gel implants, 11 for saline, and in three reports, implant type was not stated.  There were 8 additional reports that implied but did not explicitly state that the rupture was due to mammography.  All 8 of these implants were gel.

[Slide.]

We also did an analysis of MAUDE data on breast implants and reproductive issues.  We performed this analysis because there is only limited information on this issue in the published literature.

Again, this analysis includes silicone gel and saline implants and is not specific to Inamed products. Overall, we found 130 reports that reported injury or illness in mothers or their children that were attributed by the reporter to breast implants.

The majority of these reports simply stated that the children were ill, but provided no details.  In 23 reports, some detail was provided about the child's health problem and these included gastrointestinal problems, allergies, pulmonary respiratory problems, connective tissue disease,  and a variety of other problems.

There were also 5 reports in which the reporter    attributed a child's birth defect to mother's implants and 9 reports describing difficulty nursing because of breast implants.

[Slide.]

Inamed performed a review of the literature on breast implants as part of this PMA.  We reviewed and supplemented their review.  Please note that the literature does not refer specifically to the Inamed products under review and in some cases includes both silicone gel and saline breast implants.

Also, note that there are numerous published papers that have not been reviewed here.  Obviously, for the purposes of this presentation, positive findings will be highlighted.  This presentation does not include a critique of methods from individual studies or a listing of their strengths and weaknesses or we would be here all night.

[Slide.]

There are case reports describing children of women with implants with scleroderma-like esophageal disease and children with autoantibodies and clinical rheumatic complaints.  A Swedish cohort study of 5,874 children of women with breast implants indicates that these children were not at an excess risk for esophageal disorder, rheumatic disease, cancer, congenital malformation overall, or perinatal death compared to 13,999 children of women with breast reduction surgery.

The results were similar for offspring of Danish women with breast implants compared to breast reduction and plastic surgery controls.  However, in the Danish study, there was a twofold increase in risk for esophageal disorders and for malformations involving digestive organs in children born after maternal implant surgery.

There were similar increases in children of women in the comparison group undergoing other plastic surgery or breast reduction surgery leading the authors to conclude that the association is with some aspect of the surgery rather than with breast implants.

There are two issues in breast feeding - first, whether breast milk is safe, and second, whether women are able to nurse after implantation.  A single study indicated that there are not increased levels of silicon in breast milk from women with implants compared to control women, and that infant formula and cow's milk had higher levels of silicon, not silicone.  Other studies indicate that women with breast implants are less likely to attempt nursing an infant and that women with breast implants or other types of breast surgeries are less likely to successfully breast feed.

[Slide.]

Numerous studies have evaluated whether connective tissue disease is more likely to occur in women with breast implants.  Using a statistical technique to combine the results of these studies, Janowski et al. found that while the adjusted relative risk of all connective tissue disease combined from 14 studies was slightly elevated in women with implants, that analysis for individual specific connective tissue disease, including arthritis, lupus, scleroderma, dermatomyositis, and polymyositis were not increased significantly.

There was a slight increase in the summary adjusted relative risk for Sjogren's syndrome.  When they excluded the results from one large study that relied on self-reported connective tissue disease, the results were no longer statistically significant.

Likewise, based on their analysis of the literature, the Institute of Medicine concluded that there was not evidence to support an elevated risk for these diseases in women with breast implants.

[Slide.]

While numerous studies have consistently not found an increase in connective tissue disease, there has been concern over an atypical disease syndrome.  The Institute of Medicine forcefully refuted conjecture over a new atypical syndrome associated with silicone breast implants.  Some fibromyalgia experts have proposed that atypical syndrome reported in women with implants may be fibromyalgia.

However, studies evaluating fibromyalgia in women with implants have had mixed results.  This slide shows three studies in which fibromyalgia was not statistically increased in women with implants, albeit in two of these studies, fibromyalgia was increased but not statistically significant.

[Slide.]

The study by Kjoller, et al. from Denmark actually reports on two cohorts of women with implants,  one a public hospital cohort and the other a private clinic cohort.  In each of these cohorts, unspecified rheumatism including fibromyalgia were significantly increased compared to sex specific national hospital discharge rates.

The study comparison group, women undergoing breast reduction or other plastic surgery, also had a significantly increased risk for fibromyalgia leading the authors to conclude that it was the plastic surgery rather than the breast implants associated with fibromyalgia.

The second study on this slide is one that I performed here at FDA and shows the relative risk for fibromyalgia in women with extracapsular silicone detected by MRI compared to other women with silicone gel breast implants.  There is a near three-fold elevation in self-reported fibromyalgia in women with extracapsular gel.

[Slide.]

Breast cancer risk for women with breast implants has been examined in numerous studies and these studies do not support an association.  This slide shows several studies on cancer other than breast.

Highlighted in red are cancer sites for which there was a statistically significant increase in more than one study.  Blue numbers are for sites statistically significant in one study.  NS is evaluated but not statistically significant and empty boxes indicate that this was not evaluated in the study.  Sites which were elevated in any analysis in more than two studies are lung, vulva, and all sites combined.

[Slide.]

As already discussed, adverse event reported to the FDA and case reports from the literature indicate that breast implants may be ruptured by compression during mammography.  Although breast implants are tested to withstand the compressive force of mammography, there are changes to the implant in vivo which may contribute to vulnerability to rupture.  It is not known how often implant rupture during mammography occurs since there are only case reports.

Silicone gel is radiodense and obscures part of the breast during mammography.  Implants also decrease the compressibility of the breast.  Special techniques are required to image breast tissues when implants are present and these techniques require additional views, so increase exposure to radiation, time required for examination, and cost.  In some women, capsular contracture may preclude the use of mammography and adjunct methods of cancer detection are necessary.

Several studies have reported delay in breast cancer detection in women with implants.  One study reported that although there was a delay in detection, breast cancer mortality was not increased in women with implants.

Another study found no difference in tumor size, axillary lymph node involvement, or histopathology in women with implants, but this study also reported that treatment options for women with implants did differ with mastectomy being preferred over breast preserving treatments.

[Slide.]

Two cohort studies have found a slight elevation in risk for neurologic disease in general among women with cosmetic breast implants.  In these studies, there were similar findings of increased risk for neurologic disease overall in their comparison group of women undergoing breast reduction surgery.

They concluded that neurologic disease is unlikely to be causally related to the implants, but rather appears to be associated with characteristics of women undergoing aesthetic breast surgery.  They did not delineate these characteristics.

[Slide.]

Three studies evaluated mortality from various causes in women with cosmetic breast augmentation.  This slide shows the standardized mortality ratios and 95 percent confidence limits of the implant population versus the comparison population.

In the first study shown here, death from all causes was significantly reduced in women with breast implants compared to the general U.S. population.  This reduction was attributed to the generally healthier status of patients seeking plastic surgery.

However, there was a significantly increased mortality rate for both suicide and brain malignancy in women with implants.  Similar studies performed in Sweden and Finland found either slightly increased risk for mortality from all causes or no difference.

However, both studies also found an increase in death by suicide in women with implants.  The reason for this increased mortality is not clear.  Pukkala, et al. have suggested that this may be due to underlying psychopathology in women seeking cosmetic augmentation.

Studies seeking to characterize women desiring cosmetic augmentation or after cosmetic augmentation have been limited and seriously flawed leaving room for conjecture on this point.

[Slide.]

In 1999, prior to the submission of this PMA, based on their review of the literature, the Institute of Medicine concluded that reoperations and local and perioperatve complications are frequent enough to be a cause for concern and to justify the conclusion that they are the primary safety issue with breast implants.  In several reports, the rate for reoperations is around 33 percent.  This is likely an underestimate based on study methodology.

Also,  I made a mistake by putting Gutowski, et al., under here.  This is saline breast implants.

[Slide.]

The Institute of Medicine reported that these risks for reoperation and local complications continue to accumulate over the lifetime of the implant and that quantitative data on this point are lacking for modern implants and deficient historically.

The IOM also said that information concerning the relatively high frequency of local complications and reoperations is an essential element of adequate informed consent for women undergoing breast implantation.

[Slide.]

Numerous studies have reported high rupture rates based on implant status at explantation.  These studies were thought to be biased because they represent women who were having implants removed because of problems.  A meta-analysis of 35 of these studies by Marotta et al. indicated that the failure rate was 30 percent at 5 years, 50 percent at 10 years, and 70 percent at 17 years.

Two studies of breast implant rupture as assessed by MRI imaging are shown on this slide.  The first bullet shows results of a study which I conducted here at FDA.  In this study, the median age of implants was 16.7 years.  Fifty-five percent of implants were evaluated as ruptured and 22 percent of ruptures were extracapsular.

In a similar study performed in Denmark, the median age of implants was 10 years.  Twenty-six percent of implants were ruptured in this study.  The differences in the two studies may be due to implant age, differences in rupture rate for different manufacturers, which would have differed between these two studies, and problems with the MR imaging at one of the facilities in the Danish study that imaged over one third of the women.

[Slide.]

When breast implants rupture, silicone gel may migrate out of the scar capsule and through tissues.  Across a series of studies, gel migration was observed for 11 to 23 percent of ruptured implants.  There have also been cases of gel migration from intact implants.

These may be cases of silicone oil migration from intact implants or possibly gel migration from minimally ruptured implants.  The frequency and severity of distant gel migration are unknown.

Migration may result in gel or oil being found in lymph nodes, intraductal extensions of gel in the breast, granuloma formation, transcutaneous leakage of gel, ulceration, tissue destruction, and scarring.  Silicone in the tissues for these studies are typically confirmed by microscopic examination of tissues, granulomatous response, or by imaging.

In summary, there is extensive literature on breast implants.  Not all of this information is specific to silicone gel implants or to Inamed products, and there are conflicting results from some studies.  The literature offers a rich source of information on breast implants in general and should be considered.

Commander Allen will now discuss the proposed post-approval study plan and labeling.

Thank you.

Proposed Post-Approval Study and Labeling Overview

[Slide.]

CDR ALLEN:  I will now present an overview of the proposed post-approval study plan and labeling.  We have panel questions regarding these two topics.

This information and the panel questions should be viewed in the context that FDA has not made a decision on the approvability of the device at this time.

[Slide.]

The core study protocol, as well as informed consent, currently requires yearly follow-up with a physician through 10 years, including MRI screening for silent rupture at 1, 3, 5, 7, and 9 years.  If their device is approved, Inamed is proposing a two-phase post-approval study.

Phase I involves patients in the core study continuing with their evaluations as per the current protocol through their 5-year evaluation timepoint.

Phase II involves a mail-in survey completed by the patient each year from the 6 through 10-year timepoint.  This survey focuses on the safety outcomes of breast pain, capsular contracture, implant rupture, and reoperation, as well as patient satisfaction. No other endpoints, including MRI assessments, would be captured.

As I mentioned, FDA has a question for you regarding the proposed post-approval plan.

[Slide.]

With regard to labeling, the Directions for Use would be packaged with the device and be posted on Inamed's web site.

The Patient Brochure would not be distributed as part of the packaged device.  Instead, it would be distributed to surgeons who could then provide it directly to their patients during their consultation visits.  It would also be posted on Inamed's and FDA's web site.

Inamed is also proposing to conduct a focus group study of the Patient Brochure to improve its quality.

[Slide.]

This concludes FDA's presentation.  Thank you for your attention.  I will now turn it over to the panel.

DR. WHALEN:  If the FDA review team could assemble at the table of inquisition here to entertain panel questions, please.

Who would like to begin?

DR. ANDERSON:  We have heard a mixture of testimony.  This morning we heard anecdotal cases from individuals that said they had problems related to connective tissue disorders, and they attributed them in their anecdotal cases to rupture.

Then, we heard testimony from the company that provided us data about the overall group of people who have had implants, but did not give us specific data about the outcome of the patients that had rupture.

What I am hearing, and I would like confirmation or tell me if I am missing this, what I am hearing from your review, in your review of their data, in slide 99, Dr. Brown, your study suggests that--I am sorry, this is literature data--that, in fact, there could be heightened rates of fibromyalgia among ruptured patients.

I also heard, in a slide earlier, that there were eight factors that people self-reported in terms of connective tissue disease/disorder symptoms, and actually, all eight were higher post-implant than pre-implant, and that that seemed unlikely to be related to chance alone.

So, is it true, then, that we do not have resolution about what happens to these ruptured patients in terms of the connective tissue disorders?

DR. BROWN:  As I pointed out, the studies about fibromyalgia and breast implants have had mixed results.  In the study that I performed at FDA, it was not women with ruptured breast implants who were more likely to report fibromyalgia, but rather women who had extracapsular silicone gel as determined by MRI.

As far as signs and symptoms, I will let Dr. Dawisha comment on that.

DR. DAWISHA:  I guess the answer is yes, that this issue is not resolved at this time.

DR. LI:  I have a couple of questions for Ms. Allen.  The first question I have is on the retrieval study, they listed 158 of the 339 retrieved devices as being functional.  As I go back and read the fine print, they said that 17 percent of all the devices that reported ruptured, were also classified as functional, and 23 percent of all the devices that reported non-ruptured were rated nonfunctional.

So, given that, what does functional mean, because I would think that a ruptured implant, regardless of anything else, is really a nonfunctional implant.

CDR ALLEN:  I will give it my best shot.  Maybe Inamed will have to fill in, but my understanding is that the ruptured designation was made by the physician at the time of explantation, and what could have happened was that for the example of a physician noting it as ruptured, it goes back, and it is deemed functional, is that possibly they removed it and it was nicked after the fact and they found--excuse me, I said that wrong.  It was deemed ruptured by the physician, but the laboratory found no opening, so actually I can't remember Inamed actually addressing that.

The opposite way is that they believe that during the explantation approach, that there might have been an opening made at that time.

DR. LI:  I was just a bit confused over this, I am not quite exactly sure, maybe Dr. Spear can answer why you would have a situation where the surgeon would report an implant ruptured at explantation, yet, when it goes back to the company, the company finds it not ruptured, and conversely, how were implants turned in as being labeled as non-ruptured, and yet the company will find 23 percent of them as being ruptured.

CDR ALLEN:  I would let Inamed speak.

DR. WHALEN:  From the point of view of process, if you feel that really, Inamed has to clarify something very specifically right now, we will entertain it, otherwise, they have a summation period tomorrow morning where they can comment.

DR. LI:  Maybe my question to Ms. Allen is are you satisfied with their analysis and discussion of things that are ruptured and things that aren't, and then their classification is functional?

CDR ALLEN:  That's a good question.  Inamed performed the study as per the protocol, and the protocol is something we have looked at and we accepted.  I would say that it's discouraging, though, that the goal of the retrieval study, which was to determine the mode of failure, to assess what is going on, we didn't really get there, so I don't know if that is a fault of the design of the study itself, and I don't know where to take that from there.

At this point, we don't have any questions posed specific to the retrieval study they have.  I think it's a bigger question for us to look at.

DR. LI:  This question has been bothering me all day.  This has to do with the crosslink density of the implant and for surgical colleagues, why I think the crosslink density is important besides determining the physical properties of the implant, it would also control things like the ability for molecular species to flow through it, its diffusion characteristics.

It will also, the crosslink density will also be related to the amount and type of things you can extract out of it, but I am a little concerned that the only test for the crosslink density appears to be assessing the physical properties of the implant, which don't address the crosslink density.

In particular, the thing that confuses me is that I believe it's a three-layer structure, and if I did my reading right, the middle layer is supposed to be a barrier layer presumably to reduce diffusion of low molecular weight materials, and it is supposed to have been more highly crosslinked.

Yet, the penetrometer data and the mechanical testing data show that the mechanical properties of the middle layer is exactly the same as the base layers, so on one hand, you are telling me the middle layer has a crosslink density and you are going to assess the crosslink density by mechanical properties, but yet your mechanical property test shows that between two layers, you should have different crosslink densities, they are the same.

DR. AREPALLI:  As a matter of fact, I don't think even you have the barrier layer, you can completely eliminate the diffusion.  So, regardless what kind of barrier layer you put in, there will be some diffusion.

I don't have the details here, but as far as I remember, penetrometer is used for the gel, the crosslinking was used for the shell.

DR. LI:  Excuse me.  My question is are you comfortable with the use of mechanical properties alone to assess crosslink density?

DR. AREPALLI:  They use the mechanical properties to make sure that the degree of crosslinking is uniform across the lots, like elongation and break force, if I remember correctly.

DR. LI:  I understand that's the goal, but I personally don't think that is a good connection between mechanical properties and crosslink density.  So, maybe my follow-up question is, what correlation or association can you show me that penetrometer data and/or mechanical property data is actually linked to crosslink density?

DR. AREPALLI:  All I can say is the mechanical properties, because they are the same across the lot, I would say that the degree of crosslinking is uniform across the lots, but if you ask me the proof for it, I don't have any evidence for that.

DR. WHALEN:  Dr. Witten.

DR. WITTEN:  I just wanted to say that some of these questions may really be questions for the sponsor, like whether or not the assessment, there should be another assessment of middle layer, that kind of thing.  We are probably best answering questions about our presentation, what is the presentations.

DR. WHALEN:  To follow that up, are you postulating that we are going to work into the schedule, another question session with the sponsor since we already did that?

DR. WITTEN:  That is up to the panel during the panel discussion, if you have some questions that you need addressed.

DR. EHMSEN:  Mr. Chairman, we certainly would appreciate an opportunity to comment on some of the comments that were made by the FDA.

DR. WHALEN:  Beyond your summation tomorrow that is already in the schedule?

DR. EHMSEN:  Okay.

DR. WHALEN:  No, I am asking.  You need some time beyond your summation period?

DR. EHMSEN:  Yes, please.

DR. WHALEN:  We will take that under advisement and we will continue with as many questions as we can through to FDA since we are probably going to be going until midnight the way it is with the schedule as it is written, if we want to try to get through tomorrow before people have their travel commitments, but we will take that under advisement and deal with it during a break period.

Dr. Miller.

DR. MILLER:  I wonder if the FDA could help us interpret these increased frequencies of these collagen and vascular disorder symptoms because I mean I have, after sitting here today, an increased frequency of all these symptoms.

[Laughter.]

DR. MILLER:  So, it is hard for me to know how to interpret this piece of data, especially in light of all the other data that has been presented today which sort of doesn't leave the question wide open as I think was just said.

DR. DAWISHA:  I think we would all agree that without a control group, we are all left with that question of how do you interpret this data.  The data are the data.

DR. WHALEN:  Dr. McGrath.

DR. McGRATH:  I have a question for Dr. Brown first.

Under MAUDE analysis where you are looking at the breast implant rupture during mammography, do you know how old those implants were that you were picking up during the last decade, that ruptured during mammograms, do you have any idea of what their age was at the time they ruptured?

The second question is do you have any MAUDE reports of rupture following blunt trauma?

DR. BROWN:  In response to the first question about the age, I don't recall the exact age of these, and it is not always provided in the report, but as I recall, it's around 16 years.

DR. McGRATH:  So, they are quite old.

DR. BROWN:  Yes, and that is not on all of the implants that were in there, that is only given the ones where they did provide an age.

As far as blunt trauma, yes, there are reports of blunt trauma, as well as there are reports of blunt trauma in the literature as a source of rupture, as well.

DR. McGRATH:  Coming into your MAUDE analysis, you are getting something.

DR. BROWN:  I have seen them over the years, yes.

DR. McGRATH:  But not numerous.

DR. BROWN:  I have never enumerated them, so I don't know how many there are, but I have seen them.

DR. WHALEN:  Other questions?  Brent.

DR. BLUMENSTEIN:  I am particularly struck by slide 45 using the FDA presentation.  The CTD, I forget what that stands for, connective tissue disease, so what we have here is showing that pre-implant to post-implant, at 2 years, for 386 patients, these are the same patients?  In other words, this is a paired analysis?

DR. DAWISHA:  Yes.

DR. BLUMENSTEIN:  In every one of those, if I have scanned it correctly, there is an increase.

DR. DAWISHA:  That's right.

DR. BLUMENSTEIN:  This seems rather alarming to me, I mean because you are not censoring, you don't have a criterion set up, there is no regression to the mean issues going on here.  This is just simply at two points in time, you have assessed these two sets of features.

DR. DAWISHA:  That's right.

DR. BLUMENSTEIN:  Has there been any attempt to make a composite outcome or some other conclusion like that, I mean some other endpoint point like to test the significance of this?

DR. DAWISHA:  What do you mean a composite outcome?

DR. BLUMENSTEIN:  Maybe a multivariate outcome, maybe a score that combines outcomes.

DR. DAWISHA:  You mean to combine all the data?

DR. BLUMENSTEIN:  I am interested if we have a statistical null hypothesis is there is no change, I am interested in what the statistical significance of that hypothesis is.

DR. IRONY:  As I pointed out in my slide, some of these changes were statistically significant.  There was a Bonferroni correction and, you know, the requirements to detect statistical significance were very conservative in my opinion.  However, there was no control group to compare those, so one might say that the same increase in frequencies will happen in the general population after two years.

That is why we have some problems in interpreting the changes clinically.

DR. WITTEN:  Excuse me.

DR. WHALEN:  Dr. Witten.  Go ahead, please.

DR. WITTEN:  I just want to mention that as Dr. Dawisha had said in her discussion, that the purpose of this wasn't actually as a study per se to examine a specific hypothesis about connective tissue disease.  It was to collect the signs and symptoms to assist the physician in referring a patient to a rheumatologist if needed.

So, that may explain the limited amount of analysis that was done on this, or correct me if I am wrong, Dr. Dawisha.

DR. DAWISHA:  No, that's right.

DR. BOULWARE:  Dr. Dawisha, can you clarify for me what the first row muscle is versus the fourth one, muscle pain and joint versus joint pain?

DR. DAWISHA:  The muscle is a general category that incorporates a variety of signs and symptoms.  This is the way the sponsor had provided the data.  Joint pain is one specific question of joint pain, and muscle pain is one of the questions of muscle pain.  Muscle pain came from the muscle category and joint pain came from the joint category.

DR. WHALEN:  For everybody using mikes, I am pretty sure these are directional mikes, and if you come from the side, it really doesn't work very well, so try to help out the people who are trying to know what we are saying.

DR. LEITCH:  Dr. Brown, your study about the extracapsular silicone, I was wondering how those patients were identified as having extracapsular silicone, at what time frame in which they were, you know, years were they diagnosed, and it is self-report of fibromyalgia diagnosis, it wasn't evaluated any other way?

DR. BROWN:  Right.  This was a study of breast implant rupture that was based on a cohort that was put together by the National Cancer Institute, and we used one of the sites at that study and had MRI performed on women who remained in that area, and there were 344 women who had MRI examination.

So, the extracapsular silicone was based on the evaluation by three radiologists who then voted as to whether the implant was ruptured, indeterminate, or intact, and they also voted on whether or not there was extracapsular silicone.

Before they had the MRI examination, we did a survey with these women during which they were invited to the MR examination.  At that time, they were asked whether or not a doctor had ever diagnosed them with fibromyalgia, so this is a self-report of doctor-diagnosed fibromyalgia.

DR. LEITCH:  And no other documentation?

DR. BROWN:  No, that is the extent of the documentation, it is a self-report, but we did ask them whether a physician had diagnosed them with fibromyalgia.

DR. CHOTI:  I was intrigued by the way you presented the data of the increased rupture rate in the reconstruction compared to the augmentation, something like 4 times the rate, I think, if I am not mistaken.

In addition, the silent rupture was significantly higher, and it was clearly underestimated given that most MRIs were done only at one year.  So, it is a little bit of concern what a high potential silent rupture rate is present in the reconstruction.

First of all, perhaps a little speculation as to why the rupture rate is higher in the reconstruction.  We heard from the sponsor perhaps it's surgical manipulation folds.  It is not something, at least to me, why it would be so much higher in the reconstruction compared to the augmentation.

Another question of the group, perhaps to comment on, is there any reason to suspect that the extracapsular spread be any different in the reconstruction compared to the augmentation.  They both, I am sure, have capsule formation.

Dr. Brown, perhaps in your study, did you look at reconstruction versus augmentation?

DR. BROWN:  The cohort that we used was purely an augmentation cohort because one of the outcomes for this study at the National Cancer Institute was breast cancer, so therefore, they only were enrolling women in the study who had augmentation mammoplasty.

As far as there being a difference between extracapsular in reconstruction versus augmentation, I don't know the answer to that.  I don't think that I have ever seen it broken down that way.

DR. CHOTI:  But I am concerned about such a potential high leak rate in the reconstruction group.

One other question, Dr. Brown, regarding your review of mammography.  We talked about potential risks, missing a cancer, but one issue that came up is did you find any evidence in the literature to suggest that women with implants for augmentation are less compliant with mammography in a fear, whether it is valid or not, women's fear that they may rupture their implants during mammography and, thus, not comply with cancer screening?

DR. BROWN:  I have not seen such a study in the literature, however, in our analysis of the MAUDE data, there were several reports of women who feared rupture during mammography, and they also feared the pain during mammography based on capsular contracture or breast pain, which, of course, women without breast implants may fear also.

They also feared--we are talking now about fear, so they are reporting fear--some women also feared having breast cancer missed because of their breast implants.  So, the answer is yes, in the MAUDE reports, we did see fear of getting mammography, fear of pain during mammography, and fear of not being diagnosed with cancer in a timely fashion.

DR. CHOTI:  But if they are afraid of increased cancer rate, then, perhaps their compliance may be greater.

DR. BROWN:  I do not know.  Again, as I pointed out, MAUDE and data from surveillance system are not the same as a study, a clinical study.  There is no denominator, so I do not know whether that is actually going on in this population.

DR. WHALEN:  Dr. Conant.

DR. CONANT:  Back to the rupture issue, I am interested in what you feel about this underestimation of rupture potentially because of the small population screened with MR for silent ruptures.

I did some quick math.  I think the reconstruction cohort was particularly interesting because 50 percent of it was actually screened with MR, which may reflect the higher number of ruptures detected.

But I am wondering, you know, when we look at that rate of rupture, how do you think that compares to your literature search if you take into the fact that many of them are actually screened?

DR. BROWN:  Well, as Dr. Dawisha pointed out in each of the slides where she showed you the MR, the results of the MR study, there were up to two-thirds of the women from each cohort who had not been imaged, so you can do a rough calculation to guesstimate what may have happened if those women had, in fact, been imaged.

DR. CONANT:  It is sort of interesting to look at those numbers and compare the rupture rate or estimated rupture rate.  I also believe that there is documentation here about the group of age for mammography that showed a fairly low compliance with this in the study cohort.

I am wondering if someone had a comment on that.

DR. WHALEN:  Apparently not.

Other questions?

DR. MANNO:  I would just like to comment that on the control that was mentioned or the lack thereof, I think just calling it "control" is a problem.  I think they are going to have to be age matched especially if you are going to go with all three groups or at least the augmentation to the cancer group that had the implants.  I think it should be implied.

DR. WHALEN:  Other questions of the FDA?  Yes.

MS. GILBERT:  With respect to the connective tissue disease summary on the core augmentation, I wanted to come to a question of if a patient reported muscle pain, would they also be reporting that they had a problem in the muscle and maybe morning stiffness, so would those get captured in the same complaint where they get captured two or three places?

DR. DAWISHA:  Yes, that's true.  I mean there were several questions, about 50 specific signs and symptoms that were asked, so there could be multiple reporting.

MS. GILBERT:  So, when there is concern that seven indicators are going in the same direction, it may not be seven independent things going in the same direction, is that correct?

DR. DAWISHA:  That is true although the categories are exclusive of each other.  So, each of these major categories have independent questions that are exclusive of the other categories, but some of the symptoms could be related.

MS. GILBERT:  It is more that I am asking if I had muscle pain, for example, would I answer the question, do you have muscle pain as yes, then, muscle, yes, and then morning stiffness, yes.

DR. DAWISHA:  You might.

MS. GILBERT:  Okay.  Thank you.

DR. WHALEN:  Dr. McGrath.

DR. McGRATH:  One other thing, breast pain.  This has bothered me all day that we don't talk about this.  I realize we don't have control groups, and so forth, but as you looked at this, was there any prior question ever about whether patients had breast pain before the implants were put there, and then also did you notice any changes in the number of people reporting breast pain over time?

DR. DAWISHA:  Breast pain, you are talking about the breast pain in the Kaplan-Meier risk grades after?  Those complications were only assessed after implantation.

What I did look at was the sponsor asserts or included breast pain in a general surgical category of complications.  What we were able to get was the median duration of the breast pain, which was about a month, but all those questions that are in the Kaplan-Meier results were only asked after implantation.

DR. WHALEN:  Dr. Li.

DR. LI:  Can you help the non-radiologist here?  For using an MRI to detect silent rupture, is there any data that shows how accurate that test is, you know, the percentage of false positives, false negatives, or how big the rupture has to be before you can see it?

DR. DAWISHA:  In this study, the ruptures that were identified via MRI were considered the classic signs of rupture, like the Linguini sign or air within the implant, there are specific signs that you look for on MRI which indicate a rupture.

I believe the sensitivity and specificity of MRI is pretty high, 98, 96 percent.

DR. LI:  Thank you.

DR. WHALEN:  I would say that linguini is an effective segue to our delayed dinner hour.  We will break for 45 minutes and reconvene with open panel discussion at 8:20 p.m.

[Whereupon, at 7:40 p.m., the proceedings were recessed, to be resumed at 8:20 p.m.]

E V E N I N G   P R O C E E D I N G S

[9:00 p.m.]

Open Panel Discussion

DR. WHALEN:  We are about to engage the open panel discussion of the PMA.  Prior to doing that, Dr. Witten would like to make a brief comment upon the definitions of safety and effectiveness.

DR. WITTEN:  Yes, I just passed out to the panel, I think someone is going to put up an overhead, that gives you the formal definition in our CFR of safety and effectiveness.

I will read it off while it is being projected.  I just provided these to the panel, like I said.

[Slide.]

The definition of safety is:  "There is a reasonable assurance that a device is safe when it can be determined, based on valid scientific evidence, that the probable benefits to health from use of the device for its intended uses and conditions of use, when accompanied by adequate directions and warnings against unsafe use, outweigh any probable risks.

"The valid scientific evidence used to determine the safety of a device shall adequately demonstrate the absence of unreasonable risk of illness or injury associated with the use of the device for its intended uses and conditions of use."

[Slide.]

The definition of effectiveness says:  "There is a reasonable assurance that a device is effective when it can be determined, based upon valid scientific evidence, that in a significant portion of the target population, the use of the device for its intended uses and conditions of use, when accompanied by adequate directions for use and warnings against unsafe use, will provide clinically significant results."

It is these definitions that we want the panel to keep in mind when you are answering our questions about safety and effectiveness.

DR. WHALEN:  Thank you, Dr. Witten.

We will be, as may be obvious, doing this last session without further break.  We can potentially anticipate, so the audience is aware, going as late as until midnight, and that will be pretty much our absolute cutoff for tonight.

Before I forget to announce it to the panel members, please be aware, when we adjourn tonight, that this room will be totally cleaned out, so all materials should be taken with you and returned tomorrow morning.

As far as the question that arose before the prolonged dinner break as to sponsor having another session, the way that that will be handled is that if there are specific questions that any of the panel members feel need to be addressed, they will bring that up for either sponsor or FDA, and we would ask that if those specific questions are asked, that one, and only one, individual from either of the two organizations go, not to the table, but to the podium and to, as concisely as possible, address the question that has been raised.

By random selection, but because I like to hear her talk, I have decided we will start this open panel session with Dr. Nancy Dubler.  We will then progress around the table towards me going back over to the end of the table.

Dr. Dubler.

PROF. DUBLER:  Thank you, Dr. Whalen.

I thought I would begin by sketching out the difference between an ethical dilemma and the tensions that ethical situations sometimes present, and then, very briefly, list the seven component tensions that I see in the panel deliberation.  So, first, just a paragraph to set it up.

An ethical dilemma is a clash between two goods, that is, between two obligations of equal weight that are accepted as appropriate behavior for a moral individual or a moral society.

An ethical dilemma presents a situation in which competing obligations cannot both be satisfied at the same time.  Whatever choice an individual or society makes leads to an action that is bad or wrong.

The classic example is the tension between the unresolvable tension between a physician's obligation to preserve life and a physician's equally strong obligation to support patient autonomy, which plays out as an unresolvable conflict in physician assisted suicide.

This panel is, in contrast, faced not with an ethical dilemma, but rather with a set of ethical problems or tensions.  In contrast to dilemmas, most tensions can be solved and resolved with a completely satisfactory solution by people of good will, creativity, and flexibility.

Seven tensions that emerge from the discussions today:  The first is society's obligation to protect its citizens, and that is what the FDA does in asking this panel to suggest to it, and in its decision, that silicone breast implants are safe and effective.

It acts because it would be unreasonable for us to assume that citizens, individuals, would have the medical expertise and the epidemiological sophistication and the scientific background to make the kind of judgments that will be made by the panel and the FDA.

The second component in the tension is the right of individuals to make autonomous choices and, in fact, to decide what risks they are willing to court and those that they are not willing to assume, and that is, to balance the risks and benefits of the actions that they take.

Three.  The process that they use in doing that, in executing the personal idiosyncratic calculus that underlies individual choice is what we have labeled the process of informed consent.

So, choice can only be as good as the data that are presented, the choices that are made by the FDA, and the material that is communicated in comprehensible form to the person who will choose.

Four.  The notion of safe and effective, as I understand it, does not mean that an action is without possible medical consequence or risk.  Indeed, most medical interventions promise some benefit and propose some risk, so, in fact, safe and effective might in certain circumstances require someone to balance the possible negative consequences against those benefits that the person will see.

Five.  This is peculiar to the issue we are now talking about.  That is the role of medicine in the developing field of enhancement surgery.  This is a difficult matter.  Physicians are always assumed to be controlled by, and loyal to, the ethic of do no harm, and it was assumed that patients came to them with a problem which they could solve.

The problem with augmentation surgery in the context of enhancement surgery is that many physicians see their role as creating a market for what they have to sell, and the ethic of the market is buyer beware, not do no harm. That is a tension that we have to think about as we go forward.

Two final tensions.  One, real autonomy encompasses the right to make foolish decisions and, indeed, I may think it is a foolish decision for a woman to have augmentation surgery.

On the other hand, the set of values which undergird a woman's decision to seek that surgery is not available to my particular set of values, and whether I think it is right or foolish, if it is within what society says a woman can choose, that it supports her right.

Finally, the final issue is I have been rather startled today by what I see as the lack of obligation that the sponsor felt to pursue a better product over the decade that I have been engaged in these discussions.

I find it strange that this obligation of wise stewardship of the product that they are producing didn't leave any room for at least attempts to improve it in the last decade.

Thank you.

DR. WHALEN:  Dr. Dubler, if there were more lawyers who thought like you, the world would be a better place.  Thank you.

Dr. Newburger.

DR. NEWBURGER:  I was impressed with the excellent quality of follow up with the studies the sponsor presented in that such a large percentage of patients actually were followed through the course of the greater period.  I am troubled by the device failure rate in terms of morbidity of second operations.

As a dermatologist, I don't put in breast implants, but I certainly see the repercussions of some of these ruptured implants in my patients, specifically breast cancer patients who have had reconstruction where they will be in for a body exam and we will find hard nodules in the region of the skin around the breast that has been reconstructed and we have had to sample these to see whether it is metastatic disease.  With a patient who has had reconstruction, those hard papules or nodules have been silicone granulomas, these having developed early after their implantation.  Of course I don't know when the rupture was.

I do see that sometimes the impact of the silicone, which has traveled a distance, may not become manifest for a decade or more.  I also feel a little uncomfortable right now about the short duration of the study in terms of long-term safety, so I would feel real happy knowing the longest follow up which we have with real data on any of these patients.

That's all I have completely from a clinical perspective.

DR. WHALEN:  Dr. Boulware.

DR. BOULWARE:  I will limit my comments to my specialty of rheumatology, primarily the issue of connective tissue diseases.  Let me start by saying that one connective tissue disease is not the same as all connective tissue diseases.

In the practice of rheumatology, the patients that we would typically see with, quote "connective tissue" disease can vary from strain, sprains, to severe rheumatoid arthritis, polymyositis.

Most of us tend to categorize these into--I will give you my own classification, there is very of theme.  They are immune-mediated, inflammatory diseases with systemic or local effects, rheumatoid arthritis, systemic lupus erythematosus, ankylosing spondylitis, to an extent scleroderma, and not in our specialty, inflammatory bowel disease, and other types of inflammatory diseases that have these kind of manifestations.

Then, there is another area as you continue the spectrum that it goes down to fibromyalgia, and this gets very uncomfortable even for rheumatologists in defining who does and doesn't have fibromyalgia.

For the first list of diseases I gave you, the inflammatory ones, the American College of Rheumatology and a number of organizations, the European community have come up with a list of criteria, so that we can even agree what those are, and those sometimes are not always agreeable, but we have data on sensitivity, specificity, so we are fairly comfortable.

When you look through the fibromyalgia literature, the criteria is much newer and now there are people who are being included in that who are in these studies we have seen today that are self-reported, may not have fit criteria, may not even have, they may have been very loosely diagnosed or not diagnosed at all, so I am very uncomfortable.

As I get further down, we have included the look at symptoms in broad categories of muscle, which includes, as I looked at the category of the study done by the sponsors, neck pain, back pain, any kind of pain, weakness, et cetera, which I think many of us, if we filled out today, would have to admit that we had those.

So, as we go from the very--fairly well agreed-upon inflammatory diseases down to almost symptoms, I become more uncomfortable.

I am very comfortable with the data thus far that has been presented, not necessarily by the sponsors, but somewhat by the sponsors, reconfirming other published literature.  That first category of inflammatory diseases, I am fairly comfortable that that is probably a safe area, that there is not a very good relationship.

As I get further down, I am less comfortable, but I know in my practice, very often--and in rheumatology--while what we commonly see, where there is smoke, there is fire, if we equate smoke with a symptom, it doesn't necessarily mean there is a musculoskeletal problem underlying that or a fire.

Very often I will see patients who come in with complaints much like what was listed on here, which turn out to have very different diseases that are not musculoskeletal based, depression being a very common one, and it becomes very cloudy and difficult.

So, I am not so comfortable in that area at this point, but I am comfortable that the relative frequency which they were reported is not very different from what I would think a control population would have for that.

In fact, I was pointed out to a study done in Denmark, reported in 2000, comparing the patients who had cosmetic implants with plain cosmetic surgery without implants, showed very little difference in terms of their complaint level, suggesting that the implant was not the likely cause as much as the population in which the surgery occurred.

DR. WHALEN:  Dr. McGrath.

DR. McGRATH:  I am going to comment as a plastic surgeon user of these devices and I guess start by saying that I really came to this panel eager to hear the data because I have been waiting to hear these numbers and to hear the information, so that I could start to think based on the science that was going to be presented.

Where I am right now is I have been some things that I don't think are issues and some things I really still have questions about.  The thing that I do not think are issues from my own experience is I do believe patients benefit enormously psychologically on both sides, both on the cosmetic side, as well as the reconstructive side, from the use of breast implants and that there is a positive benefit to the gel product.  It does give a more natural feel under the skin particularly in slender patients, and so forth.

I absolutely support the concept of patient autonomy.  I don't feel paternalistic about telling patients whether they are doing the right thing by choosing to have a reconstruction or to enlarge their breasts, but I do feel strongly that they have to have accurate information when they are making that decision.

That then takes me to three areas that seem to be a little bit surprising to me from the numbers that I have heard just over the last 12 hours.  First of all, the reoperation rates.  Now, we know that they are running about 20 percent over three years on the core augment group, they are running about 45 percent over three years on the reconstruction group.

Somehow I had been going along thinking this was--or I had heard or the word was that this was because patients wanted to change their implants to get bigger ones or smaller ones or they didn't like the way they look.  That is not what the data is saying now.  What it is saying is that it was because of complications.

In the augmentation group, 70 percent of them went back for an implant, a second operation because of complications, and 92 percent of the reconstruction patients who went back, went because of complications, not because of some wish to change size or the type of implant.  I think those are high numbers.

The second thing is the rupture rates.  I don't understand why they rupture, and I don't feel like I really got a clear answer yet from the manufacturer.  Perhaps tomorrow morning they can readdress that again, because I think if we don't know why they are rupturing, and I don't buy the argument that it is surgeon trauma because the data that we did have said that they could find something that suggested that they were sharply damaged in 18 percent of the cases either at the time of insertion or removal, and I know in my own case, I am much less careful when I am removing one than when I am inserting it obviously.

So, I don't know how that holds up, but that still leaves me with that other 82 percent that it is very hard to account for why these products aren't more durable.

Also, the question of how to detect these ruptures, to put that financial burden on patients of going for an MRI at frequent intervals, it is a very high price for them, and then if it's ruptured, whether and when to remove it.  I don't have answers to these questions yet.

This question of rupture has enormous implications for my patients.  This puts a financial burden on them, particularly the ones that aren't operating under health insurance.  It's a health risk, and I think it's a psychologic burden, too.  I think as people go back and deal with these problems over and over, it is very emotionally wearing.

The last thing I would bring up is I just haven't seen any new data on mammogram interpretation, and I would like to see some of that, too, in the core augment group.

My last thoughts are that if the manufacturer goes further, I would sure like to see more fleshed out on the informed consent.  I certainly think other than simply perhaps the mention of a potential focus group, I certainly would like to see the opportunity for something much more elaborate and formal in terms of patient education.

Lastly, I think a registry would have to be something that the manufacturer is just going to have to address.

DR. WHALEN:  Dr. Leitch.

DR. LEITCH:  Well, as a cancer surgeon, I, of course, do the surgery that makes a woman have to think about having a breast implant, so I have to be concerned about what her options are for reconstruction and does she really have a choice.

I must say in the prior climate with the silicone gel implants, even though breast cancer patients theoretically had a choice to have them, many of them didn't because they said, well, gosh, it they are not safe for augmentation, why do they think they are safe for me.

So, they would select the saline implants and then be unhappy because they had rippling effect of the implants and the cosmetic result wasn't as nice as they had hoped for.

But these concerns that are being raised today are the things our patients think about and have heard about, and we have to give attention to them.

Now, the other things with respect to cancer that I would like to address are the issues I think that we can give our patients good competence about, and that is that having silicone gel breast implants does not increase their risk for developing a subsequent breast cancer, it is unlikely that it contributes to other cancer site development.  The two sites that were listed most frequently, lung and vulvar cancers, have other lifestyle issues that may contribute to risk in that aspect, so I don't think that has been clearly shown to be associated.

The other thing is mammography with breast implants, and it is clear that adequate mammography in these patients requires additional views, with the Ecklund views, and visualization can also be improved if the implants are placed in the subpectoral location, which is more common these days.

A lot of the studies that were looked at in the IOM Report were studies where mammography, you know, looking at detection of cancer in late diagnosis, mammography was done prior to wide use of Ecklund views.

So, I don't think that data can be stated to be exactly what the data is at the present time where careful Ecklund views are done in implant patients.  The other thing is even in those studies, interestingly, anywhere from 8 to 20 percent of cases were DCIS or the earliest stage with breast cancer that can be found.

So, even in that time when mammography wasn't ideal for implants, still, they were able to detect this earlier stage of cancer.

These days also, women do have other imaging options if their breasts cannot be adequately imaged with implants in place, there are other options with sonography and MRI, which as is being mentioned is an expensive alternative, but nevertheless, it exists as an alternative which we use for people who have dense breasts for other reasons.

I think the big thing that has come out probably for all of us is the issue of implant rupture, and this really I think is where a lot of the discussion is coming up.  I think the sponsor is saying, well, look, you know, we got the saline implants approved and the complication rate in some areas is less with the silicone, it stands to reason that these should be approved, as well, you might think that would be the case.

So, I think the differential that we are really looking at here is the issue of implant rupture and what that means to the patient.  For example, with mammography, it makes a difference if it's an extracapsular versus an intracapsular rupture because if the silicone gets out of capsule into the substance of the breast tissue, then, that opaque material obliterates the breast tissue, so you can't  examine it well.  Even if you do the Ecklund views, you can't push away that extravasated silicone, so it does cause a problem in that regard.

Extravasated material can infiltrate the muscle of the chest wall and be very painful for the patient, I have certainly seen that, generally, in patients with older implants.  I must say among my breast cancer patients, I really have not experienced a lot of complications with the silicone gel implants.  Most are from patients who have implants many years ago and then have extravasation problems.

So, these are the issues.  I think some of the things in terms of rheumatologic issues, maybe it is the case that people who have the extracapsular extension are the ones that become more sensitive to this although certainly I think all of us have seen patients who had actual silicone injections into their breasts and while they have a lot of local issues in terms of lumps and bumps, they don't have a lot of systemic symptoms.

So, it may take us several factors to come together to make that extracapsular silicone be an issue, but that may be in the study that Dr. Brown cited may be an example of that although it is not a perfect study to look at it.

That summarizes my concerns at the present time.  I do think the issues about the follow-up study we need to go into if the product were to be approved.  I think all of us do have concerns about the follow-up, as Dr. McGrath mentioned, in terms of MRI and the role it would play, so I think we need to address those types of issues.

DR. WHALEN:  Dr. Chang.

DR. CHANG:  I make my comments in the context of my being a practicing plastic surgeon.  I believe that the sponsor, in presenting their data, have really shown due diligence in gathering the preclinical studies, as well as in the excellent follow-up of patients in the time period of one and two years, and most of the three-year data.

I believe they have shown that there is efficacy and a high rate of patient satisfaction.

Some other comments.  Remaining unanswered questions, as Dr. McGrath said, the important question is why do ruptures occur.  It is a fairly high rate for reconstructive patients, 6 percent, why are they caused, how can we improve and try to diminish these numbers.

Another question I have is I am just wondering, how did the sponsor decide on the exact thickness to use for the implants, where they may range from 13 to 18,000ths of an inch, and would it have made a difference in pliability and rupture rate if it had gone to 20,000ths of an inch, so that is a question which may or may not be addressed in closing statements by the sponsor.

The last major question I have are what the repercussions of extracapsular gel migration when it occurs.

I was again impressed three years ago by the due diligence and presentation of data for one-year, three-year, and five-year follow up of saline implants, sponsor has related the complication rate to saline implants, data that were presented three years ago.

I was impressed that they had physician follow-up of these patients, so then I ask, and my statement is I believe there is an irrational optimism on the part of sponsor to think that a Phase II study using mail-in survey is going to come anywhere near the success rate that sponsor has had in gathering significant data to see what happens on a long-term basis.

So much of the public comment stressed the importance of this long-term data.  So, I really don't think that a mail-in survey would be appropriate if one wants to achieve the success that has been shown in presenting data during this presentation today.

Lastly, the issue, if the FDA were to approve this device, is the concerns about what kind of labeling and the patient education of the reoperation rate and local complications.

Thank you.

DR. WHALEN:  Dr. Choti.

DR. CHOTI:  Perhaps to reiterate some of the points made, just as a clinician, to a little bit of what I have learned in my concerns, one is regarding the device.  I am concerned about the performance of the device.

I think we have learned that there are some issues related to it, particularly the reoperation rate, the local complication rate, particularly of concern, the pain, capsule formation, and particularly, as mentioned over and over, the leakage rate, which it is clear that it's a delayed effect, or it continues to go up.

So, while the systemic effects are less clear, I think, and I am perhaps somewhat more comfortable with the fact, or at least the evidence is less clear that there is a strong correlation with the systemic effect, I think that is also still a concern and it's an open question mark, difficult to address, it's a complicated question to look at the systemic effects, but I think to say that free silicon from a ruptured extracapsular implant, that the presence of free silicon isn't totally inert, I think it is still difficult to say and still a concern.

I think, though, the local effect of a rupture is a potential problem, and as I said, the performance of the device results in a significant potential for delayed rupture and leakage.  So, these are some of the questions that I am still grappling with.

I think the study was well done, I think it was as designed.  I think the follow-up was excellent, I mean the compliance was excellent with follow-up.  The problem is really the short-term follow-up.  I think that a lot of the concerns we have are related to the delayed effects, and that is really not addressed by two- and three-year follow-up.

The review of the peer-reviewed, although compelling, and certainly there were a lot of groups that came out and suggested that the concerns perhaps are not as great as one should have thought about 10 years ago or 12 years ago, but I still think that looking at this evidence, it is still not clear in my mind.

For example, I am concerned why we really haven't seen much data out of Europe, for example.  There are really no good registries, so it is still not clear from large population-based studies as to the performance, long-term performance of these devices.

I am looking forward to having some of these questions addressed hopefully.  Those are the main points that I would like to make.

Thank you.

DR. WHALEN:  Dr. Blumenstein.

DR. BLUMENSTEIN:  The form in which the data from the sponsor was provided to me was problematic.  It was given to me as a CD-ROM, contained a heap of PDF files with one PDF file showing the contents of the CD-ROM as a list of topics along with PDF file names.

Furthermore, the PDF files contained scanned images of pages rather than converted documents, therefore, were not searchable.  A collection of PDF files without navigation tools and searching capabilities is much less than an optimal replacement for paper.

It was therefore difficult and time-consuming to assess the data on the CD-ROM.  It is surprising to me that the full capabilities of the PDF technology were not exploited and the sponsor would condone this type of barrier to easy access to the data.

From a statistical viewpoint, the presentation of the PMA data is deficient and unimaginative.  The FDA guidance document was inadequately implemented, and there was no effort to go beyond the underlying principles of the guidance document.

There were no substantive efforts to employ supplemental analytic techniques to enhance the presentation of the data.  For example, minimal graphs or other visual aids were included on the CD-ROM.  Also, the use of time, as in Kaplan-Meier plots, could help by illustrating where there might be higher hazards or non-constant hazards of events.

An example of an inadequate implantation of the FDA guidelines is the absence of a comparison of the characteristics of women who have follow-up to those who do not.  Another example is the absence of proportional hazard regression models, that is Cox regression, using time-dependent covariates.  Both of these are in the guidance document.

There is no presentation of an overall composite failure outcome analysis, such as bad-thing free survival. Such an analysis would be the single most important characterization of device risk and performance.

Instead, the sponsor presented three composite outcome analyses partitioning bad things into three classes of bad things.  The analysis by the sponsor dilutes the overall risk and negative aspects of the device's performance.

Also, the presentation of the sponsor's composite risk analyses is inadequate, for example, the sponsor's material has no graphic aids, such as the Kaplan-Meier plots.

The constellation of side effects that were prospectively collected to potentially detect increases over two years that are notable, but difficult to interpret because of the absence of a control group.

Specifically, the increase in frequency of the side effects over two years is confounded by aging.  Statistical modeling using all three patient groups and patient age at entry as a covariate might provide a means of clarifying the meaning of these data.

The three groups differ with respect to mean age and therefore can provide an estimate of the trend over age in the reporting of these side effects.  The length of follow-up is inadequate.  The amount of follow-up does not meet the minimal requirements in the FDA guidance document.

In addition, it is my judgment that the minimal follow-up specified in the FDA guidance document is inadequate.

DR. WHALEN:  Brent, next time try not to sugar-coat it so much.

Dr. Conant.

DR. CONANT:  That's a tough act to follow.

I would like to start by speaking of my concern about the underestimation of rupture based on the small number of patients screened with MR and only for such a short interval.

I am concerned about the concept of silent rupture and that since we are not detecting these cases unless these women are screened, we may need to create some prospective or some guidelines for screening of women who have silicone implants placed.

I am concerned also about the lack of evaluation of the individual rupture cases.  From my experience in the operating room helping take out ruptured implants as an imager with an ultrasound.  You can't see the free silicone, so by just visually inspecting, saying there is no extracapsular silicone, I am not sure that is an adequate endpoint, and it worries me in terms of local complications.

I am going to leave the connective tissue disorder discussion to my rheumatology colleagues.

So, I am concerned that not only are we underestimating the ruptures because of lack of long-term follow-up and small numbers screened even at one-year and three-year, that we are not fully evaluating the extracapsular extent of silicone.

I am also concerned about the possible simplification of the patient satisfaction measurements. There are so many other aspects of satisfaction than are you happy with your implant, and I would love to see a little bit better breakdown of that, for example, are you concerned about cancer screening with your implant, are you concerned about the need for implant screening and the cost of that.

Getting to my specialty breast imaging, I would just like to make a few comments about that, that our additional imaging now really does image the breast tissue better than it used to.  Our limitations are really those patients with severe capsule contraction.

I personally have not seen extravasation cause a limitation in detection, because extravasation or free silicone is usually focal or thin sheets streaming along the pectoralis muscle or along the margin of the implant, but it rarely gets in the way of cancer detection, but when a lesion is detected with breast imaging, our options of how to approach the lesion with biopsy are often limited.

For example, with capsular contraction, we can't do as many of our percutaneous simple stereotactic core biopsies or even ultrasound-guided biopsies because of the limitation of space around the implant and the inability to manipulate the breast, so it becomes a more complex situation with more risks than the normal patient.

And what else - I am a little concerned also about the screening compliance, which we talked about earlier.  I personally have seen patients refuse screening mammography because they have their implants and they are worried about that.  I would hate for those women to have the burden of more expensive screening techniques.

DR. WHALEN:  Dr. Lawrence.

DR. LAWRENCE:  Thank you.

While I agree with many of the things that have been said here this evening, I am going to focus on one issue, which is my particular area of interest, and that is the child.

Therefore, we find that mothers who get implants fall into two major groups from that standpoint, those who are outside the childbearing years and therefore it becomes a decision and issue for the mother alone, and perhaps choice is appropriate.

For the woman who is still in her childbearing years, it is a very important decision because we are talking about putting a foreign body in and some interesting chemistry into the breast.

I think the data is not very robust as to the effect on the child who is born to a mother who has an implant, because there have been no controls and the follow-up has not been very extensive.

One would need to look first at the exposure during pregnancy, and of all of the toxins and non-toxins that are in a woman's body, the exposure is far greater transplacentally than it is through the breast milk, so that even collecting cord bloods would be an interesting thing to do, and certainly to follow children born to mothers with implants and without implants for a much longer period of time, a minimum of 10 years.

Then, with respect to lactation, it was also stated that it had no effect on lactation.  Well, the only way one could confirm that would be if you were studying women who had lactated before the implant, then, had an implant and failed to lactate.

One has to look at the reason for having an implant, and for some women, it is because their breasts are rather abnormal, not only small, but abnormal.  The potential for that woman to be able to produce milk is much lower than the average woman, so the fact she has an implant doesn't necessarily explain the reason that she is unable to lactate when she has a baby. Therefore, that issue has to be discerned.

Also, in the surgical procedure for placing the implant, it should not be done peri-areolarly.  That would cut all the appropriate nerves for the let-down reflect, and it can be inserted otherwise, and I am not speaking as a surgeon now, I will quickly reassure you, but it is an important issue for women of that age group just how you operate on the breast.

Then, with respect to what gets into the milk, the American Academy of Pediatrics did say that implants were not a problem for breast milk.  It is interesting because to my knowledge, they have never studied it.

We do know that if you give mothers two tablespoons of Maalox with simethicone you can raise the D2 levels in the breast milk rather dramatically, so silicone as an inert material most of the time, does that make any difference, will it be absorbed from the baby's stomach if it is in the breast milk?

So, those questions have not been answered.  They could be answered by some very simple studies.

DR. WHALEN:  Dr. Lieberman.

DR. LIEBERMAN:  Well, I think that looking at the data, I am concerned at the high rate of local complications, and I think that they are serious.  I know that the saline implants have been approved with comparable rates of local complications, but I think the question is are those entirely comparable situations.

I don't think we know if rupture is comparable to deflation and saline, and how serious that might be.  It is not clear to me how often there is extracapsular silicone that is out there and then what that means.

I think Dr. Brown's study, starting to look at the effect of extracapsular silicone versus when the silicone stays intracapsular raises some real questions about an association with things like fibromyalgia, and other studies, bigger studies that look at the entire group of women with breast implants wouldn't detect that because it's a small proportion of the women, so negative studies in larger populations don't particularly reassure me with regard to that.

I think there is also an increase--it kind of goes along with that increase in symptoms that we saw that is very hard to interpret because there is no control group, and as we get older, we all get more symptoms, but some of the rises over a two-year period look impressive.  I am not quite sure what to do with that.

I think there is clearly a need for much more data and for much longer term follow-up to get more data on some of these outcomes, such as connective tissue diseases, cancer.  I would like to see the data from all of the studies with cancer, and not discard the ones that aren't significant, to look if the relative risk is about the same in all of them and the difference in significance is only because of the number of people in the study.

The other thing I am concerned about is should this get approved, the issues of informed consent, of making sure that women are truly educated and truly able to make educated choices is very important, and I think the other question will become how do we monitor these women over time, both to learn from their experience, but also to assure their safety.

DR. WHALEN:  Dr. Manno.

DR. MANNO:  I will make some comments as a toxicologist, but I am also concerned on experimental design in some areas.  I pretty much agree with what the physicians have said here with the concerns.

I am concerned about the rupture problem, the etiology of the problem, what is happening to that silicone. There was a comment that it is not moving out of the capsule, but yet it looks like, according to the physicians who work with these things, they can't see it.

There was also a comment made concerning, well, we can't ask the surgeons to look too far, you can't tell them what to do.  That offends me.  If you are doing any kind of a research study, you have to have the cooperation, you have to have the design, and the design has to be followed.

I think if criteria were given and the principal investigators at the various sites were told this is what they are going to do and how they are going to do it, they would do it.  If they didn't, you need to find some who will.

One thing comes to my mind in terms of the questionnaire asked about the patient's satisfaction with the implant.  I don't like a yes/no, because you only have 50-50 there.  If many of the questions asked lend themselves to questions that they could grade them on a zero to 5, pain, overall, is your pain--if you don't have pain, it's zero, but is it excruciating, as some of the women described today, that could be your 5.

There are tools out there to do this.  Then, you could quantitate your response a little bit better.

In terms of the controls, no toxicology study is worth its salt without a set of controls, and that disturbs me a great deal.  I am concerned in some places with the gel, studies that were in the package that was sent to us that they are comparing or wanting to compare historical data with more current data.

That bothers me a little bit because your controls are there to be run at a time that your substancing question is being tested.  Yes, it may be the same chemical formulation, there is still a different environment in which this is going on.  They are possibly different people doing the exposure, and so on, and I think that that is a weakness in this area.

Going back to the question that the gel is not migrating from the site based on the fact that it was injected in the animals and it stayed there, one thing that bothers me a little bit that came up from the discussions today, was that if I remember correctly, the injection of the silicone was given in the area of the back.

The comment today about any changes that may occur in the stickiness as an implant stays in, et cetera, you are putting it in a different tissue milieu when you are putting it in breast tissue.  You are putting it in lipid tissue, which will make it a different environment in the body, and this hasn't been looked at.

There is also a difference--this came to mind also in the fact of the rupture that came up with the comment today about the testing of the implant in terms of there was no test for shear effect on these implants in terms of rupture.  It just came to my mind has there been any correlation to the first implant rupture in a patient as to their handedness.  That may be one point that could be looked at, if you have any data like that.

You are obviously, if you are right-handed, you are going to be doing more motion of all directions with your right hand.  If it tends to be the rupture occurs earliest that you find it, on the right side, then, there may be a problem with shear or some other type of abrasive movement involved.

I think another area that gives me a great deal of concern is in the reproductive studies with the two generations that were done.  When the F1 generation was looked at for birth defects, for example, and anything there, was looked at a few hours to a day before the babies would have been born or the pups would have been born, delivered.

The same is true with the F2 generation.  We are concerned in terms of safety not only with birth defects, but we are looking at problems that may develop in the future like occurred with the DES where it didn't show up in the offspring until they were in their late teens, I believe it was, or early 20s.

So, I think that we are missing that segment of information as to long-term effects on the offspring.

I am interested also because I am in the Department of Psychiatry, go back to the incidence of suicide and depression that might be occurring in these individuals.  Granted, you perhaps did not have that information at the beginning, but it would not be unreasonable, in my mind, as one of the instruments of evaluation on intake of patients into these things would be to have a psychiatric evaluation similar to what is done with organ transplant patients, that give you some idea of their emotional stability at that point, and if there is any underlying problem that can be determined professionally for having later instances of depression or suicide occur.

I am also thinking in terms of the non-cancer reporting and the cancer reporting that has been in the literature in terms of genetic histories on these people especially in view of the more recent development of the marker for breast cancer.  I think it would not be unusual to include looking at genetic histories even by questionnaire and then, in certain instances, decide whether or not the genetic testing for that particular marker be done.

I think I will stop at this point.

DR. WHALEN:  Dr. Olding.

DR. OLDING:  I tend to be a lumper and an organizer into little areas, so for me, it is really about systemic problems and about local problems.  I have been surprised that I have been reassured in some areas that I thought that I was not going to be, and I have been also equally surprised that I have been concerned about areas that I had not expected to.

I respond to this as a plastic surgeon and as an advocate for my patients.  The concerns you have heard about already, the local complications, not only the reoperation rate and the rupture rate, but also what about the long-term follow-up of those patients, what is an appropriate follow-up, what is the length of that follow-up, what is the financial burden to the patient.

The things that I have been reassured about are the suicide rate, about the probable lack of association with breast cancer, with collagen vascular diseases, neurologic disorders.  I had really expected that to be the focal point of the anxiety that we heard about or the tension that we heard about at the very beginning.  I have much less tension, much less anxiety about those than I did before I reached this area.

I believe we do need to think about education, not only for our patients, but for the people who operate on those patients, and the more educated we all are, I think the happier we will be.

More importantly, I am also anxious to hear the rebuttal of our concerns by the company.

Thank you.

DR. WHALEN:  Ms. Brown.

MS. BROWN:  As a person who has been in companies responsible for development projects, I have been trying to put myself in the company's shoes as to how would you conduct a study on a gel implant today in order to get it approved.

First of all, the FDA has laid out a guidance document for testing silicone gel implants.  It includes chemistry, toxicology, mechanical data, and how to conduct the clinical studies, and I think that the company has done a good job of following the guidance document, which is where everybody should start in order to develop products like these.

The other starting point is that saline implants have already been approved, so they set a risk/benefit ratio that has already been determined to be acceptable.  So, for me, the question is what is the additional risk, what is residual risk that might be associated with the gel implants.

I think that with respect to effectiveness, one of the previous panel members commented on that maybe the effectiveness parameters weren't adequate, and I believe the guidance document pretty much laid out what the effectiveness were that were supposed to be collected in the study, and that all would have been agreed on by FDA and the company probably years ago in the protocol with respect to outcomes, and once you set those, they pretty much become something that the company and the FDA walk forward with together.

With respect to is this device safe, it appeared to me that the company presented data, the complication rates between the saline implants, the silicone implants appeared to be similar, if not, lower for the gel implants with the exception perhaps of the rupture.

I am interested in hearing the company's comment on the relationship between disease and extracapsular rupture, as well.  To me, the reoperation rate is troubling, but it was always troubling even with the saline-filled implants, but I am like Dr. Dubler, I think a patient should have a right to choose with adequate informed consent whether or not they do this, especially cancer patients.

I have a sister and a sister-in-law who both were cancer patients, and I know the tram procedure doesn't cover both breasts, so what do those patients do.  So, I am a real advocate of the right to choose.

I also think, though, that the company has an obligation to have extremely good informed consent, and I don't know it takes to get adequate informed consent where patients really, really hear the message that they really are signing up for a potential reoperation, but I think that whatever--I think I saw in the materials that the company was going to do an evaluation of the informed consent process or the patient advice process, and I think that that really should be an important focus to make sure that patients are adequately informed.

With respect to extent to follow-up, two-year follow-up, I think two-, three-year follow-up is extremely consistent with other devices.  It is probably two-year follow-up is consistent with other implants that the agency has approved.

I know that the FDA is always concerned about level playing field, and so I know of no other implantable device that has had more than two-year follow-up, and correct me if I am wrong, someone from FDA, but to ask for more than two-year follow-up, I think is fairly extraordinary for an approval process, not for--I think the company has already signed up and the guidance document already specifies a 10-year follow-up, and that is certainly appropriate.

Those are my comments.

DR. WHALEN:  Ms. Gilbert.

MS. GILBERT:  Thank you.

I think the consumers simply need to have more answers, more honest answers to some of the concerns raised. The sobering testimony provided today was a contrast to the findings of the sponsor.

I strongly recommend long-term studies to address failure and rupture, follow-up, delayed effects of gel migration, again, the follow-up concerns and the need for recommendations.  The patient and the consumers deserve to have honest information, and we need to have facts and consistent educational information.

That is all I have to say.

DR. WHALEN:  Going back over to Dr. Miller.

DR. MILLER:  Thank you.

I want to say first that I listened carefully to everybody who testified today as best as I could track, and I appreciate everybody who came today to speak to us and the passion that I saw, sincerity that I saw in everybody who spoke.

Let me say that my career as a plastic surgeon has spanned the entire implant controversy from the beginning, and I will match the passion that I saw in this room today with my own passion for the safety of my patients and for providing good care, and that is where I come from on this whole issue.

I am very concerned to understand what happens with these implants and the impact that they have on the health of my patients.

When I look at the mandate we have to look at the safety and effectiveness, when I think about the effectiveness, it is clear the implants are effective.  I mean that is clear.  The implants are designed to make the breasts larger or to create a breast where there was no breast before.

Nobody would argue that the implants don't successfully do that, that they are very successful at doing that.  The benefit of the implants, because it overlaps many areas beyond that very strict criteria, creates a lot of opportunity for controversy, because suddenly we not are asking the implants make the breast bigger, we are asking should you make a breast bigger.

A lot of different opinions begin to creep in about the appropriateness of doing that, and this begins to shade this cost/benefit, this risk/benefit analysis, and that is why I think this problem is so difficult to resolve and so unique.

I am convinced of the effectiveness of the device. As far as the safety goes, I think it is okay to hold the implants to a slightly higher bar than other devices because of some controversy possible over the benefit.  I mean if there is a questionable benefit, then, there has to be maybe a higher level of safety, I am okay with that.

I am satisfied with the safety that I have seen in terms of the long-term effects of these implants.  I think there are questions that exist, but I think that the experience of 40 years, I think that the nice analysis of the data, and the published studies that were done by both the FDA and the sponsor summarized the literature very well.

To me, it is fairly convincing that there is no evidence that these things harm people in the long term with systemic problems.  It is hard to make a case for that.  I am a little dismayed at the data that has been presented in terms of the symptomatology because that is so hard to interpret and all it does is cast doubt about how systemically safe these devices are without answering any questions, it is just another shadow of doubt is thrown in, and that, I didn't find very helpful, but the data is there.

I am troubled by the number of people who genuinely feel harmed by these devices.  When I listened to these stories and hear about the level of personal life disruption and injury, that people sincerely believe has been caused by these implants, I have to wonder why is this.

I think we need to try and answer that question dispassionately to find a solution, because these people are suffering, there is no question about, and not suffering for no reason.  So, what is the reason?

I am unconvinced it's the breast implants.  I think it does them no service to just endorse the idea that the breast implant is the cause, because if the breast implants are not the cause, then, we are not searching for the right cause in these people.  I think we need to work hard to study these individuals and see why are they having these kind of problems.

There was a very interesting paper I want to refer to the rest of the panel and the rest of the audience by Barsky and Boris out of Boston, talking about the functional somatic syndrome, and I don't think every patient who we saw today is afflicted with the functional somatic syndrome, but I think a large number of them can be categorized in this category, and I would encourage a consideration of that, an explanation of that.

A couple of other points.  The issue of reoperation as a complication, I mean aesthetic surgery and reconstructive surgery with the device in some ways it is a process, and with other procedures and other operations, a reoperation is a failure and a complication.

It is not really fair to consider a reoperation in a patient who you have implanted something to change their breasts necessarily as a complication, because you cannot make the breast like it was originally, and putting the implant in, you have introduced an element to try and create the final product of a shape.

Now, rarely, and I tell my patients this, is this going to be done perfectly with one go-around, and it depends how high your expectations are and how close to the perfect idea you want your result to be that will determine how many operations you are going to require.

So, some of the patients, the higher standards of expectations require the most operations, but in doing them, it does not represent a failure for that patient, it's a continued pursuit of that patient's goals.  So, that clouds a little bit the data when we just consider reoperation a failure.  I disagree with that view.

One final comment is I think that these devices, placing one of these devices is deceptively simple.  As a surgeon, it seems so simple.  You create an incision, you create a pocket, you put a device in, and you are done.  This invites a host of surgeons of all types to decide to do this procedure.

There are technical issues with putting an implant in properly to avoid wrinkles, to place it properly, to not cut it with a needle, all these things, to properly inform your patient, to follow them responsibly.  These are things which make a successful operation and an unsuccessful one.

Not every surgeon is really cut from the same cloth in this regard.  I think that perhaps some element of considering who is qualified to place these devices should be considered in terms of considering their safety and efficacy.

Thank you.

DR. WHALEN:  Dr. Anderson.

DR. ANDERSON:  As a breast cancer surgeon, I came here this morning thinking that what I thought I knew, I believe in the efficacy.  If you have handled a silicone versus a saline implant, there is no question that silicone is a better implant.

I believe in the right to choose.  I think my patients should have that right to choose.  Life is not without risk, so there are risks that come with all of the things that we do, there is no reason that this should be an exception.

The problem that I have hit, I must say that it was the presentation of the company that actually made me back off from my confidence in the device, because I didn't hear the answers to the questions that we were looking for in terms of the local complications of these, and also addressing these possible systemic effects in very small groups of people, because at a minimum, I don't think we can say that it is impossible, but at a minimum, we need to be able to say to our patients, well, this only happens in X percent, because then we can properly inform the patient that this is what your risks are.

So, we need that from you.  What I also didn't hear from the company were guidelines about how these should be used.  When the question came of what do you do, are we supposed to take them out at 10 years, do you wait until they rupture, there was no answer, and that should have been thought through upfront.

So, on the one hand, if this device is approved by the committee, then, we need to have a real plan about how the follow-up is done, how the data is collected, and not window dressing, but something that is enforceable and realistic.

On the other hand, if we don't approve this device at this time, then, it is the obligation of the FDA to be very clear about what it is that these people are supposed to do to prove that this device is efficacious, because it does sound like they actually followed the rules of what was laid out by the FDA.

So, if we are going to pull back on it now, we can't do that multiple times.

DR. WHALEN:  Dr. Li.

DR. LI:  My comments are basically circled around the same worries everybody else had, which is rupture and gel bleeding, so to those two regards, let me cover four areas - the manufacturing issues, extractables, mechanical testing, and retrievals.

Because we don't seem to understand or be able to duplicate the issue of rupture, this kind of raises the question of what is important and what is not important, so in the absence of actually knowing what causes the rupture, it kind of brings everything into question.

Some of this data, the company may already have, the FDA may even have it.  I am just going by what I could find and glean out of the data.

For instance, Dr. Chang mentioned the thickness of the implant, so given a particular style of the implant, the thickness could vary a factor of 4 just for any given style, from 18,000ths to 80,000ths, so certainly one would think that the 80,000ths thickness wall would be the preferred one because diffusion would be more difficult, however, from the fraction mechanics point, a thicker implant doesn't necessarily guarantee you better rupture, in fact, depending if it's permanent set, a thicker wall actually may work against you in terms of rupture, so it is just not clear,  out of these rupture mechanisms, what the influence of thickness is.

The other thing that is not discussed is the sizes of these implants.  We have talked about different models of implants, but for every model, there are a variety of sizes. I think the volumes vary by a factor of 4 almost, anywhere from 100 or 90 up to 750 milliliters, and the effect of thickness combined with volume has not been broken out.

Again, one wonders in a classic fraction mechanic test, do these design issues play a role in the rupture.

As far as manufacturing consistency goes, they propose using mechanical tests like penetrometers or modulus mechanical testing to give an indication of crosslink density, which they appear to have done for the shell, at least given a particular crosslink density.

They seem to have a penetrometer reading that reflects that crosslink density, but I am unclear if they know what the sensitivity is of changes of crosslink density to the penetrometer data.

For instance, if the number of crosslinks per molecule tripled, would you still get the same penetrometer reading or would you not?  It is unclear to me that the mechanical tests are actually a good indication of changes in crosslink density, and it doesn't appear, at least in the data that I saw, that their testing of the crosslink density of the gel is nearly as extensive as they are of the two base layers.

Extractables.  This is the gel bleeding concept.  This, you did a substantial amount of work that, just as a body of literature, just confused the heck out of me, for two reasons.  One, as has been pointed out, the extracting conditions are not in-vivo conditions.

One of the problems is that I found at least four different solvents used in these extracting tests depending on when they were done and who did them.  There was hexane,  there was dichloromethane, there was isopropanol, there was ethanol.  In some cases, the batch of implants done where the lot of implants was made under a label of an experimental production lot, they weren't pulled out of the regular lot.

Now, maybe they were the same, and maybe not, but certainly the workers making those implants knew they were making a set of implants that were going to be tested differently than the regular batch, so that certainly raises the issue of was that lot particularly good, particularly bad, or was it, in fact, representative of all the rest.

Depending on the solvents you used, you have got a wide range of molecular weights, and you also got a different set of molecular entities.  You have got both linear silicones, you have got cyclic silicones.

Now, because we don't really know what these local and broader effects are, we actually have no idea which are the most biologically important ones - is it the low molecular weight, is it the high molecular weight, is it the linear, is it the cyclic or the diphenyls, we actually have no clue.  Circumposed on that is we have no idea what the dose-response is, if there is one, for these materials.

As far as epidemiology goes, it seems to me, as the non-epidemiologist, and I know just enough statistics to be dangerous, their assessment that there doesn't seem to be any systemic effect is based on a large number of groups, however, we do know there are a subset of patients where these implants have ruptured, in fact, in many cases, we have heard that sometimes the implant stays ruptured in the patient for some length of time.

I haven't seen any breakout of what happened, what the epidemiology of that subgroup is over time, so maybe if you look at the whole thousand women where small fractions of them have ruptured, you can't see it, but if you actually followed one particular group that you know had ruptured, I don't know if that breaks out or not.

The mechanical testing, they have satisfied the ASTM and FDA requirements, but just as a comment to the rest of my panel colleagues, neither of those groups of tests are performance-related criteria.  In fact, the ASTM guidelines almost in every particular test says you ought to do these to ensure that you making the right thing, but if you make them this way, we don't really guarantee you, you will get better performance.

Now, this isn't the fault of the sponsor or the FDA, that is just the nature of the ASTM tests.  My problem with the mechanical tests is they don't seem to reflect our clinical experience.  In your fatigue tests, you can't seem to break them, but patients seem to do it kind of regularly. Your fatigue tests would say it is impossible to break these things during mammography, yet, at least we have a series of patients who appear to have done it at that time.

There doesn't seem to be any attempt to actually verify or solidify your hypotheses of where these cracks come from or these ruptures come from.  Typically, the approach in this is that cracks either occur where the implant is weakest, or they occur where the stresses are highest, and I have seen really no engineering or modeling in a classical sense to demonstrate that, in fact, what the increase is in stress out of fold, and something that has a very low modulus, it could be the increase in stress at the fold is minimal or not, but I haven't actually seen that calculation.

Again, that would be a function of thickness.  So, although it seems kind of intuitive that if you have a fold, it ought to be easier to tear, that may not be true if you have something very soft and pliable like an elastomer.

So, the problem with the mechanical tests is there seems to be a disconnect between the testing and the actual failures that we are all worried about, in this case, rupture.

That brings us to the last comments I want to make on the retrieval.  So, first, let me congratulate on you for even attempting to do the retrieval stuff.  I have been doing this for 10 years and it's a very difficult thing to do, and I think it is great that you attempted to do that.

Now, that being said, let me criticize it.  The data seems to be at least reported, if not collected, with kind of a bias toward the outcome.  In other words, you may have collected more data than you presented, which seems to be the case.  For instance, I heard that there was weights of these implants taken, but I didn't see any of that data in my panel packet.

So, the problem is the data doesn't seem to be a general study of looking at all the factors that one could measure on retrievals, and then trying to associate those factors with rupture, for instance, factors that are not described, at least in my panel packet, are the size of the implant in addition to the model, the length of implantation of that particular device, the actual direction of the crack and the length of the crack, and if there is any association between the length of implantation and whether or not there was a crack.

Kind of arguing against the fact that it's surgeon-induced is that some of the models appear to have cracks on the anterior of the implant, and other models had cracks on the posterior of the implant, so unless the surgeons are going to tell me somehow these two different styles have a different surgical technique, it is a little hard to imagine why, in one particular model, the cuts seem to be on the front and the others, they seem to be on the back.

Again, the biggest bias comes, in my view, it comes from kind of the smooth versus sharp edge rupture description.  That, in the kindest terms, is a very non-classical approach to implant failure.  In fact, unless you can do some modeling or demonstrate to me these two, in fact, have different mechanisms occurring, in fact, they could be manifestations of the same failure mode, but because of the direction of stress or the location of stress, one appears to be curved and one appears to be straight, so it just basically puts in a question of I actually don't understand at all what is going on there.

There seemed to be an inconsistent collection of data and some of the implants, they seem to be weighed and other ones they weren't weighed, but there was no assessment, for instance, if an implant was ruptured, was there an assessment of okay, well, you know, the implant now weighs half of what it did or maybe it only weighs less than what it did.

I am not even sure what the variation is in brand-new implants.  If I got 100 style 10's, I don't really know what the weight variation is in that.  I don't know if it's inconsistent or if that is something that you can actually measure in the retrieved components.

Lastly, there seems to be a general belief that as the implant stays in longer, the number of mechanical problems seemed to increase, and that whole area hasn't been addressed at all.  So, if that is the case, two things come to my mind, either it's a fatigue issue and the longer it is in, it actually just sees more number of cycles, the other, more insidious mechanism could be, in fact, that there is some kind of chemical change going on in the implant, which some people reported as kind of a stiffness or a color change, or a physical change in the material of the implant, and again, that is a whole area that we just have no idea about.

So, there is a tremendous amount of testing that left me no better understanding of why these things rupture or, in fact, even a discussion of actually how much they rupture, or why, and then this issue of the gel bleeding.

The extraction data is interesting, but it is so far from the in vivo condition, I am not really sure what to do with it.

DR. WHALEN:  As chair of this panel today and in the past, I don't often articulate my analysis or opinions on matters that are being considered.  My job is somewhat more mundane to do things such as keep us on time, for which I get an F today.

However, I would like to, on the one hand, congratulate Inamed, because I think they did a phenomenal job in answering the questions that FDA asked them to answer.  I remain astounded at the questions that FDA asked them, and with all due respect to this being the normal process and we never ask for more than two years, et cetera, that has been articulated and is the situation as it stands, we have been getting decubitus ulcers today over a matter of a product that was taken off the market a decade ago for concerns about long-term issues in these patients, a decade has passed, and we sat here today talking about two and three year data, and I am just flabbergasted by that, that we are going to be asked to give a cogent answer to safety and effectiveness when the data wasn't asked for, nor was it given.  That's my opinion.

With that in mind, let us go into the questions, please, Commander Allen, if you would present those questions and guide us through them, and each of the panel members will be giving a brief insight into how they feel about each of those questions, following which I will attempt to give a synopsis of the panel's comments to Dr. Witten, question by question.  No. 1.

Open Panel Discussion and Questions

CDR ALLEN:  FDA is going to seek the panel's input on the following seven questions.

No. 1.  Please discuss the adequacy of the information to determine the safety of this product with respect to asymptomatic rupture.

Please start your discussion.

DR. WHALEN:  We will start going around the table first with Dr. Newburger.

DR. NEWBURGER:  I don't feel that there is information available to us to determine the safety of the product with respect to asymptomatic rupture.  There certainly is a difference in symptoms associated with rupture in that augmentation patients seemed to be much more symptomatic than reconstruction patients.

I am thinking that that may be in light of the scarring that one would expect to see at the surgical implant site with the reconstruction patients, but I am not comfortable with addressing that question.  I don't have information that makes me comfortable discussing any safety.

DR. WHALEN:  Dr. Boulware.

DR. BOULWARE:  Can I clarify that with what was asked for them in terms of a timeline?

DR. WHALEN:  You certainly can.

DR. BOULWARE:  Because I think in terms of a short-term two to three year, I think they have answered that and the data is adequate, but the question has become greater, I guess, with the long-term outcomes for this.

So, I am comfortable, I guess, with the short-term data,  but not the long-term, but I want to be fair to the company in that.  I don't know what was asked of them in terms of providing information about safety with respect to asymptomatic rupture relative to implantation time.

Does that make sense?

DR. WHALEN:  Thank you.

Dr. McGrath.

DR. McGRATH:  Good question.  Are we going to stop and answer that before I say anything or after?

DR. WHALEN:  I personally consider that an answer in and of itself saying the short term, yes, long term, no, and then I will try to put that into the synopsis to Dr. Witten, and she clearly has a right to reject the answer that I give her.

DR. McGRATH:  Okay.  The only thing I would add to it then, if it were short term, and they were responding to a request for three-year data, it sounds like not much of it was three-year data, most of it is one-year data in point of fact, as we understood it today, therefore, that is another issue that comes up that it really isn't even three-year data at this point in time.

DR. WHALEN:  Dr. Witten.

DR. WITTEN:  Perhaps I should clarify the question before you all go around and answer it, and that is that yes, the sponsor was asked, what the sponsor did was they proposed a 10-year study and proposed to come in to FDA with two-year data, which is what they have done.

But our question to you in the end is going to be is there enough safety information to provide reasonable assurance of safety and effectiveness.  So, what is in their application is their study, their preclinical testing, which we have heard about the limitations of, the information from the literature giving us some idea of the long-term performance of these implants as they have been reported in the literature, not specifically to this implant, so we are going to be asking you whether the information is adequate to determine reasonable assurances, to provide reasonable assurance of safety and effectiveness.

So, I think you are going to also need to tell us whether this short-term study, coupled with the other information, will be enough to provide reasonable assurance of safety for this product.

DR. WHALEN:  In view of that elucidation, Dr. Boulware, Dr. McGrath, do either of you have any further comment or can we move along?

DR. McGRATH:  No.

DR. WHALEN:  Dr. Leitch.

DR. LEITCH:  Well, to say the safety of an asymptomatic rupture, by the aspect that it's asymptomatic, it is not bothering the patient at the moment, so in that point, I think is where they are in this study because it is early on, that those ruptures are not causing the patient's problems currently, so really, the question is not so much what is the safety of an asymptomatic rupture, as what is the long-term outcome of a rupture that is identified in the asymptomatic state, what is the long-term outcome of that.

Data wasn't presented to that other than the data that is being reported with respect to systemic symptoms and signs which there seems to be general consensus that a lot of these concerns that people had about connective tissue disorders are not highly associated with breast implant patients, but what other people have asked here is, well, if you cull that down to those with rupture, what is the rate of that.

So, while in the short term, I think the sponsor has presented what they should, we do need the long-term follow-up, which I think is what we are saying here, that if approval were to occur, that the follow-up data must be required and rigorous for that long term and posing to the sponsor a question that they can answer, so that they can generate this data for the long term.

DR. WHALEN:  Dr. Chang.

DR. CHANG:  It is hard to change the rules in midstream or to try to predict the future without having the information, so to answer the short-term question, rates of 2.7 percent were given for ruptures by MRI studies, and in collecting or looking at extracapsular rupture, the sponsor testifies that there have been no extracapsular, no findings grossly on clinical examination by surgeons of extracapsular gel leakage.

So, it's a qualified yes, but it begs the question of what happens to again patients with ruptures long term, if they had a rupture, it was asymptomatic, and the patient kept it in for a year and had this amorphous gel in a subpectoral plane for one year, we haven't had the opportunity to examine any data of what happens to those patients, so I think we are asked to make a conjecture without data.

DR. WHALEN:  Dr. Choti.

DR. CHOTI:  I think with respect to this question, I think we don't have adequate information regarding the safety with regards to asymptomatic rupture.  I think, if anything, the information is of concern.

I think rupture is bad, and I think asymptomatic rupture is bad, as well, even though it's not asymptomatic, I think it's likely there is a concern that is associated with higher risk of extracapsular rupture perhaps.

So, it brings up other issues in addition, well, the other point is that it is higher in the reconstruction patients perhaps, it appears to be, and why is that the case, and is it less safe in the reconstruction patient for that reason.

Also, other issues somewhat less related to safety are the screening issue, how then do we follow patients at risk for asymptomatic rupture and who pays for the costs.

So, I think that I do have concerns that the information provided brings up concerns of safety, and it is not adequately addressed.

DR. WHALEN:  Dr. Blumenstein.

DR. BLUMENSTEIN:  The number of events is too small and the follow-up is too short.

DR. CONANT:  I agree.

DR. LAWRENCE:  I would echo the same and say that it is much more critical for a women in the childbearing years.

DR. WHALEN:  Dr. Lieberman.

DR. LIEBERMAN:  I would say that I agree, there are too few events and we also have to keep in mind that only about a third of the patients were screened with MRI, so the rate of asymptomatic rupture in the whole population is really much higher than what we know.

I don't think we know in particular what I would call the shape of the curve, does it increase linearly, does it at a certain point, does the rate of rupture go way up, and that is very important in terms of thinking about monitoring these women and about recommendations about replacement after certain periods of time, and all the other issues that have been raised related to that.

We also don't know the consequences of detecting an asymptomatic rupture sooner rather than later, is it more likely to result in extracapsular silicone, and what is the importance of that,  so we don't know that either.

So, I think that we have a lot more that we need to learn.

DR. WHALEN:  Dr. Manno.

DR. MANNO:  I think I pretty much agree with everybody so far.  There is insufficient information on the short term, there is only one year's worth of data completely analyzed and part of the third year group, and nothing on the long term.  So, I think there is insufficient data.

DR. WHALEN:  Dr. Olding.

DR. OLDING:  This question really asked about the safety.  In the short term, as far as we know, since they are all intracapsular rupture, we can't really question or we don't have any evidence that there is any short-term problem with that.

We don't have the long-term evidence, but at least right now I would have to assume that the safety short term is certainly verified, and I am not certain about the long term.

DR. WHALEN:  Ms. Brown.

MS. BROWN:  May I ask the company a quick pointed question?

DR. WHALEN:  Well, since you brought it up, for this particular instance, we will.  In the question period, we would prefer not to do that.  That was during the open panel discussion, but since you have brought it up, if there is a quick question that they can give a quick answer to, go ahead.

MS. BROWN:  I hope they can answer quickly.  Can you comment on the relationship between disease and extracapsular rupture quickly?

DR. WHALEN:  If you can do that, boy, you should get an award.  If you go to the podium, please.

DR. McLAUGHLIN:  We will answer this in terms of this issue of this alleged relationship between extracapsular rupture and fibromyalgia, and I apologize for the overheads, but that is what we were told we have.

[Slide.]

The FDA summary of this issue of fibromyalgia relative risk estimates, I believe is mistaken.  The Wolfe and Anderson paper, in 1999, the correct relative risk in that paper is 1.22, not 6.1, and in our papers that we did in Sweden and Denmark, the Nyren paper and the Kjoller paper, the correct relative risk is 1.0 and 1.2.

For the Nyren paper, that is because we had a medical review and we excluded cases that weren't really fibromyalgia, and we also used the control group that we chose for the study, and in the Danish study, the Kjoller study, or as they say in Denmark, "Keeler," we used the control group that we chose for the study.

The Wolfe and Anderson, the 6.1 is in error because three of the six fibromyalgia cases in this risk ratio had their fibromyalgia before their breast implantation, so it's uninformative as to the association.

DR. WHALEN:  Thank you.

DR. LIEBERMAN:  Could I ask one quick follow-up?  Is that for all fibromyalgia and all implants?  That doesn't  particularly address the extracapsular silicone.

DR. McLAUGHLIN:  No, we have other slides.

[Slide.]

This slide shows the percentage of women with fibromyalgia by implant status in a paper by Dr. Brown and colleagues, and you see that in the intact implants, it is about 15 percent had fibromyalgia.  Ironically, in the intracapsular rupture, the lowest rate was to be found, which is the 8 percent, and then in the extracapsular rupture, they have a 25 percent rate of self-reported fibromyalgia.

DR. WHALEN:  What was the length of follow-up in that paper?

DR. McLAUGHLIN:  This is a cross-sectional survey, so they weren't following anybody, they just interviewed them.

DR. BROOK:  Some of the cases might have been diagnosed before implantation in this study, as well.

DR. WHALEN:  How long were the implants in?

DR. McLAUGHLIN:  Sixteen years, the average was about 16 years.

So that what happened, if you compare the extracapsular rupture rate to the intact implant rate, which I think is the more traditional epidemiologic comparison, you don't get a significant difference.

In the Brown paper, what happened was the intracapsular rate, the lowest rate, was put in with the control comparison, which then produced a statistically significant difference.

[Slide.]

We tested this hypothesis in a study, a similar study in Denmark, and this slide shows that in the intact implant, the percentage of women with fibromyalgia was less than 1 percent.  For the intracapsular rupture, it was 1.4 percent, and for the extracapsular rupture, of which we had 23, not one woman had fibromyalgia.

DR. BROOK:  The average age of the implants was 14 years.

DR. WHALEN:  The comments need to be at the microphone by the person answering the question, please.  For any question of the sponsor, we will have one person answering the question.

DR. McLAUGHLIN:  So, I, as an epidemiologist, and someone who has published over 45 papers in the field of the epidemiology of breast implants and their health outcomes, I don't think the relationship between fibromyalgia and extracapsular rupture has been established at all.

DR. WHALEN:  Ms. Brown, does that answer your question?

MS. BROWN:  Yes, thank you.

DR. WHALEN:  Thanks.  Ms. Gilbert.

MS. GILBERT:  I will hold my questions in regards to this, because I don't think there is enough information even going through what they just went through, but frankly, I don't think the--the short term, yes, one year, yes.  For long term, no, I don't think we have enough information.

DR. WHALEN:  Dr. Miller.

DR. MILLER:  I think that we don't know the rate of rupture beyond one year, but we also don't know the implications of rupture should it happen.  There certainly doesn't seem to be any health risk with the implants that can be demonstrated, so I mean it's hard to know what to say about silent rupture.

I mean it is silent, it's asymptomatic, we don't know how often it happens, we don't know the implications. It would be nice to know, but I don't know if that is a safety issue.  I don't see anybody being harmed by a silent rupture.  I mean I don't like having a ruptured implant in a patient, but I am not convinced that they are harmed by it.

So, is it safe?  It hasn't been proven unsafe.

DR. WHALEN:  Dr. Anderson.

DR. ANDERSON:  I am going to disagree in a qualified way with the majority of the panel.  I think we do have long-term data, it's not in the company's study.  We heard reference to the Mayo Clinic study, which has 7.8 years follow-up for the implant group, and in the Nurses' Study, it was 9 years, and these are not prospective studies, but there was no difference between control and intervention groups.

So, I think there is more data, it is just in the literature as opposed to what was cited here.  I agree that getting more data is better than having less.

DR. WHALEN:  Dr. Li.

DR. LI:  I will just answer from a device standpoint.  It really just echoes what everybody else said, that it appears that the rupture rate is really kind of unknown beyond one year.  It's unlikely it will get better with time, so I am a little taken aback that I don't know how we can answer the safety question when we are not quite sure how to answer the simplest question of how much they rupture.

DR. WHALEN:  Dr. Dubler.

PROF. DUBLER:  I have been a lawyer hanging out in medicine for 27 years, and it is rare that something shocks me, but I sat on the panel in '92, and that was 11 years ago.  How we could have come from 11 years ago, where we were going to collect data, to a point where we have a year's data simply boggles the mind.

I am sure from what I have heard today that it's some combination of FDA and sponsor relationship, but I simply don't understand it.

I am also shocked by a second point, and that is that clearly, the rupture of these implants is a very serious issue, and yet the company over the 11 years, the sponsor over the 11 years between the first hearings and now, has made no attempt to make this a better implant, so that it might have a lower rupture rate.

So, having said I am shocked, I feel like Claude Raines at the end of Casablanca, but given that and given that I don't think we have anything more than a year's data on asymptomatic rupture, I would say that we can't make a determination on safety.

DR. WHALEN:  Dr. Witten, in regard to Question No. 1, relating to the adequacy of information to determine safety of the product with respect to asymptomatic rupture, there is not an entirely homogeneous response to that issue.

I think that the panel pretty much uniformly feels that there is, in the short term, really measured as one plus or minus years, that there is adequacy of information to assure that it's safe at that juncture.

Going beyond that interval of time to a longer term, there is some discordancy of opinion as to whether or not there is any data to demonstrate that, but there is also some elements of thought on the panel that it remains open-ended as to whether or not if, indeed, there is a higher rupture/leakage rate longer term, what, if any, implications that has for the individual patient.

So, for the short term, it is demonstrated to be safe.  I would say for the long term, the majority opinion is that that remains a question to be answered.

Does that answer the question adequately?

DR. WITTEN:  Yes, thank you.

DR. WHALEN:  Commander Allen, Question No. 2.

CDR ALLEN:  No. 2.  Please discuss the adequacy of the literature and the preclinical testing to determine the safety of this product with respect to long-term and general health effects.

DR. WHALEN:  Thank you.

We will start with Dr. Boulware.

DR. BOULWARE:  I am actually reflecting back down at the bottom of the definition of safety that the FDA has given us to use.  In considering the literature that has been out there from the IOM, other large epidemiologic studies, I am actually very comfortable with this, that they have met the safety issue.

DR. WHALEN:  Dr. McGrath.

DR. McGRATH:  That last sentence you asked us about preclinical testing and about literature.  Preclinical testing was done according to the usual formulas, as I understand it, for ASTMS standards, so I am happy with that.

As far as the literature goes, we could go through each of these categories and discuss them, but to just lump them all together, again, we haven't heard it presented today, but certainly out there in the historical published literature, and that pulled together by other very concerned medical groups that have looked at this, the Institute of Medicine and the other mega-analyses have not suggested that there is any risks in terms of cancer risks and connective tissue disease, and so forth, so I agree with Dennis on that, safety on historical data, and yes also on the preclinical testing.

DR. WHALEN:  Dr. Leitch.

DR. LEITCH:  I would agree that with respect to systemic illness and the literature and preclinical testing that has been done, that we have good evidence for the safety.  I think again one of our main issues remains the local complication of implant rupture in the long term.

DR. WHALEN:  Dr. Chang.

DR. CHANG:  With one reference to a published literature that suggests that peri-areolar incisions may cause increased difficulty with breastfeeding, there is yes to both questions.

DR. WHALEN:  Dr. Choti.

DR. CHOTI:  I am not sure I can say that with such confidence.  I think the preclinical testing, while encouraging, I don't think reflects necessarily some aspects of the clinical safety that we have discussed regarding the gel bleed and the mechanical strength, so I am not sure it addresses all the issues about long-term safety.

As far as the literature, again,  I think that all the data is not there.  So, I still have some concerns about long-term safety based on both preclinical testing and the adequacy of the literature.

DR. WHALEN:  Dr. Blumenstein.

DR. BLUMENSTEIN:  I have been around medical studies and clinical trials long enough to know that we are always surprised.  I just don't think that the data collected from these two modes are sufficient, and I can think of a lot of reasons, for example, you can't ask laboratory animals how they feel, so I don't think this is adequate.

DR. WHALEN:  Well, you can, but they won't give you an answer.

Dr. Conant.

DR. CONANT:  I also am concerned about the laboratory, the preclinical testing for things like gel bleed and gel migration, that they really don't mimic the in-vivo circumstances, so I am not sure I have enough data either for long term.

DR. WHALEN:  Dr. Lawrence.

DR. LAWRENCE:  I think as a general comment that there is probably not enough information to answer all these questions, some are more significant than others, but I am concerned, as was mentioned earlier, about some of the animal information and the way reproductive and teratogenic effects were assessed, so I don't think we know that answer.

DR. WHALEN:  Dr. Lieberman.

DR. LIEBERMAN:  I think there are some things about long-term effects that we don't know.  I am not sure that all of them are knowable or can be known in terms of considering approval.

There are some suggestions about some kinds of cancer, but again, I think we know it is probably not very big, connective tissue disease, maybe there is a subgroup where there is an effect, but the overall literature is very reassuring.

I think we don't know a lot about gel migration and extracapsular rupture, so there are clearly still some questions that we don't have all the answers to.

DR. WHALEN:  Dr. Manno.

DR. MANNO:  I pretty much agree with what has been said, however, I have some concern, and it may be my lack of understanding of the historical, the literature, that has been given here, but it is my impression, if you are going to use historical data in terms of controls, I got the feeling that we are dealing with non-Inamed specific implants, and we are comparing apples and oranges because the uniqueness of the Inamed device is the construction of the outer or the components of the shell, and that the gel is different than has been used before.

So, we are looking at a whole new ballgame where that is concerned.  I could stand to be corrected if I am wrong in my interpretation.

DR. WHALEN:  Could I just follow up and maybe I wasn't listening to you right, so I will apologize upfront,  but in regards specifically to the literature end of it, do you feel that is adequate or not in answering the long-term question?

DR. MANNO:  It is not specific towards the product that we are discussing, we don't have enough information.

DR. WHALEN:  Thank you.

Dr. Olding.

DR. OLDING:  Again, we are talking comfort level here, and I would ask those members of the panel who are not comfortable with the adequacy of the literature, what would, in their own minds, make them comfortable with the literature.  I think it's an individual thing and it is difficult when you try to look at a lot of different articles that are not, as other panel members have said, are specific, but when there appears to be an overwhelming support or refuting of some particular general item to be discussed, then, I think one has to pay attention to that, therefore, I am, by and large, comfortable with the adequacy of the literature and also with the preclinical testing.

DR. WHALEN:  Ms. Brown.

MS. BROWN:  With respect to the literature, I agree with Dr. Olding, and with respect to the preclinical testing, the company, I believe has complied with the pre-agreed-upon standards that FDA had issued, so I believe the preclinical data, according to that standard, is adequate.

DR. WHALEN:  Ms. Gilbert.

MS. GILBERT:  I agree, but I disagree.  I think there is a lot of information, and I think that this Question 2 outlines a lot of other areas that we don't have answers to, so I think we need more information.

DR. WHALEN:  In terms of the literature part or the preclinical part?

MS. GILBERT:  In terms of the literature part, yes.

DR. WHALEN:  Thank you.

Dr. Miller.

DR. MILLER:  I think the literature is pretty convincing that there is no systemic health long-term problems, and the preclinical testing appears to be in accordance with what was required for them to do.  Even though it does leave some question open in terms of whether the tests--if the tests that are done don't mirror what happens in vivo, then, perhaps the tests that are done need to be modified, but that is not what has been required of the company.

So, I think that both preclinical testing and the literature are consistent with safety.

DR. WHALEN:  I am sorry, repeat that.

DR. MILLER:  That they would be safe based on those criteria.

DR. WHALEN:  Dr. Witten wants to say something.

DR. WITTEN:  Yes, I just want to comment, maybe just because I am from FDA that it bothers me, the use of the terms "required" and "agreement," since neither of them are really the case.  We have guidance, and that means we provide our recommendations to sponsors, and we have discussions with them about what our best recommendations are about testing to address the issues in that application, but it is not requirements, and "agreement" has a specific meaning to us, where we have an agreement that if they do this, that will get them their approval.

Maybe it's just a minor technical FDA point, but I just wanted to clarify our terminology that we have provided guidance.

MS. BROWN:  And I apologize for mischaracterizing that.

DR. WHALEN:  Dr. Anderson.

DR. ANDERSON:  I don't see that additional laboratory animal studies or studies of the material are going to answer our questions.  The questions are clinical ones.  These are already being put in women around the world, and so I think we need to stay in the clinical realm from here, and I think the literature is adequate, as well, so I am satisfied with both the literature and the preclinical testing.

DR. WHALEN:  Dr. Li.

DR. LI:  I believe the preclinical testing that was done was necessary, but not sufficient to predict the safety of this device given that we don't know the mechanism of the rupture.  So, they are necessary in the fact that they followed the guidance documents, but as we now have discussed, the guidance document does not necessarily allow you to predict clinical behavior.

So, I don't think that the preclinical testing actually addresses the safety issue of the product.

DR. WHALEN:  Dr. Dubler.

PROF. DUBLER:  The large epidemiological analyses seemed to me to answer some of the questions that have been posed over the last years about association with connective tissue disease.  I am concerned that we don't know what the chilling effect of implants might be on a woman's decision to seek mammography, and there have been some issues raised about the ability to interpret the results, number one, and to engage in more focused interventions.

So, with some discomfort, I would say yes, the literature is adequate.  I have no ability to comment at all on the preclinical studies.

DR. WHALEN:  Thank you.

Dr. Newburger.

DR. NEWBURGER:  I believe the sponsor has fulfilled answering Question 2 and that the preclinical and literature search, testing and search, are adequate.

DR. WHALEN:  Thank you.

Dr. Witten, in regard to the second question which addresses the adequacy of the literature and preclinical testing in determining safety of the product with respect to long term and general health effects, there again is not a unanimous opinion of the panel.

About half of the panel in regard to the preclinical testing, perhaps more the clinicians than others, reflective of the earlier discussion of the perhaps non-linkage of in vitro and in vivo data, remain somewhat hesitant to say that that would directly correlate to evidencing safety, but others are reassured by it.

In the same light, there are many who feel that the literature does adequately address some of the long-term complications, if not most of the long-term complications, but others remain hesitant to commit to that being demonstrated.

DR. WITTEN:  Thank you.

DR. WHALEN:  Question No. 3.

CDR ALLEN:  Given these data shown in the bullets above and that the augmentation patient generally has breast implant surgery at a younger age which includes childbearing years compared to other indications, is there reasonable assurance that the device is safe for augmentation patients?

DR. WHALEN:  We begin with Dr. McGrath.

DR. McGRATH:  Since you gave us three bullet points, I am going to go ahead and break it into the three bullets points as I comment.

The first one, is there reasonable assurance that it is safe with regard to the local complications as have been reported in the studies, today, I am going to say no to that,  and I am basing that on the fact that there is a 20 percent reoperation over three years and that 70 percent of those reoperations are for complications, not for patient choice.  Therefore, I think there are issues here about safety.

No. 2.  With regard to the asymptomatic/silent rupture information, do we have reasonable assurance that the device is safe based on the fact that approximately 30 percent of the patients, and so forth, have data, we have sort of addressed this already with the first question.

I think the answer from many of us was that we don't know.  We don't know enough information with that snapshot in time.  We don't yet really understand the prevalence of asymptomatic rupture over time, the best way and most cost effective way to detect it, and the implications of this in young people.

The third bullet point, is there reasonable assurance the device is safe based on published historical literature and the animal data to address the long-term effects, to that I would say yes, these studies are large. Someone mentioned that they are not necessarily Inamed's products, but I think that we have to remember that the incidence of these problems, these long-term problems are so small that it has to be pooled data from many different sources trying to give us the information, so I don't see that as a difficulty.

The only thing that I would focus on particularly in the young group is the issue of mammography, and we certainly have information from the long-term studies although I don't think it was drawn out here today, that mammograms with extra views, as was mentioned before, can detect changes in the breast in younger patients, and we clearly have good historical evidence that breast cancer is detected at just as early a stage in patients with breast implants.  So, I would yes to the third bullet point.

DR. WHALEN:  Dr. Leitch.

DR. LEITCH:  Well, I would just like to preface my remarks by saying, you know, we go on this Question 4, and I am not trying to answer Question 4, but my comments sort of reflect those two things.

Certainly, patients who have implants reconstruction for breast cancer can be of childbearing age, so I think this distinction of childbearing age should kind of go under both groups.

I think, as Dr. McGrath has said, we have talked over and over again about the local complication rate and reoperation issues that are related to local issues, so those issues are of concern, and certainly in an augmentation patient where she is not "required" to have the surgery to restore her to her original state as a reconstruction patient would have, her motivations to undertake augmentation in this circumstance where she has a risk of having reoperation, may not make it as valuable to her in that circumstance.

So, I think we still have questions about these local issues.  Again, on the third point of the literature about long-term health effects, I think I agree that we demonstrated that well, and just to reiterate my point, I would not separate out so much these patients with respect to younger age because we do have childbearing patients in the reconstruction age.

DR. WHALEN:  Dr. Chang.

DR. CHANG:  I see Question 3 as do the benefits to health far outweigh the probable risk, and with the caveat that the patient has adequate instructions for its use and is truly informed about the local complications, so two bullets, I say yes, and that is, first local complications, rates for this PMA are comparable to the historic local complication rates of the saline implant presented three years ago.

We don't take even a 1.2 percent rupture rate lightly, but it is within the range of a previous study with saline implants, which have passed muster.  Bullet 3, I have already stated I believe the published historical literature does address concerns about systemic disease.

Bullet No. 2 is qualified because of lack of long-term data.  So, that is the part that is inadequate.

DR. WHALEN:  Dr. Choti.

DR. CHOTI:  Well, obviously, this question is trying to tease apart the first two questions and asking the question specifically regarding augmentation.  The question is can it couch by saying younger patients, childbearing, so that perhaps we don't get brought into the moral issues of the risk/benefit regarding the indications for augmentation versus reconstruction.

So, in my mind really, regarding safety, they are equal whether it's reconstruction or augmentation, and I don't thing necessarily the younger age or childbearing should impact that much.  As I mentioned, generally, I am concerned about the local complications in either group, I am concerned about asymptomatic and silent rupture in this group, as well as reconstruction, as well as historical and preclinical data.

Perhaps the one difference is that younger patients, there are some issues regarding informed consent.  Perhaps for this indication it is somewhat more difficult and since there are risks, and the patient really needs to understand carefully the risks, one needs to think about in the younger augmentation patient, the motives as it relates to are they really understanding the risks before they take this on.

DR. WHALEN:  Dr. Blumenstein.

DR. BLUMENSTEIN:  I find myself having to think about the efficacy here and the balance between efficacy and risk, and I don't have much question that this device is efficacious in a subset of the people in which it is applied to.  It is the subset that has the risks in which it is not efficacious or at least not immediately so, and so balancing those two factors is where I am having difficulty coming up with a yes or no answer, I have to say no.

DR. WHALEN:  Dr. Conant.

DR. CONANT:  I feel very strongly that I shouldn't be separating out the indications for this procedure and treating them differently except for the duration and longevity of how long that implant may be in there, and therefore the younger patients come into question, and I am concerned again about the long-term follow-up for local complications, for silent ruptures, and that we just don't have that data particularly when maybe we will be looking at these--we could be looking at these young women for many years.

I am also concerned about the ability to truly provide informed consent without longer term follow-up, particularly to educate a young woman choosing implants when she may not have a clue about her lactational interests 10 years later, or the importance of breast cancer screening.

So, I would say no.

DR. WHALEN:  Dr. Lawrence.

DR. LAWRENCE:  I think in this situation that it is not a matter of why it was done, the risks seem to be the same for women in the same age group, however, the risk/benefit ratio varies depending on your interpretation of benefit.

I think that the literature is still lean on the long-term outcome of rupture and the long-term outcome of any offspring.

DR. WHALEN:  Dr. Lieberman.

DR. LIEBERMAN:  I have great concern about the local complications especially in a young woman who is going to have this implant for a long time and probably going to have to have, even if nothing goes wrong, repeated surgeries for replacement since they don't last forever.

On the other hand, even with that concern, given that the bar has been set and the rates are similar to saline, and given issues of choice for women, I have to say that it is probably acceptable if we can get really good informed consent.

On the second bullet, I think there is a lot we don't know about asymptomatic and silent rupture, and I am concerned about that, and I would have to say no.

For the published literature, as I said, there are probably some sub-areas and some subgroups where we don't have information, we definitely need more information, but I think that is going to need to come from large epidemiologic studies, not from the sponsor.

DR. WHALEN:  Dr. Manno.

DR. MANNO:  I think effectiveness is a non-issue here.  I think it is pretty broadly presented that I can recognize the effectiveness of the implant.  I agree with Dr. Lieberman on the first question.  I think we don't have enough information on the long-term aspects of the asymptomatic rupture.

I think other than what I have already mentioned on the third point, what I have already mentioned on what I see as problems with the offspring in the reproduction when we lack that little bit of information.

DR. WHALEN:  Dr. Olding.

DR. OLDING:  My comments pertain to Question No. 3 and No. 4, because I, like some of the other panelists, cannot separate these two out for any reason.

I believe that, to address the third bullet first because for me, it's the easiest, I believe that there has been good historical literature and animal data.

The first bullet, local complication rates, I am concerned about the local complication rates, but in the long term we haven't shown that those local complication rates, other than in and of themselves, and the problems which they create by being present at that time, don't really cause any long-term problems.

I would ask myself the same thing I asked of an earlier question of the other panelists, what would make me comfortable or happy to answer the second bullet, and I think that it answers itself.  I think that I would be much more comfortable with some more screenings.

DR. WHALEN:  Ms. Brown.

MS. BROWN:  With respect to local complications, it seemed that the largest concern was reoperation.  It is a concern we are already living with, with the saline-filled implants, so I think that if we are willing to accept the reoperation rate for saline implants, we should be willing to accept it for this.

With respect to long-term data, actually tomorrow I would be interested in clarification from the company, not tonight, tomorrow, clarification on what their long-term data collection plan is for the 10 years, because I thought I saw almost two different plan presentations, and I am very curious as to what the real plan looks like, but I think that is very important.

Then, with respect to the definition of what constitutes reasonable assurance of safety, built into that definition is when accompanied by adequate instructions for use and warnings against unsafe use, so I do think that for the augmentation patient, it is very important that they get good informed consent.

DR. WHALEN:  Ms. Gilbert.

MS. GILBERT:  I think this is a personal issue.  I am not going to go down the three bullets.  I am going to say is there reasonable assurance that the device is safe for augmentation patients, no.  I don't think that women understand, nor should they accept the fact that when they get reconstruction or augmentation surgery, that they are going to be on a merry-go-round of surgery after surgery after surgery.  I mean they need to be handed those facts upfront.

I mean the high reoperation rate speaks for itself.  That is all I can say.

DR. WHALEN:  Dr. Miller.

DR. MILLER:  I think the local complications are well characterized by the current studies.  I think that if a patient is informed of what their risks are, that those risks don't necessarily outweigh the benefits in that person's mind.  If the person will receive high benefit from the procedure and understands the risks and wants to proceed, then, I think that is a valid choice.

So, I would say yes on No. 1, I would say yes on No. 3 because of all that we have talked about, and because of the yes on the No. 3, I would say  yes on No. 2, because we don't know the implications of the rupture, so if long term there is no health effects seen with breast implants, in some ways, if they rupture, if they don't rupture, there is no long-term health effects seen.  In terms of the safety, I am not sure that it bears on the safety.

DR. WHALEN:  Dr. Anderson.

DR. ANDERSON:  FDA has provided us with four studies that address nursing problems with women with implants.  I think that in the childbearing years question, it is an informed consent issue, and should be stated explicitly in the consent form that women who have breast implants may have trouble breastfeeding in the future.

DR. WHALEN:  Dr. Li.

DR. LI:  Actually, I have nothing to add.

DR. WHALEN:  Dr. Dubler.

PROF. DUBLER:  It is wonderful to watch the focus and the cadence of a discussion change.  When we sat on this panel, some of us, and we considered the saline implants, we were challenged by a member of the public to say how could you possibly approve them given the rupture rate, on the one hand, and the local complications and resurgical intervention rate.

I was comfortable then, and I am not comfortable now.  It seems to me that the one year data that we have don't provide us sufficient information on bullet 1 and 2, and the published epidemiological studies answer bullet 3, and I would give that answer both for No. 3 and No. 4.

DR. WHALEN:  Dr. Newburger.

DR. NEWBURGER:  I agree with Dr. Dubler and without a prospective long-term study, I am uncomfortable about the impact of what will ultimately turn into probably extracapsular silicone, and without that information, and this isn't something that I saw in the literature search, without that information, I don't see how a patient can really give an informed consent.

DR. WHALEN:  Dr. Boulware.

DR. BOULWARE:  I have nothing further.

DR. WHALEN:  Before giving the synopsis to the answer to Question No. 3 to Dr. Witten, I would like to ask the panel's indulgence and advice.

With the possible exception of Dr. Anderson, it was my sense from all of your answers to Question No. 3, that that would be the answer to Question No. 4, that we are not discriminating again with the possible exception of Dr. Anderson, with some warnings about the breastfeeding population, that we are not discriminating on the question of safety between the augmentation and reconstruction population.

I open the floor to anybody who disagrees with that.

DR. CHANG:  I would change one bullet.  Rather than saying yes to two bullets on 3, it is yes to the third one because I am just too uncomfortable with the rate of 45.9 percent reoperation rate with breast reconstruction.

So, I think that is an inordinately high number, so that is my question.  That doesn't change my answer in terms of should a person needing reconstruction, should they have the option.  I believe they should have the option, but I am sure that I am convinced that that is absolutely safe.

DR. WHALEN:  Dr. Li.

DR. LI:  My only addition to Question 4, I don't disagree with what you said, Dr. Whalen, the only issue I would want to bring up with Question 4 on reconstruction, is that we don't appear to understand why the rupture rate appears to be higher in the reconstruction, so we should not be surprised if, as the study goes on, that, in fact, the increase in rupture as more and more patients are studied, it actually doesn't go up linearly between the augmentation and the reconstruction group.

DR. WHALEN:  Dr. Leitch.

DR. LEITCH:  Well, the other thing, the point Dr. Miller made, which I think we do have to really take into account, on the reoperation rate on the reconstruction patients is that those reoperations can be part o the completion of the reconstruction, and I think it was noted somewhere in our booklet that it might have been coded that removal with replacement was actually taking out the expander and placing the permanent implant at that time, so again, that's the stepwise process which the patient would be informed about in advance.

So, I think if you just look at it and say, oh, my gosh, look at all these operations, they have to be presented that it is a stepwise process in many circumstances, and that accounts, I think, for, you know, maybe 20 percent of these reoperations, so just to make that point.

DR. WHALEN:  Before going on to the next comment, let me admit my abysmal failure in my judgment because obviously, people look at these two questions differently, so I will answer No. 3 and then we will go on to No. 4.

My interpretation had been that everybody was being consistent with the first two bullet points with Question No. 1, in that there was no agreement upon that, and everybody was being consistent with the third bullet point about historical data in that there was no agreement with that, and all of your analysis for this augmentation population fell into that same line of thought, so that there was a discordancy in the degree of safety for the augmentation population.

For Question No. 3, Dr. Witten, does that satisfactorily answer your question?

DR. WITTEN:  Yes, thank you.

DR. WHALEN:  We have already had some comments about Question No. 4, and we are just going to go through everybody else who hasn't yet answered, and anybody who wants to add to what they have already said.

We will just start over at that end with Dr. Miller and work our way down the table.

DR. MILLER:  I think the reconstruction patients are already a different kind of patient, and when you talk about safety, the way it is defined here, it is really a balance between the risks and the benefits.

A reconstruction patient, many of them, those who elect to have reconstruction, the benefit is high, so you can tolerate perhaps a higher complication rate because the ratio is what we are talking about here.

I do many operations that have a 45 or 50 percent complication rate, which sounds horrendous, but when you consider the alternative to those operations, it becomes very acceptable.  We shouldn't accept a 45 percent incidence of problems with these operations, we should always strive to lower that, but that doesn't mean that they are not safe operations.  We just have to improve them, but they are still safe because the balance of risks and benefits is acceptable.

I think the reconstruction patients, we can tolerate a little higher risk of problems, and I want to reiterate the issue of the reoperation in these patients, it is difficult to group all those reoperations together as a problem, because many of these patients have revisions by their choice, some want to exchange an implant.  You are more prone to capsular contracture, and it is very visible with the very thin tissues over the implant.

I have been tempted all evening to speculate as to my theories as to why you may have larger rupture rates visible with these implants, but I will refrain from that, unless you want me to, but I won't prolong the evening.

But I would say yes on all three of the bullet points, that it is a safe procedure by the criteria of balancing the risks and the benefits, and the complications are well known and can be explained to the patient.

DR. WHALEN:  It follows logically, but just for the simplicity and in view of the hour and my current mentation, if we just focus on is it safe for the reconstruction population, because that subsumes the other bullet points, and it is quite clear you are saying yes in that regard.

DR. MILLER:  Yes.

DR. WHALEN:  Dr. Anderson, anything further?

DR. ANDERSON:  Can you come back to me?

DR. WHALEN:  Okay.  Dr. Li, anything further?

DR. LI:  No, nothing more.

DR. WHALEN:  Dr. Dubler.

PROF. DUBLER:  Nothing further.

DR. WHALEN:  Dr. Newburger.

DR. NEWBURGER:  No questions.

DR. WHALEN:  Dr. Boulware.

DR. BOULWARE:  Yes, yes, yes.

DR. WHALEN:  I will just take that to be stuttering since I just wanted one yes.

[Laughter.]

DR. WHALEN:  Dr. McGrath.

DR. McGRATH:  The same as for augmentation, just one issue, and that is that since you don't need to do mammography in the reconstruction group, I think that that eliminates some concerns that one might have about visibility of the breast tissue and also the question of whether mammography can injure the implant.

DR. WHALEN:  Anything further, Dr. Leitch?

DR. LEITCH:  No.

DR. WHALEN:  Or Dr. Chang?

DR. CHANG:  No.

DR. WHALEN:  Dr. Choti.

DR. CHOTI:  Just one point about the two different groups.  I personally think that they should be, you know, the approval is comparable, yes, the risk/benefit is different between the two different groups, but that should be the informed consent aspect, the patient should make that decision based on their risk/benefit, but for approval, I think the safety profile should be comparable for both the reconstruction and the augmentation group, and I still have concerns in this group, as well.

DR. WHALEN:  Dr. Blumenstein.

DR. BLUMENSTEIN:  No comment.

DR. WHALEN:  Dr. Conant.

DR. CONANT:  I have the same response to 4 as 3, but I just wanted to ask for some guidance.  Through Question 3, I heard a couple of people comment, saying that, well, we think it's okay for saline implants, so because the complication rate is no more, then, it must be okay.

I thought we weren't supposed to be considering that as a comparison, but looking at this independently.  Can you provide some guidance?

DR. WHALEN:  I will defer to Dr. Witten.

DR. WITTEN:  Well, I think that really relates to that product is approved and what the clinical experience is with it, and if that helps people put it in perspective, then, I think that's your prerogative, if it helps you put it in perspective to give an answer, then, that is up to you.

DR. CONANT:  Thanks.

DR. WHALEN:  Dr. Lawrence.

DR. LAWRENCE:  No additional comment.

DR. WHALEN:  Dr. Lieberman.

DR. LIEBERMAN:  Nothing further.

DR. WHALEN:  Dr. Manno.

DR. MANNO:  Nothing further.

DR. WHALEN:  Dr. Olding.

DR. OLDING:  Nothing further.

DR. WHALEN:  Ms. Brown.

MS. BROWN:  I did have a comment just with respect to informed consent for this patient population.  If there is any way to address the issue of reoperation as a function of the process as opposed to complications, I think that would be helpful, so if a patient could know that X percent were due to the fact that a tissue expander is being removed as opposed to there has been a problem with the implant, if hat makes a sense.

So, I think that would be helpful in the informed consent decision.

DR. WHALEN:  Ms. Gilbert.

MS. GILBERT:  I know it's late, but I have concerns about this.  Do I think it's safe?  Yes, for a short period of time.  Do I think patients are being educated?  Absolutely not, no, I don't.

I myself went through this.  I didn't get all the facts.  I didn't have all the information in front of me.  I don't think that we are giving women the information they need to make an informed decision.

It may be safe for the short period of time, but long term, no.

DR. WHALEN:  Dr. Witten, in regard to the question of there being reasonable assurance that the device is safe for reconstruction and revision patients, I would state that there is a larger element of support for safety among panel members than there was for augmentation patients, although there remain some significant elements of reservation based upon the same tenets of the data that have been previously discussed.

DR. WITTEN:  Thank you.

DR. WHALEN:  Question No. 5.

CDR ALLEN:  Based on the patient satisfaction and health status/quality of life data, has Inamed adequately demonstrated reasonable assurance of effectiveness of the implants for each of the augmentation, reconstruction, and revision indications?

DR. WHALEN:  Starting with Dr. Leitch.

DR. LEITCH:  Yes, I would agree that they have.

DR. WHALEN:  Dr. Chang.

DR. CHANG:  Yes.

DR. CHOTI:  Dr. Choti?

DR. CHOTI:  Yes, I do.

DR. WHALEN:  Dr. Blumenstein.

DR. BLUMENSTEIN:  Yes.

DR. CONANT:  Yes.

DR. WHALEN:  Dr. Lawrence.

DR. LAWRENCE:  Yes.

DR. WHALEN:  Dr. Lieberman.

DR. LIEBERMAN:  Yes.

DR. WHALEN:  Dr. Manno.

DR. MANNO:  Yes.

DR. WHALEN:  Dr. Olding.

DR. OLDING:  Yes.

DR. WHALEN:  Ms. Brown.

MS. BROWN:  Yes.

DR. WHALEN:  Ms. Gilbert.

MS. GILBERT:  Yes.

DR. WHALEN:  Dr. Miller.

DR. MILLER:  Yes.

DR. WHALEN:  Dr. Anderson.

DR. ANDERSON:  Yes.

DR. WHALEN:  Dr. Li.

DR. LI:  Yes.

DR. WHALEN:  Dr. Dubler.

PROF. DUBLER:  Yes.

DR. WHALEN:  Dr. Newburger.

DR. NEWBURGER:  Yes.

DR. WHALEN:  Dr. Boulware.

DR. BOULWARE:  Yes.

DR. WHALEN:  Dr. McGrath.

DR. McGRATH:  Yes.  Dr. Boulware and I are laughing because we are looking at page 27 and page 63, which are the summary of the selected health status/quality of life measures, and for every parameter they are statistically different in the pre- to two-year post-implant score, but every single one, they are worse.  Am I not mistaken, are these not the quality of life information that we are talking about here?  So, I can't say that I think that it has been shown in the data that has been submitted by Inamed.

On the other hand, I think there is historical literature out there that hasn't been discussed here or presented, but definitely supports that there is a positive benefit to patients from having these implants for both augmentation and reconstruction.

DR. WHALEN:  Dr. Witten, the overwhelming preponderance of the panel, which may be reflective of the fact that it's 25 minutes to midnight, is that efficacy has been established for this with one salient observation by a panel member that perhaps some of the data does not overwhelmingly demonstrate that, but some other elements of data do.

DR. WITTEN:  Thank you.

DR. WHALEN:  Next.

CDR ALLEN:  No. 6.  Given the information in Question 1, please address the following with respect to the labeling for the device:

First, provide your recommendations for the frequency and method of screening for asymptomatic rupture, given that prospective screening for asymptomatic rupture is not currently routinely performed..

Second, provide your recommendations for the necessity of explantation of asymptomatic implant ruptures.

DR. WHALEN:  Dr. Chang.

DR. CHANG:  I am going to take my cue from Dr. Spear's comments during the sponsor's presentation, that monitoring every two to three years would be very reasonable.

Is it necessary to explant a known rupture?  In order to avoid potential for extracapsular migration and just the thought of the gel being outside of the capsule, and being slimed, I think yes.

DR. WHALEN:  Just fill us in with regard to the method of the screening.  You have stated the frequency, but  I don't believe you stated the method.

DR. CHANG:  Screening by follow-up by the physician.

DR. WHALEN:  Dr. Choti.

DR. CHOTI:  This is an extremely hard question to answer.  I mean we have not been shown any data to provide you any guidelines regarding these recommendations.  We have heard some testimony.  Some said that patients are examined every six months, one comment made annual screening, who pays for it, and so forth.

I think, although there is no evidence, I think it is reasonable to recommend annual screening, and as far as the recommendation for explant, again, we don't have any good evidence, but I am concerned that rupture is bad and probably even if it's asymptomatic, it should come out.

DR. WHALEN:  Dr. Blumenstein.

DR. BLUMENSTEIN:  This is an opportunity to do a randomized clinical trial embedded in this study where you would give the woman who has discovered to have an asymptomatic rupture the chance to be randomized to have it explanted or not.

Explantation requires additional procedures, whereas, not explanting it avoids those procedures.  It seems like it might be a reasonable choice and a reasonable randomization to do, assuming that there is going to be an increase in the number of asymptomatic ruptures discovered over the next few years as this study progresses, I think it is possible that there might even be enough sample size to do reasonable statistical testing.

Even if it is not, it could be a descriptive study with comparable groups.  I would also, at the same time, increase the surveillance frequency to six months.  To me, as I inspect all of the things that bother me the most about what I have been talking about tonight, and answering these questions, and so forth, it's these asymptomatic ruptures that bother me the most and what the significance of them are.

So, I feel like the opportunity to do an embedded randomized clinical trial of the consequences of explantation versus not, is going to give us an awful lot of data.

DR. WHALEN:  Dr. Conant?

DR. CONANT:  This really is a very tough one, being an imager.  To Bullet No. 1, I am not sure we have the data from Inamed with this implant, or for these implants.  However, if we look at other data that shows approximately 50 percent rupture at a ten-year point, in a perfect world, you would want these women to have some sort of screening which I think, at this point in time, I think is M.R.  But I can't see requiring--possibly recommending, that.  It is very expensive.  I am not sure that insurance will pay for that, so I have real problems with that and I don't have a good answer to it.

In terms of recommendations, I disagree about a randomized trial because women are not going to let you randomize them with this and I am very sorry, but I cannot imagine that.  However, anecdotally from my experience, I have seen quite a few women who I have, on many sequential years told them, your implants are ruptured.  There is free silicone.  And they do not want to do anything about.  And it stays very static for a very long time.

However, in those women, it is never going to get better and it is only going to get possibly worse.  It may stay the same but it is complex.  It is scarring to remove it, often and these women need to be told about their status.

I think the right thing to do is recommend explantation.  However, I think many women won't do it because if it looks good, it doesn't bother me, I am not going to do anything about it.  But I think recommending is correct, explantation.  Complex answer.  Sorry.

DR. WHALEN:  Dr. Lawrence?

DR. LAWRENCE:  If we had an ideal way to screen that was relatively inexpensive we could really make a more intelligent response to this question although, with mammographers becoming more efficient, perhaps, in the near future, we will be able to do it with ultrasound or something a little less expensive than MRI.

It certainly is appropriate to counsel any woman with an implant as to the risk of rupture and the understanding of that.  If a randomized controlled study is going to be done, we don't want to spoil its possibility of getting through the research committees by saying we know that everybody needs to have it removed.

DR. WHALEN:  Dr. Lieberman

DR. LIEBERMAN:  I think this is the hardest question.  I have to agree with Dr. Blumenstein.  This is the place where I have been most bothered and the thing where we have the least adequate information.  But, given that we know that eventually these "wear out" is the term that is being used, I think it would be prudent to recommend screening probably on a yearly basis at least until we have some more information and we can gauge things and a few of those recommendations could be changed.

We also don't know the significance of an asymptomatic rupture although, on the face of it, it would seem logical to remove it and replace it.  But it is very difficult with the limited information we have to make good  recommendations.

DR. WHALEN:  Dr. Manno?

DR. WITTEN:  Excuse me.  Before you go on.  It is Dr. Witten.  Could you just also state the method of assessment that you would recommend for this screening?

DR. LIEBERMAN:  Yes.  I'm sorry.  I guess MRI is the standard right now and I think that is what we have to do since that is the only way to pick up asymptomatic ruptures.

DR. WHALEN:  Dr. Manno?

DR. MANNO:  I would think that MRI screen should be done at a minimum of one, three, five, and it definitely should be done at ten, because we don't know the real life of these things and when the trouble will be developing.  That would be my response for that first question there.  I think the--I am having thought problems here.  I don't know if I am having an autoimmune response with loss of memory, or it is the lateness of the hour.  It probably is a little bit of both.

At any rate, I think that, in terms of the necessity for expantation, I think that it is interesting to me that it was pointed out that there are patients that, in spite of knowing they have a rupture, won't have the implant removed in some instances.

I think there has to be good information and education for the patient.  There is not a whole lot you can do if they won't submit to having them removed.  But they do have to have the information provided that some of the complications that they may develop won't be seen for a long period of time, a matter of years.

So I am kind of ambivalent whether there should be just automatic removal of the implant.  I guess I am just kind of in the middle of the road on that and it is a little out of my area.

DR. WHALEN:  Dr. Olding?

DR. OLDING:  It sounds like equipoise to me.  I will answer the second part of this first because it is the easiest of the two and I can maybe respond to it, that is provide my recommendation for the necessity of explanation of the asymptomatic patient.

We know that there are--at least in the studies presented, there have been a relatively small number of extracapsular rupture and a lack of long-term side effects relating to that.  So, therefore, for me, this one becomes simply a matter of choice, a woman's choice.  We present the data as we know it.  We give them the alternative whether or not to.

So, to me, this is really toward the second part of the--

DR. WHALEN:  Could I just ask you, Dr. Olding, is it safe to, then, say that, as regards to the specific question of its necessity, you are saying no.

DR. OLDING:  Correct.  It is not an absolute necessity.  The first one, I don't know if I can really answer that.  I would have liked to have had some additional guidance from the sponsor's studies.  I am not sure we have that, but what we do know is that the rate of rupture increases with time.  So, whatever I would recommend, or we would recommend, or whomever would recommend, for the first time period, I would continue to recommend that same methodology.

I stated before, I think it is very difficult for me, as a clinician, to determine whether or not a patient is ruptured from my examination.  So I think it has to be more than simply my examination, but I don't know how much more.

DR. WHALEN:  Ms. Brown?

MS. BROWN:  I have no comment.

DR. WHALEN:  Ms. Gilbert?

MS. GILBERT:  I think annual exams by a woman's physician or plastic surgeon should definitely--I mean, you should definitely maintain your relationship with that provider, especially if you are not seeing a specialist in this field that can tell you if there is something going on that you need to be alerted to.

I also think MRI is something that a lot of women should be aware of as an option, three, five, seven years because we don't know the rupture time line there.  The second question, explantation, yes, because I think that is something that should happen.  I don't think that we have the confidence or the knowledge to make an informed decision on whether or not there are long-term side effects.  We don't have that information.

DR. WHALEN:  Dr. Miller?

DR. MILLER:  I think a yearly or even every-other-year physical exam is good for follow up followed by an MRI if you suspect a rupture on physical exam.  And if I confirm a ruptured implant in a patient, I generally remove it.  I insist that we remove it, but that is probably an emotional thing because we don't know the implications of a ruptured implant.

That is the problem.  There is no long-term health problems associated with the implants that we can detect.  And so, if the implant ruptures, it is not associated with a long-term problem.  We are not sure what it is associated with, if anything, but it is more pleasing to get it out.

DR. WHALEN:  Dr. Anderson?

DR. ANDERSON:  I feel more strongly than that, that if we know of a rupture, we should recommend that it be removed because an asymptomatic rupture can progress to becoming a symptomatic rupture which is the problem that we are trying to get the handle on.

So, if we remain neutral, well, you could take it or not take it out, well, who in the world would take it out.  Why would I do that if the recommendation is not being made.  In my car, it is recommended I change the oil every 5,000 miles.  It is actually my decision about whether I act on that or not, but that is their recommendation.

In terms of the screening, how do we screen for this, I think we have to defer to the plastic surgeons about how that screening is done.  I think suggesting MRI is unrealistic.  It is way too expensive a test for this.  But it is prudent to suggest breast exams at one-to-two-year intervals because that is what we recommend for routine breast-cancer.  So I think that we should stay with breast-examination recommendations that are the same as what we do for women without implants which is clinical exam every one to two years.  It depends on which source.

DR. WHALEN:  Dr. Li?

DR. LI:  I will defer to my surgical colleagues.

DR. WHALEN:  Dr. Dubler?

PROF. DUBLER:  Really, I have to defer to my clinical colleagues to set what would be the appropriate medical approach and procedure.  But having said that, it implicates immediately what we say to patients and it underscores the lack of our data at this point to permit us to craft that message.  So I see a message that says if you sign up for augmentation or reconstruction, it really does mean a series of surgeries, not one, with the danger of rupture and if it is a silent rupture, we recommend strongly that it be removed.

If you are really a prudent patient, you will get an MRI.  The notion that the cost, in terms of public policy, would affect the clinical standard is a difficult one for me.  So I think that No. 6 is difficult because it raises some hard clinical issues on which we don't have data, but it is additionally difficult because it demonstrates how very difficult it is going to be to craft a message that is specific enough and raises sufficient issues to alert a women to the dangers but not so overwhelming that people simply stop listening.

DR. WHALEN:  Dr. Newberger?

DR. NEWBERGER:  I think it becomes even more important now that the long-term data isn't available in terms of reactions to implant rupture, extracapsular types of rupture.  In terms of monitoring these patients, I think they should be seen on a yearly basis by their physicians.  MRIs are certainly the gold standard.  But, if you have an MRI at one year, at three years, and so on and so forth, I calculate with the current rate of augmentation and reconstructive procedures that, four years from now, there will be over 1 million MRIs done just for the care and feeding of these breast implants.  And that really does seem overwhelming to our healthcare system.

So I would be content with just MRI indicated if the physician feels there is evidence for rupture.

DR. WHALEN:  Explantation?

DR. NEWBERGER:  Again, there isn't data to guide us.  It does get worse over time if there is a rupture.  I have certainly seen that, over time, silicone gel will go in some unusual locations.  I would recommend that, as soon as an asymptomatic rupture is found, that the implant be removed.

DR. WHALEN:  Dr. Boulware?

DR. BOULWARE:  I am not very confident in the sensitivity of MRI in terms of picking up asymptomatic rupture.  I am also not very confident that asymptomatic rupture really progresses on or that we have data that progresses on to be symptomatic.  I would have no recommendation for screening for this other than the usual postoperative care and would not recommend explantation at this point for asymptomatic rupture.

DR. WHALEN:  Dr. McGrath?

DR. McGRATH:  Everybody else has said it better than I but, at this point, I cannot answer Part A.  I think any recommendation for a diagnostic test has to be data driven and I just simply don't have longitudinal data about the incidence of rupture to tell me how significant a problem this is and how it should be monitored.

In answer to the questions from some at this end of the table about leaving it up to the clinicians and physical examination, by definition, then, asymptomatic rupture would be one that neither you nor the patient would diagnose by physical exam and that these are ones that are being picked up now by a diagnostic test.

I would be hard-pressed to advise a patient without any information to spend her cash for this and I certainly wouldn't want to have to be put in the position of going to the health-insurance industry or to Medicare to explain why they should pay for MRIs for patients when we don't have any data about the incidence of asymptomatic rupture and so forth over the long term.

As far as B goes, I would say yes.

DR. WHALEN:  Dr. Leitch?

DR. LEITCH:  Well, this is the reason the MRI study needs to be completed is that we would be able, I think, to provide an answer for one with respect to that mode of screening for rupture.  So I think it would be great for that portion of the study to be completed out to the ten-year point.  Of course, this one- and three-year business, that is where the risk of rupture is lowest.  But the more years out you get, the higher risk there is so the greater yield you would have from doing those studies.

But, before that is finished, then I would suggest the annual exam by the physician.  If the breast is preserved--in other words, if it is an augmentation patient or a symmetry procedure in a patient who has breast cancer, that they should have mammography.  If they have breast cancer, they should be having mammography annually.

If they are without breast cancer but 40 years of age, they should have mammography annually.  If the mammography gives a hint that there could be an implant problem, which sometimes there can be even without extracapsular extension, then that could prompt the clinician to recommend further screening interventions to assess that probability.

Then, with respect to the issue of identified implant rupture that is asymptomatic, I generally would recommend removal of that implant if the patient's health status is otherwise good.  I mean, you are talking about the person having surgery with its attendant risk as well, so, you know, you have to put that in the context of their general health conditions as to whether, again, that risk/benefit ratio is there.

The primary reason would be to prevent the local complications of an extracapsular extension of that rupture and the difficulties it could pose for the patient in the future.

DR. WHALEN:  Dr. Witten, in regards to Question 6, with the first part being recommendations for frequency and method of screening for asymptomatic rupture, there was not a strong consensus of the panel.  I would say that the majority favor simple physical examination and the majority, I had a head count of seven recommended that that be on an annual basis.

However, that being said, there were one or two salient observations being made that it is certainly not at all established how sensitive physical examination alone would be in detecting asymptomatic rupture.

As far as other methods being utilized, the most preponderant one discussed was that of MRI.  Concerns were raised about its cost, some concern raised about its particular sensitivity in detecting this and further concern raised about the very significance of even detecting what that sort of finding might mean if, indeed, it is picked up.

There is, therefore, perhaps, less consensus about this first arm of this question than any other question that we have thus far dealt with tonight.

The second part was the recommendation for the necessity of explantation of asymptomatic ruptures and, by my count, I found seven panel members saying they should be removed, three saying they should not and three being rather on the fence about that.  So the preponderance was that they should be removed but, again, there were some reasonable questions raised as to what the significance of such ruptures are and I think there were some very good issues raised about how you would convince somebody with an asymptomatic rupture to undertake the explantation to begin with.

Does that answer your question?

DR. WITTEN:  Yes.

DR. WHALEN:  Even though we are at midnight, I would prefer to go on to Question 7 since we are in the question mode, since that is our final question.  Commander Allen?

CDR ALLEN:  With regard to the proposed postapproval study plan, first please comment on the method of data collection, the mailed survey, from the six- to ten-year time points and, second, please describe any other specific endpoints which should be captured as part of the postapproval study.

DR. WHALEN:  Dr. Choti?

DR. CHOTI:  Well, I think that follow up longer than ten years should be recommended, perhaps by examination and/or whatever screening methods we recommend.  As far as the specific endpoints, I mean, I think where guidelines really aren't based on sufficient data, it is still not clear with rupture rates exceeding 70 percent at 20 years, should there be a time in which implants should be recommended to removed regardless of rupture just because--it is still unclear.

But I think if the consensus is that they should stay until which time there is evidence of rupture, which it sounds like is probably reasonable based on the data, then I think guidelines should be removed if there is any sign of rupture and routine explant in that situation and routine evaluation, perhaps, with examination.

DR. WHALEN:  Dr. Blumenstein?

DR. BLUMENSTEIN:  I don't have much to add.  I think that active follow up is going to be better with respect to compliance than a mailed follow up.  I just wanted to say one more thing.  It is a shame that, with all the device experience in Europe, that we don't already have the data that addresses these issues.

DR. WHALEN:  Dr. Conant?

DR. CONANT:  I also think the follow up needs to be extended longer than five years and I think direct follow up is important.  I would like to see the completion of the MR study at least out to ten years.  In terms of other specific endpoints, I would like to have the population of ruptured cases confirmed, ruptured cases better evaluated as to intra- and extracapsular and potential systemic problems.

And then I would also like to know farther out the increase or the additional surgical procedures these women undergo to try to figure out what that curve is, whether it is exponential or linear or whatever.

Thanks.

DR. WHALEN:  Dr. Lawrence.

DR. LAWRENCE:  I don't feel that a mail-in kind of questionnaire is adequate for the kind of information that we wish to have available for the full ten years.  So I would vote for continuing the process as it is now and looking at endpoints such as rupture and need for additional surgery and, of course, for those women who bear children with an implant, that we have absolute follow up on those children.

DR. WHALEN:  Dr. Lieberman?

DR. LIEBERMAN:  I agree.  I think that a full ten years needs to be done with in-person follow up and also to continue the MRI screening because asymptomatic rupture seems to be the place where all of us are having the most difficulty in terms of getting a handle on its significance.

I think also following up the history of the reoperations, the long-term health effects in the women and, if possible, their children would be helpful.

DR. WHALEN:  Dr. Manno.

DR. MANNO:  I am not comfortable at all with the mailed-survey approach.  I think it should be face-to-face with the physician approach accompanied by the MRIs until we get a better handle on the long-term stuff.

In terms of other endpoints, I would suggest that, if this is approved that we would be trying to capture some information on these interim visits that will continue, that, in terms of any pregnancy since a previous visit and the outcome of the pregnancy in terms of live births, total number of children, any complications with the pregnancy or the delivery and any health or developmental problems of the child.  I think this should be added to the questions asked of these people.  This is just with the child-bearing age.

Oh, one other thing that I forgot to bring up earlier that might be of value in there is ask if there are any chemical contraceptives used in these people in the child-bearing age.

DR. WHALEN:  Dr. Olding?

DR. OLDING:  Like the rest of the panel, I agree that you need to have more than a mailed survey and I think it should be examination by the treating physician through the ten-year period.  I have nothing to add about B.

DR. WHALEN:  Ms. Brown?

MS. BROWN:  I think the ten-year period is appropriate.  When I look at the list of complications that have been reported by the company to date, a lot of these--actually, all of them appear to me to be complications that the patient could self-report.  So I think that a survey or a phone questioning could be appropriate, maybe with a physician visit somewhere at the end of that ten-year period might be appropriate.

DR. WHALEN:  Ms. Gilbert.

MS. GILBERT:  Well, I think the ten-year period is something that I would like to stress but, also, when it comes to data collection, can we be creative about this?  I think that the patients really need to take an active role and I don't think a mail-out survey really does that.

I mean, what is the return rate?  Why aren't there more women participating in this?  What does it take and why aren't more companies investing in this?  It really restores the confidence of the patient in regards to these devices.  So I would definitely encourage a stronger approach when it comes to getting more women actively involved and understanding their role in research.

Also, endpoints; just the ten-year.  I really think that it is important that women--that we learn more about this ten years.

DR. WHALEN:  Dr. Miller?

DR. MILLER:  I think I agree with the comments that were made about the need for ten-year data and for physical visits and physical examination.  I would like to see the current stage just continued on through ten years with the patients that are entered already.

I like the idea of the controlled trial regarding what to do with the asymptomatic ruptures.  I think that might be something worth considering.

DR. WHALEN:  Dr. Anderson?

DR. ANDERSON:  Agree with need for ten-year follow up, agree with concern about mail data collection.

DR. WHALEN:  Dr. Li?

DR. LI:  I would like to see long-term follow up as well out to the ten-year period, but I would definitely like to be able to answer the questions we seem to be asking for a decade now, what is the rupture rate, silent or otherwise, and what are the clinical consequences of those rupture rates.  I think it would be a shame to collect data from the ten years and not be able to answer those question.

The other thing I would add about other specific endpoints is, despite my criticism of the retrieval programs that I provided earlier, I think, in the long run, the retrievals will probably be your best indication to the mechanisms of rupture.  So, to continue to gather as many retrievals as you can, but I would probably take a more biased, more generalized approach, failure-approach, method to analyzing those retrievals.

DR. WHALEN:  Dr. Dubler?

PROF. DUBLER:  Yes.  I agree with the last two comments.  My problem is one of the framing of the question.  Inamed provided a brief description of their postapproval study plan.  In the best of all worlds, which this is not, we would have ten years worth of data to support the panel's comments and evaluation of the implant process and product.

But we don't.  And I just don't know if the one-year, basically one-year, data we have is sufficient to get us to an approval that gives us an postapproval study plan.  It might be a preapproval study plan.

DR. WITTEN:  Yes.  I was going to say that the question could be taken more generically to be read as what type of study, longer term, do you think would be appropriate for these implants.  So you could just answer that question, then.

PROF. DUBLER:  I think that we need to collect data and we need to add very clear endpoints of the rupture rate, especially the silent rupture rate, and to learn more about the surgical revisions since I think we know something about them but not as much as we should.

DR. WHALEN:  Dr. Newberger?

DR. NEWBERGER:  I agree with my colleagues on the left side of the table here and I would like to ask for the possibility of another endpoint along with the retrieval of explanted breast prostheses which is is it possible to get samples of the tissue from the capsule, or even a small amount in a punch biopsy extracapsular, this might be helpful in analyzing the outcomes of some of these patients down the line.  It shouldn't be technically very difficult at all.

DR. WHALEN:  Dr. Boulware.

DR. BOULWARE:  I am worried that if we do mail-in self-reported data, we will be facing this same problem five years from now with confounding data which we cannot interpret.  So I don't think that is adequate.  I would agree with Dr. Li and Dr. Newberger about the explanted retrieval data as well as, perhaps, the capsule biopsy at the time of explant.

DR. WHALEN:  Dr. McGrath?

DR. McGRATH:  I agree also about long-term follow up and, again, I am not quite clear as we are talking this, you are saying what other specific endpoints.  Well, I am not quite clear what endpoints would be the endpoints that would be the ones that would be in the process to start with because there is bunch of others that you could think about adding beside the ones that Dr. Li and my colleagues on the left have suggested.  Certainly health status, quality of life, needs to be looked at, again, since the only data we have is at two years and it is confusing, to say the least.

A lot of other things could come out.  For instance, before we had a discussion about fibromyalgia and how it was self-reported.  I mean, there are certainly a lot of opportunities here if you are going to do this to get some of that into a different setting so that you could really look at some of these systemic symptoms more effectively.

DR. WHALEN:  Dr. Leitch?

DR. LEITCH:  I agree with all the panelists who suggested that the follow up should continue to the ten-year point with physician examination, which I would recommend on an annual basis.  I think that six- to ten-year period actually offers the point in time of the greatest opportunity to identify problems which, to the patient, may be asymptomatic but, on questioning the patient or by examining the patient, it may be evident that she has a complication of the implant with respect to rupture.

So, again, that is a high-yield time and I think that that does demand the physician participating and the follow up.  And then the things that Dr. McGrath has mentioned for the follow up of the patients I think is important.  I would continue the MRI follow up out to the ten years, at least in a cohort of the patients that are followed and, absolutely, those patients are identified to be ruptured should be evaluated for their follow up specifically with detailed follow up to find out what happens to them in additional surgical procedures that are done.

I also, which I forgot to say earlier, don't think it would be crazy to do a randomized study of what to do with the asymptomatic rupture identified on MRI because, like I said, it is our gut feeling to get it out because you don't think it ought to be there and the potential local complications, which is more my concern.

But, again, with the benefit of having surgery versus not, that could certainly be looked at in a randomized trial.

DR. WHALEN:  Dr. Chang?

DR. CHANG:  I have already expressed my concern based on data presented three years ago, that mail-in surveys are not very successful in getting a high compliance rate.  I believe ten years is reasonable follow up especially to get the completion for the MRI data.

DR. WHALEN:  Dr. Witten, despite the lateness of the hour, there was a wonderfully harmonious thumbs-down to the mail survey.  Certainly, the preponderance of opinion expressed by the panel members was that hard data should be collected for at least ten years and perhaps beyond that and the same would apply to the MRI study.

The type of data should be collected face-to-face and endpoints that should be considered to be included within this include, certainly, the rupture rate, the explantation rate, what may happen to the capsular tissue itself and the extracapsular tissue, whether the ruptures are intra- or extracapsular, a defined set of systemic problems that should be rigorously surveyed for, whatever surgeries were employed and the effects upon both the pregnancy of any woman who has an implant and their children and whatever health of developmental problems the children might experience if they are born to a mother who had an implant place.

Does that answer your questions, Dr. Witten?

DR. WITTEN:  Yes.  Thank you.

DR. WHALEN:  If it does, I would like to point out that there is no resident member of the panel so we don't have to send anybody home at noontime tomorrow, fortunately.  That being said, we can adjourn for a brief nap and be back here at 7:30.

[Whereupon, the proceedings were recessed, to be resumed at 7:30 a.m., Wednesday, October 15, 2003.]

- - -