Chapter	9

DATA ANALYSIS AND REPORTING

John M. Keller

Oak Ridge National Laboratory

INTRODUCTION

Many of the areas discussed in this chapter are the procedures, guidelines, expectations and other processes that should be well defined in a laboratory(s Quality Assurance Plan (QAP).  A general laboratory QAP should have the flexibility to both define a minimum level of data quality as well as provide guidance for higher levels of QA to meet regulatory and compliance requirements.  There are some who think all laboratory work should require high levels of QA but the problem with this position is the high cost.  Another fundamental problem that can be routinely observed in almost all environmental laboratories involves the application of high level QA requirements for processing the laboratory samples, but little or no requirements for the field sampling process.  It does not matter how good the laboratory analysis is performed if there is no effort to collect representative samples. There is a process called Data Quality Objectives (DQO) that was developed to define the levels of QA needed by the customer and this process is designed to apply a graded approach.  If the DQO process is used and properly applied it can have a significant impact on overall cost for laboratory analyses.  Since the level of QA can have a significant impact on cost, many laboratories will have multiple project specific QA plans to satisfy different customer needs.

The raw counting data is rarely useful without some type of additional data manipulation to convert the counting results to an activity or concentration.  The data manipulations can be as simple as dividing the counts per second by the detector efficiency to more complex least-squares fitting techniques to define peak shapes which are then used to calculate peak areas and/or heights.  Typically the counting data must be corrected for branching decay modes, gamma energy yields, efficiency, and sample dilutions.  Another type of correction applied to counting data includes a mass absorption correction which can be especially important for alpha counting and to a lesser degree beta counting.  Mass absorption corrections are usually less important for higher energy gamma counting but may be required for specialized applications.  Overall, there are many data manipulations performed to reach a final result which also means there are many opportunities to make mistakes.  Any complex process, such as what is involved with reporting radiological data, requires a good technical review to ensure data quality.

CHAIN OF CUSTODY & LABORATORY SAMPLE TRACKING

Any good laboratory should have control over the access and security of analytical samples received from customers.  At a minimum, this sample control should include documentation to verify when the sample was received, who delivered the sample, the customer’s name, sample identification, and the requested analyses.  A typical form to document this transaction is referred to as a Chain-of-Custody (COC) form.  A COC form is usually signed and dated by both the person delivering the sample and the individual receiving the sample at the time the sample is delivered to the laboratory.  The person receiving the sample should inspect and document the sample condition as received before accepting custody of the sample.  In this day-and-time any analytical sample could become the subject of litigation, especially if the sample is related to environmental compliance, environmental monitoring, or waste disposal.  As part of the sample receipt the inspection should also include a radiological screening to ensure level of activity does not exceed any operational limits for the laboratory.  This is also an opportunity to segregate samples with higher levels of activity from lower levels sample to help minimize cross contamination problems.  Any other information available about the sample history should be documented because it may be important for the interpretation of the results.

QA/QC RECORDS

There are various types of documents and records generated during the operation of a radiochemistry laboratory and many of these records may be needed for various audits, quality assurance requirements, and other documentation needs.  Therefore, there needs to be a document control program which is usually defined by the laboratory quality assurance program (QAP).  The types of records that need to be controlled include operating procedures, analytical methods, calibration records, sample data files, control charts, and many other types of technical documents created during the operation of a laboratory.  The records to be controlled need to be identified and listed in the laboratory QA plan.

There are many Laboratory Information Management Systems (LIMS) available that provide customized database management tools and interfaces to help manage the large amounts of information generated by a laboratory operation.  A typical LIMS system includes tools to track samples, collect data, and generate reports.  In addition to managing sample data, most LIMS systems are designed to handle quality control (QC) data such as spike recoveries, blanks, calibration checks, and instrument performance data.  Once the QC data is collected and stored in the data base it can be used to generate control charts to monitor the performance of both methods and instruments.  Quality control charts for calibration check standards and QC samples can be grouped by analyst and used to document continuing qualification of the laboratory staff for performing various methods.   Since any well managed laboratory performs routine monitoring of instrument performance, backgrounds, and method performance, this information needs to be stored in a format that is easy to retrieve, recall, and monitor for long term trends that could indicate problems.  Most short term trends are detected with weekly and daily checks performed by laboratory staff.

Although it is very useful to store information electronically, the computer database systems do not eliminate the need to retain hardcopies of information related to the sample analysis.  The laboratory needs to maintain a well organized filing system that has controlled access and is protected from fire and water damage.  The types of hardcopy information typically stored in this filing system would include chain-of-custody documents, sample preparation records, computer printouts from instruments and other forms of raw data, instrument calibration and performance data, and any other type of sample information not suitable for electronic storage.  For many laboratories it is more convenient to collect certain types of information in laboratory notebooks or logbooks.  If logbooks are used there needs to be a system in place to link or associate the information with both the hardcopy files and the sample database stored electronically; usually these associations can be done with a sample identification code and a date.  For example, if a laboratory keeps a calibration logbook for pipets and a pipet is identified on a sample preparation worksheet, an auditor should be able to go back with a date and a pipet identification code to verify that the pipet was properly calibrated.

In a regulatory or compliance environment the laboratory must document and be able to track all aspects of the laboratory operations including sample management, sample preparation, instrument calibration, standard preparation, and analyst qualifications.  Also, it is a good management practice (required for some forms of regulatory work) to participate national performance demonstration programs where a nationally recognized group prepares samples with a known composition for distribution to laboratories participating in the program.  The participating laboratories analyze the samples as unknowns and report the data.  The data for all the participating laboratories is then summarized in a report that each laboratory can compare their performance with the group average and the known values.

DATA PRESENTATION

There is no standard or historically based method on how to present radiochemical data although there have been attempts over the last 10-15 years to encourage most laboratories to use the International System of Units (SI) units. The SI units for reporting radioactivity are the internationally accepted units and will eventually be the standard in the United States; however, this is not the current practice for reporting most environmental radioactivity data.  The Becquerel (Bq) is the SI unit for activity and is defined as one disintegration per second (dps); the historical unit for activity is a Curie (Ci) which corresponds to the activity for one gram of 226Ra.  The relationship between a Bq and Ci is defined as follows:

The level of radioactivity in environmental samples is usually discussed in units of curies (Ci) and concentrations are generally expressed in terms of fractions of curies per unit mass or volume.  Since a curie represents a comparatively high activity, it is more common to see concentrations for environmental samples expressed as pCi/L or pCi/Kg (pCi = 10-12 Ci).  A list of some other common units used to express activity is given the table below.

| Symbol | Name |
| --- | --- |
| Ci | curie |
| dps or Bq | disintegration per second |
| dpm | disintegration per minute |
| mCi | millicurie (10-3 Ci) |
| Ci | microcurie (10-6 Ci) |
| nCi | nanocurie (10-9 Ci) |

The ultimate goal when reporting data is to provide the customer or data user with the information in the most useful format possible to meet the customer’s needs.  There can be a wide variety of people, with a broad range of technical background, that use the radiochemical data reported. The various types of data users typically include radiation protection technicians, project/program managers, scientists, environmental monitoring groups, regulatory/compliance organizations, and even lawyers.  Therefore, it is important to know the audience that will use the information to ensure the correct interpretation of the data and to avoid misconceptions.  Even though it is well known that there are significant sources of natural radioactivity (cosmic radiation, 40K, 232Th, 235U, 238U, etc.) in the environment, the presence of any radioactivity in the ground water, surface water, and soil is a very sensitive area when dealing with the public.  It is not uncommon to see various media outlets, in an attempt to sensationalize news, to report radiological data with units that result is a large numbers.  For example, consider a commercial reactor has a release of 1 mCi of 60Co into a stream that results in a concentration downstream of 0.0001 mCi/mL; this would likely be reported to the public by the media as 100,000,000 pCi/L.

Another important aspect that needs to be considered when reporting radiological data (or any analytical result) is the relative error associated with the result.  No measurement performed in a laboratory provides an absolute value for the result and no result should be reported without an estimate of the overall error.  The overall error includes the standard deviation of the measurement, which is a measure of the precision, along with all other sources of error.  The combination of all these sources of error is referred to as propagation of error and this subject is discussed in more detail in Chapter XXX.  The overall measurement error is especially important for environmental or low level samples because as the measured activity approaches method detection limits the analytical error can become very large.  For low level samples the counting statistics will frequently dominate the overall error.  The standard error (G) for a gross number of counts (N) is,

In a similar fashion, the standard error (g) for a count rate (N/t) would be,

If a radionuclide of interest is not detected there needs to be a means to report the minimum activity level the method could have measured.  Different conventions and terminology have been used to estimate the lower limit of detection (LLD) and the technique suggested by the Department of Energy Environmental Measurement Laboratory (DOE-EML) will be discussed here.  The LLD is defined as the smallest quantity of sample radioactivity that will yield a net count for which there is a predetermined level of confidence that radioactivity is present.  The LLD may be approximated as follows:

Where,	kα	= risk for concluding falsely that activity is present (α),

kβ	= degree of confidence for detecting the activity (1-β),

σr	= standard error for count rate (r),

G	= gross counts collected in time tG,

B	= background counts collected in time tB.

If α and β are equal then (kα + kβ) can be represented as 2kα/β. If the sample and background counting rates and count times (tG=tB) are similar, as would be expected near the LLD, then the above equation can be simplified to,

Some common values of k are given below

| α/β | (1-β) | kα/β | 2(2 k |
| --- | --- | --- | --- |
| 0.01 0.02 0.05 0.10 0.20 0.50 | 0.99 0.98 0.95 0.90 0.80 0.50 | 2.327 2.054 1.645 1.282 0.842 0.0 | 6.58 5.81 4.65 3.63 2.38 0.0 |

For α and β equal to 0.05 and when the gross count time is equal to the background count time, then the smallest amount of radioactivity that has a 5% (α) probability of a false positive and a 95% (1-β) probability of being detected is,

This equation for LLD95 is used by many laboratories; however, one must use care that the given conditions apply.  If the gross and background count times are different and/or different probabilities are used the first equation must be used to calculate LLD.  To obtain the LLD95 in units of activity one must correct for detection efficiency (E), decay branching ratio (β), and sample quantity (Q) which would include any dilutions.

Continue this discussion to include reporting negative results vs. detection limits.

DATA REVIEW

There are different types of reviews that are routinely performed on analytical data, including the radiochemical results.  One type of review that is commonly used on regulatory data packages is referred to a data validation.  Data validation is more of a clerical review of the data to ensure that a data package includes all the information and raw data needed to reconstruct and defend the final results.  These extensive data packages are required by the customer to meet regulatory or compliance requirements and typically include documents such as the raw data, instrument calibration information, certificates for standards, sample preparation information, chain-of-custody, along with the final results.  Whereas a data validation may include some verification of calculations, this type of validation typically does not include an extensive scientific review of the final results. An acceptable technical review requires much more experience and technical background than normally available for data validation.  Some suggested areas to look at for a technical review will be the topic for the rest of this section.

A good technical review of radiochemical data usually requires more than just a laboratory background and should draw from a broad range of experiences including nuclear chemistry, nuclear engineering, reactor operations, waste operations, and even some basic background in nuclear weapon testing is useful when reviewing environmental data because of the global fallout from above ground testing.  A technical reviewer should have a good understanding of the nuclear processes that produce the various radionuclides. For most laboratories the radionuclides of interest would involve natural products (i.e., 40K, 232Th, 235U, 238U, etc.), fission products (e.g., 90Sr, 137Cs 154Eu, etc.), activation products (e.g., 60Co, 63Ni, 65Zn, etc.), and decay daughters (90Y, 208Tl, 210Po, 228Th, etc.).  However, do not forget the sample history because some unusual radionuclides could be found around accelerator sites, medical isotope production facilities, university and government nuclear operations, and even hospitals with nuclear medicine treatment facilities.

To begin the review process it is always good to do a reality check to make sure the data makes sense.  For example, one would not expect to see short lived fission products (i.e., 103Ru, 131I, 144Ce, etc.) in an old waste tank that historically collected liquid waste from a spent fuel processing facility but had not been used for ten years.  Due to the relative fission yields and the half-lives for fission products, the activity in aged (5-10 years decay) fission product waste would be dominated by 90Sr (t1/2 = 29.1 y) and 137Cs (t1/2 = 30.2 y) followed by 154Eu (t1/2 = 8.59 y) and 155Eu (t1/2 = 4.71 y).  If short lived fission products were reported on this ten year old waste the data would be questionable.  Either there was an error in the identification of the radionuclides or there was much more recent use of the waste tank than originally indicated.  These types of reality checks on the data are not always so obvious and it may take a experienced reviewer to recognize basic problems with the data that do not make sense physically or chemically.  If there is some sample history available, each radioactive species identified in the sample should have reasonably defined pathway to explain why each radionuclide would be present in the sample matrix.  Any historical information about the site where samples are collected is useful for the technical review and validation of radiochemical data. The types of site information should include information on the various production operations such as waste processing, topsoil characteristics in the area sampled, surface water locations, annual rain fall, and sometimes even the average daytime temperatures may be useful in explaining sample results.  Historical information might also include radiological data where past samples were collected at or near the same site.  If historical data is available and there is some knowledge of activities performed at the site during the period between the sample collections times, there may be enough information to estimate some expected outcomes such as types of radionuclides present and even activity levels after decay corrections.

Another aspect of date review is to make sure that the chemistry makes sense, for example, the presence of thorium in a caustic solution would be suspect because of the low solubility of thorium at high pH.  Sometimes the chemistry can become very complex and hard to predict but for most cases the chemistry of the radionuclides will follow expected behavior. For example, most anionic species tend to be water soluble and are mobile in the environment. Two good examples are salts of pertechnetate (99TcO42-) and radioactive iodide (129I-); both of these examples are long-lived fission products that are very mobile in the environment. For the most part the radionuclides within the group of transition metals, the lanthanide and actinide series, and even some of the Group I and II elements that are present as cations in the environment are absorbed on soils or precipitated at high pH values.  Some metals readily form complexes with common organic acids found in the environment to form water soluble species.  However, it important to realize that soil in various geological locations can vary in composition and chemistry, which can have a significant impact on the environmental mobility of radionuclides.

The soil in east Tennessee has a strong affinity for actinides and cesium, but there is almost no absorption of strontium.  Even in areas around nuclear facilities that experienced releases of fission product waste from underground storage tanks and associated piping many years ago, there has been minimal movement of the 137Cs and actinides from the immediate area.   However, the many years of exposure to ground water and rain moving through the soil has removed most of the 90Sr from these areas.  The 90Sr moves so freely through the soil in east Tennessee that if outflows around these areas are sampled soon after a rainfall and immediately analyzed almost none of the 90Y normally found in equilibrium with the 90Sr is present.  As water moves through the soil the 90Y is continuously removed by the soil and only the 90Sr is carried by the ground water.  The identification of beta activity from 90Sr that was not in secular equilibrium with 90Y has led to the misidentification of environmental beta activity on more than one occasion.

Sometimes the sample history can help identify some expected trends in the radiological data; however, there are additional expected outcomes based upon the decay of radionuclides that must be true if the nuclide identification is correct.  Many radionuclides, both natural and manmade, have radioactive daughters.  This decay process can continue past the first daughter and result in decay chains where the activity of the progeny can be estimated from the original parent.  Some of the more complex decay chains are found in nature.  The 238U decay chain, often referred to as the uranium series is illustrated on the next page.  The other two decay chains found in nature are the actinium series which starts with 235U and the thorium series which begins with 232Th.  The neptunium series which starts with 237Np used to be found in nature millions of years ago, but all currently available 237Np is manmade.  In general, most of the manmade radionuclides do not have a decay chain as long or complex as the decay series found in nature, but many of the isotopes routinely identified in the laboratory have one or two daughters that also will be present in the sample. These daughters can be useful for confirming nuclide identification and may even provide a second confirmation of the parent activity.  To fully appreciate the nuclear decay process the reviewer needs to understand the ingrowth of radioactive decay products and the three different types of parent-to-daughter equilibriums that are possible.

A more detailed discussion of the ingrowth of decay products can be found in several of the references at the end of this chapter; Nuclear and Radiochemistry by Friedlander, et. al., provides an excellent discussion of this subject.  A set of equations can be derived, which are referred to as the Bateman equations, to calculate the number of atoms or the activity of each daughter in a decay series.  The equation for the first daughter, N2(t), as a function of the initial amount of the parent, N1(0), and time (t) is given below in terms of the number of daughter atoms,

Since A=N, this equation can be restated in terms of activity,

These equations can be expanded to include additional progeny in a decay chain, but only the first daughter in a series will be discussed here.  There are three special cases that depend upon the relative half-lives for the parent and daughter in a decay chain summarized in the table below.

| Decay Constant or Half-life | Type of Equilibrium |
| --- | --- |
| 1 < 2 T2 < T1 | Transient equilibrium, where the ratio of the parent and daughter decay rates become constant. |
| 1 << 2 T2 << T1 | Secular equilibrium, where the activity for the daughter becomes approximately equal to the activity of the parent. |
| 1 > 2 T2 > T1 | No equilibrium. |

The condition of secular equilibrium, where the half-life of the daughter is much less than the parent, is the most useful case to the radiochemist for the measurement of activity.  Both 90Sr and 137Cs are long-lived fission products routinely observed in environmental samples that demonstrate secular equilibrium.

The time required for ingrowth of the daughter to approximately match the parent activity is in the order of 10X the half-life of the daughter radionuclide. This ingrowth time is illustrated for the 90Sr/90Y secular equilibrium in the following curve.

Although an understanding of daughter ingrowth and secular equilibrium is very useful to the radiochemist doing measurements it also helps the data reviewer validate and confirm results.  There are a number of cases where the identification of a radionuclide is based upon measurement of a daughter in secular equilibrium; several examples are listed below,

For each of these examples listed, the daughter in secular equilibrium is easier to measure by gamma spectrometry than the parent.  This is especially true for the 106Ru since this radionuclide represents a “pure” beta emitter.  Frequently, the 1001 KeV gamma from 234mPa can be used to quantify uranium in soil samples.

REPORTING ACTIVITY FROM MASS SPECTRAL DATA

Because of the improvements in sensitivity, reliability, and lower cost mass spectrometers are becoming available for more radiochemistry laboratories.  Although Thermal Ionization Mass Spectrometers (TIMS) have been around for many years for the measurement of isotopic ratios for uranium and plutonium the Inductively Coupled Plasma Mass-Spectrometer (ICP-MS) has recently become popular for a wider range of radionuclides, including many of the long-lived fission products such as 99Tc, 129I, as 135Cs.  Mass spectrometers provide a measurement of the number of atoms for each isotope present which is directly proportion to mass.  Typically, results from a mass spectrometer are reported in units of concentration in ng/g or ng/mL; most users of radiological data are interested in units of activity such as Bq/mL or pCi/g.  The conversion factor between mass and activity is defined as the specific activity (SA) which has units of activity per unit mass.  The specific activity for each radionuclide can be determined from the following:

Since,

Then,

Where T1/2 is the half-life in units of seconds, NA is Avogadro’s number (6.022 x 1023), and MA is the atomic mass of the radionuclide of interest.

Some radionuclides can decay by more than one path such as 227Ac (T1/2 = 21.77 y), which is a natural decay daughter of 235U.

The total decay constant T for 227Ac can be represented as the sum of the partial decay constants  + .  Since decay constants are a function of the half-life the concept of a partial half-life can also be introduced.  Since the decay constant is defined as,  = ln 2 / T1/2, the half-life for 227Ac can be represented with partial half-lives as follows:

REFERENCES

Evans, R. D., The Atomic Nucleus, New York: McGraw-Hill, 1958.

Horrocks, D. L., Applications of Liquid Scintillation Counting, New York: Academic, 1974.

Friedlander, G., Kennedy, J. W., Macias, E. S., Miller, J. M., Nuclear and Radiochemistry, 3rd ed. New York: Wiley-Interscience, 1981.

Tsoulfanidis, N., Measurement and Detection of Radiation, New York: McGraw-Hill, 1983.

Knoll, G. F. Radiation Detection and Measurement, 2nd ed. New York: Wiley, 1989.

Ehmann, W. D. and Vance, D. E., Radiochemistry and Nuclear Methods of Analysis, New York: Wiley-Interscience, 1991.

# Text boxes

90Sr+90Y

90Y

90Sr

Example

What is the specific activity for 137Cs (T1/2= 30.17 y)?





234Pa

6.70 h













































5.30MeV

7.69MeV

6.00MeV

5.49MeV

4.78MeV

4.69MeV

4.78MeV

4.20MeV

238U Decay Chain

206Pb

stable

210Po

138.4 d

206Tl

4.20 m

210Bi

5.01 d

206Hg

22.3 y

210Pb

22.3 y

21oTl

1.3 m

214Po

0.164 ms

214Bi

19.9 m

218At

1.6 s

214Pb

27 m

218Po

3.10 m

222Rn

3.82 d

226Ra

1600 y

230Th

7.54x104 y

234U

2.46x105 y

234mPa

1.17 m

234Th

24.1 d

238U

4.47x109 y

99.8%

99.98%

0.02%

99.979%

0.021%

vw

vw

Example

If the partial half-life for the  decay of 227Ac is 22.075 y, what is the  activity per gram for 227Ac (T1/2 =21.77 y)?