Fermilab/AD/TEV

Beams-doc-1499-V0

December 10, 2004

Tevatron Beam Position Monitor (BPM) Upgrade

Timing Generator Fanout Module, Filter Card and Integrated Crate Test Plan

Dehong Zhang, Luciano Piccoli, Bill Haynes, Ken Treptow, Jim Steimel, Vince Pavlicek, Tim Kasza

# Introduction

This document specifies the test plan for the standalone quality control (QC) tests performed by the CD prep group for the Timing Generator Fanout (TFG) modules and Filter Cards (FC) as well as the integration tests performed on the fully assembled and cabled crates before they are taken to the Tevatron service buildings. It is often logical to modify the original qualification criteria. However, justifications of such changes and modified test results should be included in this document.

# Timing Fanout Generator Stand-alone QC Test plan

Introduction

The Timing Generator Fanout is a Double Width VME card’s main function is to provide a precision phase locked clock and Sync pulse (or Gate) to each of the EchoTek digitizers of the Beam Position Monitor (BPM) system. Also, the TGF provides a high resolution time stamp and can generate VME interrupts based on 16 different TCLK events. The Clock fanout section of the TGF requires only power from VME to operate and can by tested independent of the operation of the FPGA controlled timing section. After the usual verification that power/ground plane shorts do not exists, proceed as follows:

Equipment and setup needed for the stand alone quality control test:

- VME64X crate.

- Motorola MVME2400 processor board.

- VME Console Terminal.

- Frequency Counter, Spectrum Analyzer or Oscilloscope.

- Various cables.

Test procedure

Step 1:  Clock fanout section

- Insert the module into an un-powered VME Crate. Turn on the crate power and the 5.0v & 3.3v LEDs should turn on.

- Connect an RFCLK (53Mhz clock) source to the RFCLK input and note the LED labeled RFCLK should turn on. The LED turns on when the internal PLL locks to the RFCLK and produces an output clock multiplied by a 7/5 ratio.Since the RFCLK is generally 53.10468MHz then the output clock is 74.34655MHz.

- Connect a frequency counter, spectrum analyzer or oscilloscope to the

CLK0 output connector. In general an oscilloscope is adequate to test the clock fanout circuit for functionality but a frequency counter or spectrum analyzer should used occasionally to measure that the locked output frequency is precisely 74.34655MHz (unless there is beam being accelerated in the main ring). If a spectrum analyzer is being used, any side band noise should be attenuated by least 50dbm.

- Compare the remaining CLK1 - CLK7 to CLK0.

Step 2:  FPGA Controlled Section

The FPGA section consist of the VME interface, the BSync decoder and TClk Decoder to control Sync pulses the the EchoTek digitizers and interrupts to the VME CPU.

- First the FPGA firmware must be downloaded to the TGF via the Front JTAG port using the Quartus programmer software. The firmware file name is “bpmfanout.pof” and takes about 5 minutes to download.

- Set the pc board  eight position dip switch to all “on” except for switch position 2 being set to  “off”. This gives a VME base address of 0x2000.

- Check to see if the VME select LED flashes during a read to VME address 0x2000 using the TGF diagnostic software.

- Run the TGF diagnostic software “runTimingTest” for the remaining tests.

- Select menu item # 0 to verify all the TGF registers are readable and writeable from VME.

- Connect BSync to its front panel input and note that the BSync LED turns on.

- Connect TClk to its front panel input and note that the TClk LED turns on.

- Select the BSync Test menu item # 1 of TGF diagnostic software. This checks the BSync event interrupts to the VME CPU with all the events required in the full BPM System.

- Select the TClk Test menu item # 2 of the “TGF diagnostic software”. This checks TClk event interrupts to the VME CPU based on the decoding of specific TClk events and requires TClk to be connected to the TGF card. This test will try several combinations of TClk events interrupts and report success or failure.

- Select the BSync Test menu item # 3 of TGF diagnostic software. This test uses the BSync revolution marker (0xAA) for generating the trigger and Sync output pulses. Connect the front panel TRIG1 (TEST on prototype card) to an oscilloscope trigger input and verify that each of the Sync(0-7) outputs generate one pulse per trigger. The output pulse of  Sync0 will be generated first, followed by output Sync1, then Sync2, thru to Sync7. Each of the pulses is evenly spaced through one revolution of the beam (~21us).

- Select the Wait TClk menu item # 4. User inputs a particular TClk event (decimal) to be decoded. If the TClk event is not decoded in 30 seconds an error is reported.

- Run the “TGF diagnostic software” TimeStamp test. This test will measure the time interval from a BSync Clear event to the first TurnMarker after a Start event. This time interval is based on a 106Mhz counter and should be the same for each TGF card.

- Run the “TGF diagnostic software” FilterCard test. This test will permit individual relays within the Filter card to be set or cleared via the VME backplane. Writing a “01” to channel0 should be an adequate enough test to show that the TGF can communicate with the Filter Card. This command would close the contacts of the relay connecting the diagnostic signal (53Mhz) to the BPM Channel0 output of the Filter Card.  If Channel0 is connected to an oscilloscope, then the 53Mhz signal would appear after the value is written.

- To write to a particular TGF register outside the diagnostic software:  setTSGRegister  offset, value
Note: offset and value are in standard hex format, for example 0x102, 0x0203.

- To read to a particular TGF register outside the diagnostic software:  getTSGRegister  offset
Note: offset is in standard hex format, for example 0x102.

- This concludes the incoming production tests and qualifies the card for system level tests.

# Filter Card Stand-alone QC Test plan

Introduction

The Filter Card is a single width VME64X card.  The Filter Card obtains 5V, 12V, -12V power, SYSCLK, and interfaces to the Timing Generator Fanout (TGF) via reserved bussed pins from the VME crate.  The Filter Cards main function is to filter and attenuate the signal from the TeV BPM sensors going to the Echotek digitizers.  It also allows via relays for a diagnostic signal to be sent to the Echotek, the BPM sensor, or both.  The TGF card controls this through its VME interface.

Equipment and setup needed for the stand alone quality control test:

- VME64X crate.

- Motorola MVME2400 processor board.

- Two Echotek ECDR-GC814-FV-A modules.

- Timing Generator Fanout card.

- VME Console Terminal.

- Agilent 4395A Network Analyzer and 87511A S-Parameter Test Set.

- Various cables.

It requires two external inputs: Ethernet access to the 253 subnet of FCC for the PPC processor and the 53.1 MHz Tevatron RF clock for the TGF.

Here the TGF serves 3 purposes:

- Interface to the filter board to control the relays;

- Generate the 53.1 MHz diagnostic signal on the VME back plane;

- Generate the 74.3466 MHz clock signal for the Echotek boards.

One of the Echotek boards serves as the down-stream digitizer, while the other as fake pick-ups up-stream to monitor the signals sent back to the BPM pick-ups.

The PPC processor boots from “fndaut”, and can save test results to “/usr/devel/area2/vxworks/filter_board_test” on “fndaut”.  The loadable modules, Echotek setup files and the startup script can be found in: “/usr/products/devel/vxworks/tbpm_filter_board_test” on “fndaut”.

Here is a picture of the setup, except for the VME console terminal.

Test procedure:

Step 1: Visual Inspection and Initialization

- Visually inspect the card for damage, solder shorts, or missing components.

- Verify with an ohmmeter that there are no power to ground shorts.

- Insert the card into an un-powered VME64X crate.  Turn on the crate power and the 5V, 12V, and –12V LEDs should turn on.

- Download the FPGA firmware (bpmfc_control.pof) via the JTAG port using the Altera programmer software and an Altera ByteBlaster cable or equivalent.

Step 2:  Diagnostic Signal Test

Description of the tests

Currently, we foresee to perform 3 tests on each production filter board.  First we set the TGF to generate the 53.1 MHz diagnostic signal, then:

- Set the TGF to switch all the 8 relays on the filter board to “ADC”, or

- Set the TGF to switch all the 8 relays on the filter board to “BPM”, or

- Set the TGF to switch all the 8 relays on the filter board to “ADC and BPM”

Next we software trigger the Echotek boards to digitize and analyze the signals they receive, then read out the Echotek boards to check the 53.1 MHz components against some predetermined range (minimum – maximum) criteria.

The software for these tests is actually a merge and adoption of the multi-board test for the Echotek driver and the PREP filter test for quality control of the production Echotek boards.  The core part is the setup file for the Echotek boards.

We took “DDC_board_test.ch” from the PREP filter test, modified the following parameters:

- Clock rate from “80 MHz” to “74.3466 MHz”;

- NCO frequency from “3 MHz” for an over-sampled 3 MHz input to “21.2418857 MHz” for an under-sampled 53.1047 MHz input;

- RCVR_DATA_SKIP from “0” to “18” to skip the transient effect;

- OUTPUT_WORD_SIZE from “6B” to “2B” to force one RCVR_DATA_SKIP to correspond to one BURST_COUNT.

With this setup, the Echotek boards would digitize the 53.1 MHz diagnostic signals at 74.3466 MHz, down-convert their images at 21.2418857 MHz to DC, decimate and filter out frequency components beyond +/- 0.26 MHz.  The total decimation is 16.  The overall frequency response is shown in the following plot:

For each test, we take 64 samples from each channel and find the minimum and maximum.  The current cut-off range criteria are as follows:

| ECDR Channel | ADC | BPM | ADC & BPM |
| --- | --- | --- | --- |
|  | min | max | min | max | min | max |
| Proton ADC | 1200 | 1600 | 0 | 50 | 700 | 1000 |
| Pbar ADC | 4000 | 5000 | 0 | 50 | 2300 | 3000 |
| BPM Pick-up | 0 | 50 | 14000 | 18000 | 9500 | 13000 |

The 2 Echotek boards are set to A16 addresses of “0xc000” -- the one on the left, the “ADC”, and “0xd000” – the one on the right, the fake BPMs, while the filter board is set to index “0”.

Testing procedure

- Set the filter board to index “0”, insert into the crate, make sure it has seated properly;

- Connect the 8 output cables straight across horizontally to the corresponding input connectors on the Echotek board on the left – the ADC board;

- Connect the 8 input cables slightly slanted to the corresponding input connectors on the Echotek board on the right – the fake BPM pick-ups;

- Power on the crate, watch for the boot process, in the end there will be 3 times rather loud “pop” sound, while the PPC prints to the console screen like the following:

Filter Board Test ->RunTests

0x1da441c0 (tShell):

0x1da441c0 (tShell): Diagnostic 53.1 MHz  ==>  ADC

0x1da441c0 (tShell): Channel   1:      35      34      35 -- BAD !!!

0x1da441c0 (tShell): Channel   2:      35      34      35 -- BAD !!!

0x1da441c0 (tShell): Channel   3:      34      33      34 -- BAD !!!

0x1da441c0 (tShell): Channel   4:      32      32      33 -- BAD !!!

0x1da441c0 (tShell): Channel   5:      31      30      31 -- BAD !!!

0x1da441c0 (tShell): Channel   6:      34      34      35 -- BAD !!!

0x1da441c0 (tShell): Channel   7:      33      32      34 -- BAD !!!

0x1da441c0 (tShell): Channel   8:      36      35      36 -- BAD !!!

0x1da441c0 (tShell): Channel   9:      20      19      22

0x1da441c0 (tShell): Channel  10:      22      21      24

0x1da441c0 (tShell): Channel  11:      28      27      29

0x1da441c0 (tShell): Channel  12:      26      25      27

0x1da441c0 (tShell): Channel  13:      14      13      15

0x1da441c0 (tShell): Channel  14:      14      13      15

0x1da441c0 (tShell): Channel  15:      19      18      20

0x1da441c0 (tShell): Channel  16:      19      18      21

0x1da441c0 (tShell):

0x1da441c0 (tShell): Diagnostic 53.1 MHz  ==>  BPM

0x1da441c0 (tShell): Channel   1:      34      33      35

0x1da441c0 (tShell): Channel   2:      35      34      35

0x1da441c0 (tShell): Channel   3:      34      32      34

0x1da441c0 (tShell): Channel   4:      32      30      33

0x1da441c0 (tShell): Channel   5:      30      29      32

0x1da441c0 (tShell): Channel   6:      34      34      35

0x1da441c0 (tShell): Channel   7:      33      32      34

0x1da441c0 (tShell): Channel   8:      36      36      37

0x1da441c0 (tShell): Channel   9:      20      18      23 -- BAD !!!

0x1da441c0 (tShell): Channel  10:      21      20      23 -- BAD !!!

0x1da441c0 (tShell): Channel  11:      28      27      29 -- BAD !!!

0x1da441c0 (tShell): Channel  12:      26      25      27 -- BAD !!!

0x1da441c0 (tShell): Channel  13:      14      13      15 -- BAD !!!

0x1da441c0 (tShell): Channel  14:      14      12      15 -- BAD !!!

0x1da441c0 (tShell): Channel  15:      19      18      21 -- BAD !!!

0x1da441c0 (tShell): Channel  16:      19      17      22 -- BAD !!!

0x1da441c0 (tShell):

0x1da441c0 (tShell): Diagnostic 53.1 MHz  ==>  ADC and BPM

0x1da441c0 (tShell): Channel   1:      34      34      35 -- BAD !!!

0x1da441c0 (tShell): Channel   2:      35      34      35 -- BAD !!!

0x1da441c0 (tShell): Channel   3:      34      33      35 -- BAD !!!

0x1da441c0 (tShell): Channel   4:      32      32      34 -- BAD !!!

0x1da441c0 (tShell): Channel   5:      31      30      32 -- BAD !!!

0x1da441c0 (tShell): Channel   6:      34      34      35 -- BAD !!!

0x1da441c0 (tShell): Channel   7:      32      32      34 -- BAD !!!

0x1da441c0 (tShell): Channel   8:      36      35      37 -- BAD !!!

0x1da441c0 (tShell): Channel   9:      20      18      22 -- BAD !!!

0x1da441c0 (tShell): Channel  10:      22      20      23 -- BAD !!!

0x1da441c0 (tShell): Channel  11:      28      27      30 -- BAD !!!

0x1da441c0 (tShell): Channel  12:      26      25      27 -- BAD !!!

0x1da441c0 (tShell): Channel  13:      15      14      15 -- BAD !!!

0x1da441c0 (tShell): Channel  14:      14      13      16 -- BAD !!!

0x1da441c0 (tShell): Channel  15:      20      19      21 -- BAD !!!

0x1da441c0 (tShell): Channel  16:      19      17      20 -- BAD !!!

0x1da441c0 (tShell):

0x1da441c0 (tShell): This board FAILed the tests !!!

Where the last printout should declare the board to have passed or failed the tests.  If the board fails, please check the cable connections, and re-run the tests by typing “RunTests” on the console keyboard;

- Type ‘save “Serial No”’ on the console keyboard to save the test result in  “/usr/products/devel/vxworks/tbpm_filter_board_test” on “fndaut”.  The file will have a name like ‘Board_“Serial No”.txt’, and read like the following (from a different test than the above):

FRI DEC 03 14:44:55 2004

--------- FAIL ---------

======================================================

Diagnostic 53.1 MHz  ==>  ADC

Check criteria:

Proton ADC     1200 --   1600

Pbar   ADC     4000 --   5000

Pick-up           0 --     50

Proton & Pbar               Pick-up

-------------------   -------------------

mean    min    max    mean    min    max

Channel 1:    1424   1423   1426      27     25     29

Channel 2:    1393   1392   1394      26     25     28

Channel 3:    4287   4286   4288      30     29     31

Channel 4:    4119   4117   4120      29     27     29

Channel 5:    1330   1329   1332      16     14     18

Channel 6:    1363   1362   1364      19     18     21

Channel 7:    4503   4502   4505      24     23     26

Channel 8:    4535   4533   4536      22     20     24

======================================================

Diagnostic 53.1 MHz  ==>  BPM

Check criteria:

Proton ADC        0 --     50

Pbar   ADC        0 --     50

Pick-up       14000 --  18000

Proton & Pbar               Pick-up

-------------------   -------------------

mean    min    max    mean    min    max

Channel 1:      17     16     18   16239  16218  16257

Channel 2:      17     16     18   15972  15952  15991

Channel 3:      16     15     17   15650  15633  15666

Channel 4:      19     18     19   15021  15006  15036

Channel 5:      15     14     16   15057  15045  15073

Channel 6:      14     12     15   15652  15640  15672

Channel 7:      10      8     11   16091  16077  16112

Channel 8:       9      8     10   16308  16292  16331

======================================================

Diagnostic 53.1 MHz  ==>  ADC and BPM

Check criteria:

Proton ADC      700 --    1000

Pbar   ADC     2300 --   3000

Pick-up       9500 --  13000

Proton & Pbar               Pick-up

-------------------   -------------------

mean    min    max    mean    min    max

Channel 1:     859    858    861   11174  11159  11188

Channel 2:     830    829    831   11048  11033  11062

Channel 3:    2562   2560   2563   10859  10842  10872

Channel 4:    2457   2456   2459   10397  10383  10407

Channel 5:     794    793    795   10403  10392  10415

Channel 6:     813    811    814   10824  10814  10836

Channel 7:    2680   2679   2682   11150  11140  11164

Channel 8:    2707   2705   2708   11309  11299  11323

Step 3:  Through Response (S21) Test

- Using the Agilent 4395A with the S-parameter test set, measure the magnitude and phase difference between channel pairs.  Calibrate to one channel and measure and record the difference of the second.

- Repair out of spec channels.

# VME Crate Integration Test

Once the VME crate is stuffed with individual boards, loaded with appropriate firmware, following test procedures are performed. The objective of this test procedure is to fully qualify the crate before taking it to the assigned service buildings.

Equipment and setup needed for the stand alone quality control test:

Test procedure:

Step 1: Check …

Goal:

Goal of the testing …

Equipment and HW/SW setup needed for the stand alone quality control test

Pass/fail Criteria:

Performer:

Step 2: Check …

(Repeat steps as necessary)

Test result documentation (spreadsheets, screen captures etc. and their on-line locations)

# Change Log

| Version | Issue Date | Description of Change |
| --- | --- | --- |
| 1.0 |  | Original |

# Concurrence

Following persons reviewed and concurred with the content of this document.

Steve Wolbers, Project Manager	(date)

Bob Webber, Deputy Project Manager  (date)

Vince Pavlicek, Subsystem manager	(date)

# Text boxes

53.1 MHz In

74.3 MHz Out