INTERMET, Inc.

7732 Castleisland Drive

Sarasota, FL 34240

USA

Phone + 941 378 1313

Phone/ Fax + 941 379 3536

Infotiempo@comcast.net

To:	Mr. Oscar Arango B

WMO Sub Regional Office for the Americas

San Jose, Costa Rica

Subject:   The following report describes the activities, recommendations and accomplishments of the Contractor for the period June 4 thru September 28, 2004.

In General:

The Contractor has accomplished all the activities proposed for the first period of the Proposal and Terms of Reference.

Missions have been carried out to GSN and GUAN sites.  Missions to the 9 Mexican Sites and San Andres Island will be performed when funds are made available to the Contractor

The Contractor has completed technical files on all the sites visited and will mail them to the Sub Regional Office in San Jose

The files are constructed such that any future maintenance contractor or technical expert will be able to use them for their purposes

The Contractor has provided instruments to sites.

The Contractor has purchased instrument standards for comparisons at the sites.  The details of the instruments will be provided to the Sub Regional Office.  The instruments will be retained by Intermet, Inc until the contract terminates.  The Instruments will then be turned over to the Sub Regional Office.

The contractor has used Spanish speaking experts from countries in the region to assist other countries.

The contractor will review the needs of the sites with the Sub Regional Office to determine priorities for procurements and expenditure of funds

Training plans for the sites will be reviewed with the Sub Regional Office and GCOS authorities to identify priorities and timing.

Danny L. Foster____________________         September 28, 2004

Intermet, Inc.

Cc: Richard Thigpen, GCOS Officer

GCOS TECHNICAL SUPPORT PROJECT

FOR THE

CARIBBEAN, CENTRAL AMERICA AND MEXICO

1ST Quarterly Report

September 28, 2004

INTERMET, Inc

7732 Castleisland Drive

Sarasota, FL 34240

USA

CLIMAT and CLIMAT TEMP Reports

Situation:

None of 6 GUAN sites do CLIMAT TEMP reports

Only about 60% of expected CLIMAT reports were received at NCDC during Jan-July2004

4 of 17 GSN sites reported 100% of the time

9 of 17 GSN sites reported roughly 70% of the time

4 of 17 GSN sites were silent

Only 4 sites report all 4 CLIMAT groups, 111,222,333, and 444

See CLIMAT TEMP and CLIMAT Summaries below.

There are several years of upper air data and surface data available at each site that could be put into CLIMAT TEMP or CLIMAT format.

Why ?:

GCOS is not a high priority at the sites

Sites are not familiar with GCOS program

No exchange of information or rapport between the Sites and the monitoring 	centres

Staffs are not trained in data assimilation and coding the report

The existing training tool, CLIREP CD is cumbersome and not user friendlyl

Telecommunications paths and responsibilities are not fully understood by all 	participants including the RTHs and GCOS Centers.

CLIMAT AND CLIMATE TEMP REPORT SUMMARIES

Discussions with the staffs at 5 of the 6 sites, namely Barbados, Belize, Bogota, Kingston and San Jose indicate they do not prepare or submit CLIMAT TEMP reports.

Barbados 78954

Kingston 78397

The Directors are willing to begin the CLIMAT TEMP program but request some training for their staffs.  The CLIREP program was given to Kingston and to Barbados.

Belize 78583

Belize discontinued the CLIMAT TEMP program in the fall of 2003.  They are preparing to resume the program and would like some training.  The Contractor supplied them with the CLIREP program

San Jose 78762

San Jose discontinued the CLIMAT TEMP program in July 2001 because the communications link with Hadley Centre failed.  The Contractor worked with both parties and resolved the issue.  Hadley Centre should begin receiving the San Jose CLIMAT TEMP reports on a regular basis in September 2004.

Bogota 80222

Bogota has not reported Temps since March 2004 because of a communications problem and the winds component of their Vaisala GPS system failed because of local interference.  The Contractor is working with them to resume their observations without the winds and to begin using the NOAA EDIS system for transmission of their Temp messages.  The solution to the missing winds can be solved with the provision of a different type of Vaisala radiosonde.   Training in the assembly of CLIMAT TEMP data and the preparation of the CLIMAT TEMP should be considered for this staff.

Manzanillo 76654

Hadley Centre reports that Manzanillo submitted some CLIMAT TEMP reports in late 1999 and early 2001 and none since 2001.   The observatory is transmitting 27 to 30 Temp reports per month so they should be doing the CLIMAT TEMP reports.  The Contractor will investigate this and report accordingly

CONTRADICTION

CLIMAT TEMP records from Hadley Centre indicates the sites have in fact been submitting CLIMAT TEMP reports.  Nothing on a regular basis but still their records show the  receipt of some CLIMAT TEMP reports.  The Contractor advised each of the 5 GUAN sites of Hadley Centre’s records.  Excepting San Jose, the sites were puzzled and told the Contractor they did not do those reports and did not know who did those reports.  Hadley Centre indicated they were reported in the same CLIMAT TEMP bulletin as other Caribbean stations.  That bulletin is CUCA01 KWBC dtg, CLIMAT TEMP dtg.  The Contractor will continue to investigate this mystery.

The following table shows the archived CLIMAT TEMP messages at Hadley Centre.

|  | Barbados Belize Kingston Bogota San Jose Manzanillo |
| --- | --- |
| 1999 | 1 12 7 2 |
| 2000 | 7 11 8 |
| 2001 | 11 6 11 9 7 6 |
| 2002 | 11 7 10 9 |
| 2003 | 3 8 5 2 |
| 2004 | 0 6 June 4 May |

SUGGESTIONS

The Contractor firmly believes that training in the compilation of data for the CLIMAT TEMP report and the construction of the actual report should be given to each of the 6 sites.  If funding does not permit individual on-site training then perhaps English and Spanish speaking sessions should be considered in central locations.

There is a substantial amount of raw data at each of the GUAN sites that has never been submitted to GCOS.  The data is there and should be researched and put into the CLIMAT TEMP format and submitted to Hadley Centre.  The Contractor is willing to lead this effort if it is considered desirable by GCOS authorities.

The Contractor’s several years of experience in overseeing the management of the upper air sites in these locations believes the integrity of this upper air network including the submission of the CLIMAT TEMP and CLIMAT reports will benefit from close monitoring and interaction with the Directors and their staff.  Please bear in mind that CLIMAT TEMP and CLIMAT reports are not at the top of the sites priority list.  Another aspect that is just as important is establishing and maintaining interaction between the GCOS monitoring centers and the sites themselves.  I know there was some attempt made at this by identifying a GCOS focal point at each site.  That did not prove to be valuable.  To do this job properly someone has to maintain a close vigilance on these activities and be proactive in identifying and correcting problems.

CLIMAT REPORTS

The Contractor performed missions to all the assigned GSN sites excepting those in Mexico and San Andres Island 80001.   With the exception of Acajutla, El Salvador 78650 all of the sites visited by the Contractor are doing a very good job in preparing and submitting the CIMAT reports.  Some sites are only doing the minimum message i.e. parts 111 or maybe 222 even though they have the background data to do parts 333 and 444.

Parts Submitted

Bahamas Nassau     78073    111,222

Barbados Grantley  78954    111

Cayman                   78384    111,222, 333, 444

Pto Limon               78767    111,222, 333, 444

Montego Bay          78388    111,222, 333, 444

Acajutla                   78650    none

San Andres              80001    unknown

9 Mexico                               unknown

The Contractor does not see any need for training in the preparation of the CLIMAT messages.  An exception may be Acajutla, El Salvador.  They haven’t been doing the CLIMAT report and will need some training.  The training should include organizing the data in preparation for the message, message preparation, and ensuring the message gets to the proper GCOS center on time.

There has been some confusion in the exchange of CLIMAT reports between Grand Cayman 78384 and the centers.  Grand Cayman has been using AFTN as their primary communications link and apparently this method does not get the CLIMAT message properly distributed to the GCOS centers.  This problem is currently being investigated by the Contractor and others.

Nearly all the instrument fields have a variety of instruments to measure the same parameters.  In addition to the standard instruments they often have an automatic weather station with more state-of-the-art instruments.  In some cases, the observatory uses the AWS for the hourly and synoptic temperature and at the same time they use the thermometers in the instrument shelter for their max/min reports.  The station should specifically identify which instruments are used for the measurements recorded in the CLIMAT message.  I believe this can be resolved with current Metadata information.

Recommendations

GCOS Program Establish a Better Presence at the Sites

GCOS Management needs more rapport with the NMHS Directors and Managers 	through personal visits

Monitoring Centers or GCOS Management should maintain open 	communications with 	sites. Advise the sites 	when they do not receive the 	monthly reports by the 10th of each month.  Provide recognition to sites doing 	good jobs.

GCOS Centers or Management should provide monthly monitoring reports to the 	sites.  The reports should be stylized for the sites and shows how the sites in the region compare with each other.  This is a good management tool for the Directors and Supervisors and provides encouragement to the staffs.

All sites are willing to begin CLIMAT TEMP program if they receive training in data assimilation and in coding the report.

On site personal training is preferred over centralized training

Centralized training in two locations.  One Spanish and one English

Standards, and procedures for quality control of the surface data used in the CLIMAT reports should be reviewed by GCOS experts on mission to the sites.  If funding is available, centralized training courses in quality control and CLIMAT message coding should be considered.

Accomplishments

Identified and helped resolve communications problems which led to the resumption of CLIMAT and CLIMAT TEMP reports

Identified major communications problems which are being resolved to allow GSN site to enter their Synops and Hourlies onto the GTS

Identified and helped resolve responsibilities and communications problems between RTHs and GCOS centers that had caused loss of CLIMAT reports.

Reviewed the CLIMAT data quality control procedures at the GSN sites and made recommendations

GSN Sites

Instrument Fields

12 of 16 GSN sites are located on airport facilities.

Of the 6 GSN sites visited, the exposure of 4 instrument fields have been degraded because of airport construction and traffic encroachments.

All 4 instrument fields will likely be moved to more remote locations.  None of the sites have plans to do comparison observations

Instruments

In some cases, the current temperature and humidity measurements were taken from the instruments mounted on their automatic weather stations while the max and min temperatures were taken from liquid-in-glass instruments in the instrument shelter.

Calibration or comparison standards for surface observing equipment exists at only one GSN site.  Comparisons and calibrations have not been performed at 5 GSN sites in recent history.  Some barometers are more than 30 years old and have never been compared to a standard or been calibrated.

Some sites are periodically replacing sensors instead having a maintenance and calibration program.  They claim the replacement option is less costly in terms of funds and personnel.   This practice is contrary to the “GCOS Best Practices for GSN Stations”

The Contractor is working with NCDC to get more specific requirements for metadata NCDC does not have metadata for any of the GCOS sites.

Recommendations

GCOS initiate near real-time monitoring of the flow of surface observations and keep sites aware of missing reports.  This will promote enthusiasm and cooperation with the sites.

A high precision portable barometer should be purchased and taken to each site for comparison readings with the station official barometer

One set of RM Young wind system calibration tools and equipment should be purchased and kept at a central location for use at different sites.

Contractor should be proactive in working with sites that have locations for their new instrument field.  This may require purchasing some instruments for comparison observations.

Contractor should establish a project to assist all the sites in producing accurate metadata and providing that data to the Centers

Accomplishments

Contractor identified all the above problems and is in the process of assisting the sites to overcome them or reduce their effect

The Contractor identified communications problems at GSN sites and established replacement communications using internet

The Contractor raised the level of interest and enthusiasm in the GCOS program at the sites.

GUAN SITES

The Contractor has performed missions to 5 of the six GUAN sites.  Manzanillo has not been visited.

The attached individual site reports describes the operation and problems the staffs are encountering along with site specific recommendations by the Contractor

In General

The sites have too much dependence on the manufacturer for trouble-shooting, maintenance, replacement parts and repair of the system.  This is true for both the IMS 1500C and Vaisala systems.

The observers and electonics technicians at the IMS 1500C sites have never been properly trained.  The observers do not know how to make adjustments, corrections, or to manage out-of-the-ordinary events during the soundings.  This inability leads to early terminations of soundings because of the loss of the signal.

In earlier years when the sounding was reduced and coded manually the observations were carefully inspected by Asheville and error reports were sent to the sites for their information and to correct any faulty procedures.  In today’s world the automatic systems are not supposed to make mistakes.  However they do make mistakes and upper air observations are frequently rejected by the models.   The observatories are not informed of these errors or rejections as they occur and are unaware that they or their equipment are not functioning properly.  An excellent example is the wind bias problem at Barbados.  The Contractor discovered the problem by mistake in conversation with NCEP authorities.  Barbados did not know they had this problem.

NOAA/NWS is installing the IMS 2000 system in their upper air network.  InterMet Systems focus and support will be towards the IMS2000 and not the IMS 1500C.  The NWS parts and supply depot will be tailored for the IMS2000.  The IMS 1500C will rapidly become the stepchild for support from InterMet Systems and NOAA depot.  The Contractor thinks the conversion from the 1500C to the 2000 series is not that difficult although it has to be done at the factory.  The Contractor does not have solid information on this conversion.

Excepting Mexico, the sites using the expensive Vaisala GPS sondes will never be independent of external funding.

The Vaisala sites use 300 gram balloons.  The 300 gram balloon bursts much lower than the 600 gram balloon.  None of the GUAN sites exceed the 5 millibar level and only those using the  600 gram balloon ever exceed 10 millibars.

The automatic and direct transfer of the Temp message from the iMetOS Compaq computer to the Corobor or the MetLab system does not work.

This past summer proved the weak link in the GUAN network remains the production of hydrogen gas.  The 3 GUAN sites that use the Proton Generator were inoperative at times because of internal power supply failures.  Two of the three sites had 2 or more failures within a 3 months period.

Recommendations

IMS 1500C Sites

Centralized or on site training by the manufacturers for both the electronics technicians and the observers.  On site training is preferred because more staff can be trained at each location and they will use their own equipment and the software that is resident in their system.  Centralized training is less costly but accommodates fewer students.  However Centralized training allows for the exchange of information and experiences among the students.

At this time, the Contractor cannot recommend the provision of spares, specialized tools or test equipment for this system.

The practice of using Target Antennas for direction checks before each flight should be reinstated and emphasized at the site.  The Contractor will help the sites build and mount the target antennas.

Vaisala Sites

GCOS should create  2 or 3 day workshop in San Jose, Costa for observers and electronics technicians from Bogota and from Mexico.  A Vaisala representative should provide some training and be an active participant.  The Contractor can develop the agenda.

PROTON HYDROGEN GENERATOR

Training on this generator is needed at the sites.  This may be done on site or centralized with the IMS 1500C training.

Spare master and slave power supply units should be provided to each site.

Extra water filters for the Hydrogen Generator and the Aqua solutions system should be provided to each site.

UPS systems or some type of power surge suppressors must be used to protect the generators.

Proton Calibration kits and instructions should be provided to each site as recommended by the manufacturer

BALLOONS

The Contractor recommends the sites using 300 gram balloons convert to 600 gram balloons to attain better height averages.  Also the Contractor believes the use of 800 gram balloons will provide better chance for getting above the 5 Millibar level.  This was a successful strategy used during the IGY years.

INDIVIDUAL GSN SITE REPORTS FOLLOW

NASSAU BAHAMAS 78073

Number and types of Surface Observations and CLIMAT reports

Nassau Airport is providing approximately 650 hourlies and usually over 200 synops per month.  NOAA/NCDC records indicate Nassau routinely submits the monthly CLIMAT reports.  Nassau uses the NOAA Email Input Data System for its communication with Washington DC.

The May 2004 CLIMAT message follows:

CLIMAT 05004 78073

111 10185 20193 30253010 402930218 5244 60030104 7819114 8000000 9000000

Note that groups 222,333,and 444 are not reported.

Instrument Field

During the past several years the airport has grown significantly and sadly enough it has grown around the instrument field.  The screened area of the grass covered field is within a plot of ground about 100 feet square.  However black paved roads surround the field on all four sides with black paved parking lots on 3 sides.  See the diagram in the Bahamas Technical File. The influence of the pavement and auto traffic likely introduces additional heat the instrument field but this has never been thoroughly examined.  The staff feels the temperatures measured in this field is usually 4 or 5 degrees higher than the temperature taken at the nearby radiosonde station.

Specifications for all the instruments used for the GCOS CLIMAT message and program are included in the Bahamas Technical File.

Apparently the Airport Authorities want to move the instrument field to make room for a more profitable parking lot.

Instruments and Methods of Observations:

The accuracy of the surface instruments that are used for the CLIMAT report can only be judged by comparison with a certified standard.  The thermometers were compared with the handheld Lufft E200 Stand V1 high precision thermometer.  The differences were usually less than two tenths of a degree and were dictated more by the error of parallax.

A standard 8 inch rain gage is the official instrument for the measure of rainfall.  There is very little chance for error in using this rain gage.  The placement of the gage is in a good location and away from high buildings or structures that may influence the catch.  The station also uses a tipping bucket to record the rate of rainfall.

The station uses a Bendix-Friez wind velocity and recorder system.  The Transmitter is a model 120, the Indicator is a model 135, and the Recorder is a model 141-1.  The vane is located in Center Field on a 30 foot tower approximately 2700 feet from the office.  The last calibration was performed in November, 2003.  The station calibration equipment for this system is inoperative.

The sunshine recorder is the typical Glass Ball, Casella Model 7630 mounted 4 feet above the surface in the instrument field.

A Negretti Aviation aneroid  barometer, Model MK-2 Type 2236 is used as the official station pressure.  The height of the barometer is reported to be 23 feet above mean sea level.  There is no calibration sheet for this barometer.  A comparison with a barometer standard should be conducted.

Metadata

Metadata for Nassau is not on file at NOAA/NCDC

Historical Record:

The contractor was advised that they have historical records dating back to the 1930s.  However there are gaps and missing data until about 1963 when the site was established at its present location.  The contractor is trying to get more specific information on the historical records.   The site meets the Minimum Requirements related to temperature and precipitation which calls for monthly for the past 20 years.  The recent records are homogeneous is terms of continuity without missing months.   See GCOS-73 paragraph 3.6

Suggestions:

If the instrument field is going to be relocated then a program of comparative readings should begin as soon as possible or as soon as the new location is identified.  See Manual on the Global Observing System, WMO-No.544, 2.10.3.17.

The Bendix wind system needs to be routinely calibrated.  Provide needed equipment to site.  Another GSN site in the region has the required calibration equipment.  Perhaps a loan of this equipment can be arranged.

The station barometer should be compared with a standard for accuracy.

The station should update their Metadata information and provide a copy to the NOAA/NCDC

Temperature in Instrument Field needs further investigation.

BARBADOS  GSN 78954  (does not include the GUAN report)

Number and types of Surface Observations and CLIMAT reports

NOAA/NCDC statistics for the past 4 years indicates Barbados  is providing a large number of surface observations to the GTS.  They are averaging approximately 675 hourlies and 200 synoptics monthly and routinely submits the CLIMAT message to NOAA/NCDC.

The following is a sample CLIMAT for the month of July 2004.

CSBR01 TBPB

CLIMAT 07004

CLIMAT 111 10086 20151 30278013 403020250 5332 61270417 7233059

Please note that sections 222, 333, and 444 are not reported.

INSTRUMENT FIELD

The current instrument field is large in comparison to other observatories.  The field is grass covered, with ample spacing for all the instruments  and is enclosed by a high fence.  The instrument field was originally located near the old airport terminal which is about a half mile from its present location at the new terminal.  The instrument field was moved to its present location in 1979.  During January 2004, the airport authorities at Grantley Adams constructed a huge aircraft hangar within a few yards of the existing instrument field.  The Director of the Barbados Meteorlogical Service was not consulted about the location or design of this hangar and his concerns were ignored.   This hangar obstructs the entire northeast quadrant of the observing field.  All observed winds from 360 to 090 degrees are considered to be erroneous.  Also there are several instances when the accuracy of the collection of rainfall is questionable because the hangar stops the flow of the wind causing “deadspots” over the gages.

Currently (2004), the airport is undergoing major renovation and expansion which is causing the authorities to once again move the instrument field.  The new instrument field will be 1 kilometer due east of the existing field and at nearly the same height above sea level.  The new site is 100 square feet in area, grass covered and is near the east end of the runway and roughly one half mile from the ocean.  The authorities have indicated that this area will not be developed and the surrounding area will be kept free of buildings, traffic, and pavement.  While electricity will be available at the instrument field it will be vulnerable to power outages.  Some type of backup power such as a UPS should be installed at the field.

INSTRUMENTS AND OBSERVING PRACTICES

A Casella Automatic Weather Station (s/n M103159) system is located near the center of the instrument field.  The AWS is composed of a Casella wind system mounted atop the ten meter tower with a combined temperature/humidity sensor located further down the structure.  Also a precision barometer is located inside the AWS electronics box.  The AWS also includes a rain gage tipping bucket which is not used for the official CLIMAT or Synoptic reports.

A standard 8” rain gage is the official instrument for the site.  The gage is read during the standard synoptic observations and the data is recorded accordingly.

A Campbell-Stokes sunshine recorder is also located in the field.  Standard procedures are used for recording and reporting the duration of sunshine.

A standard Evaporation Pan with a wire mesh protective covering and wind counter is located in the field.  These observations are performed in accordance with standard procedures.

An instrument shelter or Simpson Screen is properly mounted and located in the field.  The shelter houses ventilated Casella psychrometeres and Casella Max/Min thermometers.  The maximum and minimum temperatures for the CLIMAT report and Synoptic reports are taken from these thermometers and not from the Casella AWS.

The observer uses the temperature, humidity and wind velocity measurements from the Casella AWS for the hourly reports.  The official barometer is a Negretta Precision Barometer Mark II ref. Met20005, s/n 8483/83 located in the observers office.  The barometer is 166 feet above sea level while The airport runway ranges between 152 and 159 feet above sea level.  The reports are inserted onto the GTS via the Corabor VSAT system.  Comparisons with a Standard Barometer should be made to ensure the accuracy of this Negretta Barometer.

All the observations are routinely stored and sorted in a local file called MONFSUM.EX.  This file performs some very rudimentary quality control checks and puts the data into a format that is designed for composing the CLIMAT message.   This is a very efficient system which promotes accuracy in a variety of ways.  The station only reports the 111 groups.

A FEW COMMENTS ON THE SURFACE OBSERVING PROGRAM

The accuracy of the surface instruments that are used for the CLIMAT report can only be judged by comparison with a certified standard.  The thermometers were compared with the handheld Lufft E200 Stand V1 high precision thermometer.  The differences were usually less than two tenths of a degree and were dictated more by the error of parallax.

The contractor did not have a portable air pressure standard to make comparisons with the official station barometer.  However, it appears the station barometer is within acceptable limits when it is compared with the barometer on the AWS and the aneroid barometer at the radiosonde office.

As indicated above, the collection of rainfall may be compromised by the nearby location of the airplane hangar.

The official wind system is a Casella 3 cup anemometer with a pointer/tail direction configuration and is approximately 2 years old.  The cup model is 032340 and the direction is model 032336.  Casella claims the system is accurate, doesn’t need maintenance and never needs calibration.  The station instrument expert agrees with this claim and indicated that Casella does not provide test or calibration equipment for this system.  The wind velocity at the GSN compares favorably with the winds observed by an RM Young system at the radiosonde office a few miles away.  However, northeasterly wind measurements at the GSN must be viewed with skepticism as indicted above.

Historical Record

The official records for the old instrument field date back to 1942.  Some of the records were lost during a hurricane in 1955.  Users of these records should note that during 1971-1972 the instruments were temporarily placed on the roof of a building which resulted in spurious temperature readings.  The instruments were removed from the roof and put back in their original location near the old terminal.

Metadata

Metadata for this site is not on file at NOAA/NCDC.  The observatory is currently taking steps to update the Metadata and send the information to NCDC, et.al..

Suggestions

A program should be initiated to begin comparing the surface observations in the new instrument field with the existing instrument field.  See Manual on the Global Observing System, WMO-No.544, 2.10.3.17

Comparisons of winds and air pressure measurements should be made with official standards.

Metadata needs to be updated and the information sent to NOAA/NCDC.

GCOS authorities should advise Barbados if they want them to begin including 222, 333, and 444 groups in their CLIMAT message

GCOS centers should note the suspect temperatures measurements during 1971 -1972 and the suspect wind velocity and rainfall data beginning January 2004 to present times.

GCOS authorities should note that current temperature and humidity reported in hourlies and synops use the AWS instruments while the synoptic reported max/min temperatures are taken from the liquid-in-glass instruments in the instrument shelter.

Grand Cayman 78384, Owen Roberts Airport

Number and types of Surface Observations and CLIMAT reports

Grand Cayman is open from 6:00AM till 10:00PM and is listed in Volume A as taking all synops except  06Z, and 09Z .  They should be averaging 6 synops per day or 180 per month.  They averaged between 85 and 90 synops and 500 hourlies per month. The NOAA/NCEP report http://www1.ncdc.noaa.gov/pub/data/gcos/GSN_sum_200101_200406.txt   indicates they haven’t submitted a CLIMAT report for the entire 4 years covered in their report.    Director F. Sambula assured me they were reliably doing the CLIMAT report and sending them via AFTN to KWBCYMYX with the header CSGC01 MWCR.  The observational data is entered into the station CLICOM system which generates the CLIMAT message.  Basic quality control is exercised within the CLICOM software.

Please note that Grand Cayman is one of the few GSN sites that include 333 and 444 groups in their reports.  Below is a copy of the CLIMAT message for July.

CLIMAT 07  4
78384
111  1//// 20150 30295007 403260252 5315  60224/13    7//////
8310000     9000231

333  03131 31307 40701 9010131

444  0030806 1028315 2033811 3022814 4083114 5//////
608//   711806

Instrument Field

The instrument field has been in approximately the same location for 40 years. The instrument field is roughly a 900 square foot grassy area with a four foot  fence.  There are no noticeable obstructions to the field that may influence the instrument readings.  Once again, as the airport gets busier the competition for land space gets more intense.  For several years the airport authorities have envisioned moving the weather office to a different location.  The move may not include the instrument field and may only apply to the upper air facility.

Instruments and Methods of Observation

The instrument shelter is a standard Stevenson Screen mounted over a grassy area in the instruments field.

The station uses standard liquid-in-glass thermometers for their wet and dry bulb as well as standard liquid-in-glass maximum/minimum thermometers.   The thermometers are not aspirated.  Comparisons were made with the Luft E-200, 5220.00  Serial 029-0404-0302.3.  The comparison indicated the liquid-in-glass thermometers were accurate within plus or minus .3 degrees F.  The error of parallax is more of a determining factor.

The official rain gage is the standard US Weather Service Model: D-100 8 inch.  The gage is favorably located within the instrument field and is not influenced by an surrounding structures.

The barometer is a Vaisala DD-50 located in the air control tower.

The aerovane is a Bendix Corporation Model 120 mounted on a 20 ft tower and not a ten meter tower.  The aerovane is located at the Northwest end of the runway.  The wind recorder is a Belfort Instrument Co. model 141-1 and the Wind speed and direction indicator is a Bendix Friez Model 135.  The station does have the required calibration equipment for this system.  ON July 13, 2004 , the contractor checked the accuracy and calibrated the system.  The system was found to be within the limits of the specifications.  See the technical report for details on all the instruments and calibration tables.

Historical Record

Apparently there were weather records dating back to 1936 as well as weather observations performed by the US Navy during 1942-45.  These records may have been lost during moves and construction at the airport.  A more organized weather observation program began during the IGY year in 1956-57 at a location about 5 miles from the current location at the airport.  In the early 1960s the USA began full support to the surface and upper air observatory in Cayman which more or less continues today.   The site meets the Minimum Requirements related to temperature and precipitation which calls for monthly for the past 20 years.  The recent records are homogeneous is terms of continuity without missing months.   See GCOS-73 paragraph 3.2.

Metadata

There is no Metadata on record at NCDC

Comments on the distribution of the Grand Cayman 78384 CLIMAT message

Problem:

Apparently the use of AFTN to send their CLIMAT reports to a center to be inserted onto the GTS does not work.  NCDC, DWD, JMA, and U of East Anglia indicate they do not receive the Grand Cayman 78384 CLIMAT reports.  However, NOAA Data Management Division indicates they receive the CLIMAT reports and routes them to NCDC.  NOAA needs to resolve this internal problem.

The AFTN is successfully used by several countries to transmit their observations and products to the GTS.  The replacement systems for the STAR4 do not satisfy the needs of the Latin America sites for the transmission of data and several use the AFTN in which they have more confidence.  The use of the NOAA EDIS system is gaining popularity because of its simplicity and everyone has internet.  Grand Cayman successfully uses the AFTN for the following messages successfully:

All messages are sent to KWBCYMYX and SBBRYZYX

Message Type			Header

TAF’s				FTCA1 MWCR

METARs			SAGC31 MWCR

SYNOPS			SMGC01 MWCR

UPPER AIR			UMGC01 MWCG

The CLIMAT message is sent to KWBCYMYX with the header CSGC01 MWCR

Grand Cayman does have full access to internet at their weather service and we are trying to arrange for them to send their CIMAT message to NCDC using the (NOAA) Email Data Input System (EDIS).   It appears that using the EDIS system is efficient in inserting the data they receive onto the GTS.  Other GCOS GUAN and GSN sites use the EDIS system and their observations routinely appear on the GTS.  The Contractor has been working with NOAA to get the Grand Cayman CLIMAT message onto the EDIS system.  On September 3rd, 2004 the Contractor was advised that NOAA Data Management will work with Cayman to get the EDIS implemented in Cayman.  After this is implemented the Contractor will try to monitor the flow of the data from Cayman to the other Centers.  This will not be easy unless the Contractor gets some support and timely response from the NOAA Data Management Offices.

At this particular time, the Contractor cannot locate anyone that knows of the observatory in Grand Cayman listed in WMO Volume A as 78383 at lat. 19 15 N and lomg. 81 25 W.  The contractor suspects this was the “Cuban Observatory” that used to operate 5 or 6 miles from the airport during the 1950s that was subsequently abandoned.  This cannot be verified until the Contractor can contact authorities in Grand Cayman.

PLEASE NOTE:  HURRICANE IVAN DEVASTED THE ISLAND OF GRAND CAYMAN.  THEY HAVE NOT TAKEN A WEATHER OBSERVATION SINCE SEPTEMBER 11, 2004.  THE ISLAND IS WITHOUT ELECTRICITY, FRESH WATER, AND TELEPHONES.  THE AIRPORT IS CLOSED EXCEPT FOR EMERGENCY AID FLIGHTS.   IT WILL BE QUITE A LONG TIME BEFORE THE PROBLEMS ADDRESSED IN THIS DOCUMENT CAN BE ADDRESSED BY THE NMHS OF GRAND CAYMAN.

SAN ANDRES ISLAND, COLOMBIA 80001

Number and Types of Observations

This airport observatory is not open for the 03Z and 06Z observations and averages roughly 160 synoptic and over 500 hourlies per month.  They are very reliable in sending in their CLIMAT messages and have not missed one message in the past 4 years.

The contractor has not visited this GSN site.

Metadata

There is not Metadata on file at NCDC

Suggestions

The station should update or create a Metadata file and send the information to NCDC

PUERTO LIMON, COSTA RICA 78767

Number and Types of Observations

There are two certified observers at this small airport on the eastern coast of Costa Rica.  The office opens at 6:00AM and closes at 7:00PM daily.  They perform the 12Z, 15Z,18Z,21Z and 00Z synoptics.  They provide between 250 and 300 hourlies per month and 60-100 synops.  They have only missed 4 or 5 CLIMAT reports during the past 4 years.

The CLIMAT report is created at the headquarters climate section in San Jose.

Typical CLIMAT report from Pto Limon follows:

111 78767 10132 20136 30250007 402850214

5266 601562214 7158097

8000000 9000000

222 07100  10128 20134 30241005

402910206 5264 6027617

7163 8000000 9000000

333 03002

31409 40600 9000000

444 0026255 1023123 2030506

3019375 4047403 60000

Instrument Field

The instrument field is roughly 50 yards from the airport terminal and has excellent exposure.  The field is only about 15 yards from the actual Caribbean Sea and provides an excellent sampling point for a maritime tropical air masses (mT).  There are no obstructions or artificial influences on the field.  The field about 20 feet on a side and is grass covered with a fence.  All of the instruments used for the observations are located in this field.  The instrument field has been in this location for at least 30 years.  Instruments have been changed and replaced over the years but there is not definitive record of these changes.

Instruments and Methods of Observation

A wooden white Simpson Screen instrument shelter is located in the instrument field.  They use Lambrecht wet and dry bulb thermometers with a wind-up fan aspirator.  The Maximum and Minimum thermometers are Lambrecht 2742/92 and 5579/02 respectively.  These were compared with the contractors temperature standard handheld Lufft E200 Stand V1 high precision thermometer.  All thermometers compared favorably.

They use a standard mercurial fixed-cistern barometer for their official station pressure.  There is some confusion over the height of the barometer.  The height was determined to be 3 meters above sea level in 1991.   An earthquake occurred and it is believed the barometer is now about 5 meters above sea level.  This height estimate is only a guess.  The barometer has been in the same location for 30 years.  The staff cannot recall any maintenance ever being performed on the barometer (like cleaning the mercury) nor has there ever been any official comparisons made with a standard.  However, the staff does closely monitor the pressure recorded by a precision barograph with the mercurial barometer.  If some difference occurs they will notify their authorities.

The wind system is an R.M. Young propeller-type anemometer with fuselage and tail wind vane (type 05106).  The vane is mounted on what appears to be a ten meter mast and is part of a Campbell Scientific automatic weather station.

They have 3 rain gages in the instrument field.  The official gage is a 200mm mouth catcher.  The other automatic recording gages are not used in the CLIMAT report.

They use a standard Campbell-Stokes glass ball for the sunshine observations.

NOTE

The observations are telephoned to the NMC in San Jose where they are entered onto the GTS.  This is not a good practice and does promote the chances for error.  The contractor worked with the Director to provide internet capability at the site so they could email their observations to the NMC in San Jose.  The Contractor used $500 GCOS funds for a 10 month contract. This is more efficient and less prone to errors.  They are currently evaluating this email method for 10 months and will then determine if it is worthwhile to make this a permanent procedure.

Historical Record

The historical records date back to 1941 and are continuous.  This satisfies part of the 
GCOS qualifications for a  Target Requirements site.

Metadata

There is no Metadata information on file at NOAA/NCDC

Suggestions

Other than the need to provide Metadata to NOA/NCDC there are no pressing actions needed for this site.

If funding becomes available it would wise to acquire RM Young Aerovane calibration equipment for the site.  Barometric comparisons with a very good standard should be performed..

The NMHS should plan funding for the permanent internet capability at Pto Limon.

MONTEGO BAY, JAMAICA 78388

Number and Types of  Observations

This airport observatory is open 24 hours per day and is expected to transmit 8 synoptics per day.  Hourly observations range between 600 and 750 per month.  They average over 220 synoptic observations monthly.  During the past 4 years they have been perfect in submitting their CLIMAT messages on time to NCDC.  Please note that Montego Bay submits all groups 111 thru 444.

CLIMAT 07004 78388

111 10142 20143 30287/// 403290246

5307 60077/09 7272021 8000000 9000000

222 05180 403130235 6005307 7257

8////00 9//0000

333 03131 10000 20000 30906

40400 50000 8010000 9000001

444 0029520 1027301 2033117 3024515

4017022 54///09 61600=

Their synoptic, metars and CLIMAT messages are sent via AFTN to Kingston AFTN and then onto Washington via AFTN.  Sometimes this fails and the synops have to be sent via voice telephone to the forecast office in Kingston for furtherance onto the GTS.

Comments

The observatory is fully staffed for 24 hours a day with 4 certified observers and the station Official in Charge (OIC).  The OIC has over 30 years experience at this observatory and is well qualified in all aspects of managing and operating this observatory.  The OIC pays close attention to the requirements of the CLIMAT reports and ensures her staff observes and records the meteorological parameters in a very professional manner.  The OIC routinely performs quality controls on all entries and calculations needed for climatological records.  This is an excellent GSN site.

Instrument Field

As with most airport observatories, the airport infrastructure grown around the instrument field.   The grass covered instrument field is probably 500 square feet and is enclosed by a 6 foot high fence.  A paved road butts against one side of the field and small structures surround two of the other sides.  The activity of auto and machinery traffic may have some effect on the temperature.  The rain gages are well located and are not affected by the surrounding infrastructure.  The wind vane is located in the middle of the runway on a ten meter tower.  The contractor believes the instrument field has been in roughly the same location since 1951.

Instruments and Methods of Observation

This observatory uses a Casella Automatic Weather Station (AWS) M103107A located in the grassy area between the runway and taxi-way  at Sangster Airport.  This AWS is used for the hourly observations including temperature, humidity, air pressure, and wind velocity.

For the synops, the observers go to the instrument field and records the liquid-in-glass Casella Max and Min thermometer readings and measures the  accumulated rainfall at a standard 8 inch rain gage.  The current temperature, humidity, wind velocity, and  barometric pressure are taken from the AWS.  The wind velocity is the standard 10 minute average.

NOTE:  The CLIMAT Max and Min temperatures are taken from glass thermometers in the instrument field and the hourlies & synops use the AWS in the runway area for all current temps and humidities.  The contractor was not able to go to the AWS site  to make comparative temperature readings between the AWS and the instrument field.

The sensors of interest on the Casella AWS are the following types:

The wind velocity sensor is a 3 cup rotor anemometer using the infrared light and interrupter.  The wind vane is the typical pointer and tail type.  The temperature sensor is a platinum resistance type and the humidity uses the capacitive system.  Both are mounted in a standard radiation screen.  The barometer is mounted 3 feet above sea level on the tower and is a Monolithic Silicon transducer with accuracy of .3 Mb. For details on these sensors please see the contractor’s Technical Report.

The hours of sunshine are measured with a standard Campbell-Stokes glass ball type instrument and is located in the instrument field.

Historical Record

Montego Bay satisfies the high level Target Requirements outlined in para 3.6 of GCOS-73 .  The continuous records date back to 1951 and include all the parameters identified in the 111, 222, 333, and 444 groups.  This is the only GSN station in the Central America and the Caribbean with these credentials.

Metadata

There are no Metadata records for Montego Bay at NOAA/NCDC

Suggestions

The AWS in Montego Bay is the same as the AWS used to support the GUAN site in Kingston.  The contractor is working with Casella to identify ways and means of ensuring the accuracy of these instruments.  If available, a set of instruments needed for AWS calibration should be provided to Kingston.  The technician at Kingston makes monthly maintenance trips to Montego Bay.

Metadata should be completed for Montego Bay and the information provided to NOAA/NCDC

GCOS management should consider using the OIC as an instructor to those sites that are lacking in their CLIMAT program.

The communications system(AFTN) thru Kingston AFTN to Washington DC is fragile.  An alternative system such as the NOAA Email Data Input System (EDIS) should be considered for this site.  Internet/Email is not being used at the weather office.

During the passage of Hurricane Ivan the electricity failed at the observatory which turned off the AWS.  Therefore there is no history of the wind velocity during the passage of Hurricane Ivan.  The tower and instruments were not damaged but sadly enough the maximum wind was not measured or recorded.   The Contractor does not know if this loss is important to the GCOS program, if so, then  some back-up power or other recording devices should be considered.

ACAJUTLA, EL SALVADOR 78650

Servicio Nacional De Estudios Territoriales (SNET)

Number and Types of Observations

This site is not located at an airport.  WMO Volume A indicates this observatory performs a full suite of observations including aviation, sea, and  hourlies 24 hours per day along with 8 synops.  In reality they are open from 6AM till 9PM and miss the 06Z and 09Z synops.  The reception of weather observations at Washington DC for the past 4 years varies considerably.  The data in the following table is taken directly from the NOAA/NESDIS  web page http://www1.ncdc.noaa.gov/pub/data/gcos/GSN_sum_200101_200406.txt.  The averages are approximate.  The staff advised me they have not done a CLIMAT report for at least 10 years.

Year/Month Hourlies Synops CLIMAT Monthly Monthly Reports Average Average

2001 Jan-Sept 0 0 0 2001 Oct-Dec 250 0 0 2002 Jan – 2004 February 400 90-170 0 2004 Feb- July 20 1 or 2 0

Instrument Field

The instrument field is approximately 50 foot square. Grass covered and surrounded by a high fence.  The field is located in sight of the shore, loading docks and piers.  The field is on a hill and is probably 30 to 50 feet higher that sea level.  The field is surrounded by large trees and a road/footpath that is heavily traveled.  The wind measurements may be affected by the trees.   There are several instruments in the field that have been donated to El Salvador over the years including the Sutron Automatic weather station.  It is believed this was donated and installed by the USA/USGS as a result of Hurricane Mitch.

The wind system is a RM Young 3 cup anemometer Model 06201 s/n WT 09233. The wind system has not been calibrated in recent memory.  El Salvador does not have the proper calibration equipment for this system.  It is believed that they normally use a anemocinemografo recording wind system but it is currently damaged.

The official rain gage is a standard Lambrecht 3400 Gottingen Plastic bottle type that the observers visits daily at 7AM and takes a measurement.  This measurement is used for the synoptic reports.  They also have a Sutron rain gage Model 5600-0425 – 2 serial 000596 that is used for local purposes.  They have a standard cotton-region type instrument screen housing the psychrometers and max/min thermometers.  The psychrometers are Lambrecht Wet 4575/94 and Dry 4629/94.  The max is a Fuess 7393-03 and the min is a Fuess 7295.  The synoptic reports use the current temperature and humidity from the Sutron AWS.   Sutron designates their sensors as the AT/RH High Accuracy 5600-0311.   The max and min temperatures reported in the synoptic messages come from the Fuess liquid-in-glass max and min thermometers in the instrument shelter.   This mix of instruments may create some confusion with researchers.

The official barometer used for the hourly and synoptic reports is a standard Fortin-type  mercurial.  The barometer is located in the observer’s office on the 3rd floor of a building near the instrument field.  The staff indicated the height of the barometer is 25 meters above msl.  WMO Volume A indicates the height is 15 meters above msl.  The Contractor recommends the height be reevaluated by a professional surveyor.   Also the Contractor suggests comparisons with a precision standard barometer be made to see if there are major differences.  The mercury in the barometer has never been cleaned and that may or may not affect the accuracy.

They use a standard Stokes-Campbell glass ball type sunshine recorder.

Communications

The observer transmits the observation to San Salvador via radio (telephone) to the Central Office S.S. (Snet/Cpm) where it is sent to Washington DC via Metlab (satellite).

PLEASE NOTE: During 2004 the receipt of observations of any type at Washington DC has dropped to nearly zero.  The Director believes this to be a problem with the GST WAFS/MetLab.  The Director has asked GST to investigate the problem.

The Sutron AWS automatically sends an observation every 3 hours.   The Contractor thinks this report is only used by the USGS but this cannot be confirmed.

Suggestions

GCOS must work with the SNET staff to establish a reliable and accurate CLIMAT program.  This should include setting up office procedures for recording data, assimilating and cataloging data and the actual creation of the CLIMAT message.  The Contractor recommends GCOS provide on site support and training.

GCOS should consider whether the use of one set of instruments for the synoptic current temperature/humidity and the use of different instruments for the reported max/ min temperatures is a problem for researchers.

The RM Young wind system should be calibrated.  This is one of three RM Young that are used in the 3 CAC sites.

Barometer comparisons with precision standards should be performed.

Mexico 9 GSN Sites

MEXICO GSN SITE PERFORMANCE 2001 THRU JUNE 2004

The following is based on information obtained from the NOAA site;

http://www1.ncdc.noaa.gov/pub/data/gcos/GSN_sum_200101_200406.txt

2001 thru 2003			Jan –June 2004

MESSAGES RECEIVED							MESSAGES RECEIVED

Site			Daily	      	Hourlies	Synoptics 	CLIMAT	   Hourlies 		Synoptics	CLIMAT

Synops		Monthly	Monthly	Received	    Monthly		Monthly	Received

Expect’d		Average	Average	 	   Average        	Average

Choix Sin 76311		7		100	100	15	   100	   	   100	  4

Monterrey 76393		8		200	200	31	   150		   150	  3

LaPaz       76405		8		175	175	35	   140		   140	  6

Colonia J. 76458		8		170	170	25	   150		   150 	  4

Guanajto  76577		8		175	175	31	   200		   200	  6

Merida     76644       		8		850            100-200	21	   850	  	   160	  3

Manzan’   76654		8		175	175	31	   160		   160	  5

Mex City  76680		8		50	50	17	   40		     40	  1

S.Cruz     76833		8		0	0	0	     0		       0	  0

The number of hourlies taken can vary considerably depending on whether the site is an airport and if they are open 24 hours per day.

The number of synoptics received monthly should average between 210 and 248  depending on hours of operation

The number of CLIMAT reports should be 12 per year or 36 for the period 2001 -2003.

The monthly averages are estimated and not done arithmetically.

Strategy and Estimated Budget for Mexico

There are 9 GSN sites scattered over Mexico and 1 GUAN site in Manzanillo which is also one of the 9 GSN sites.

Intermet, Inc.  feels it is of primary importance to meet with senior officials of the Mexican Meteorological Service to explain the purpose of the GCOS maintenance project and how the project can benefit Mexico as well as the overall GCOS program in North America.  Intermet, Inc. recommends the a WMO Representative at the Subregional Office in San Jose, CR accompany Intermet, Inc. to Mexico City MMS Headquarters and take part in the discussions.  It is envisioned that one full work day in Mexico City will be sufficient for this task.

A purpose of this meeting is to explore ideas for providing support to their GCOS GUAN/GSN sites.  Obviously it will be very expensive to dispatch experts from other surrounding countries to provide this support and assistance.  And the MMS may not be feel comfortable having experts from other countries investigating and reporting on their activities.  MMS has the technical expertise to maintain or improve the integrity of the GCOS sites in Mexico.  However the GCOS high expectations in terms of reliability and quality may cause undue strains on the MMS budget. These additional costs include increased expert travel, replacement parts, test equipment and spares, training, overtime salary as well as other unforeseen costs.  Intermet, Inc. proposes exploring the idea of providing funds directly to the MMS to offset the costs of the GCOS efforts in Mexico.  Intermet, Inc. feels this will be significantly less expensive than other alternatives and will allow more funds for equipment replacement, spares, test equipment and other items of need.  Intermet, Inc will work very close with WMO Subregional Office CAC and with MMS to manage this effort.  Intermet, Inc. will also consider contracting directly with experts in Mexico to perform the required activities.  However, this idea will probably not be as attractive to the MMS.

In an effort to conserve funds, the Intermet, Inc. Representative will travel to the GUAN/GSN observatory in Manzanillo and perform the initial mission.  The Intermet Rep will also perform the Mexico City GSN initial mission at this time.

Mission Costs		Intermet Rep.		$4500.00

WMO Rep.		  1700.00

Total		$6200.00

The following costs assume the missions begin in Mexico City and the experts are provided by the MMS.  Also the missions are for single locations because flying or finding transport from one Mexico GSN site to another GSN site is not practical and is very expensive.  Also the costs assumes the expert will spend two full working days on site and one day in travel.

GSN site				Airfare	Perdiem    Misc.	Total

76393  Monterrey			300		630	       100	1030

76405   LaPaz				500		567	       100	1167

76458  C.J.Carras Mazatlan		425		558	       100	1083

76577	Guanajuato (Leon)		250		501	       200	  951*

76644	Merida				400		504	       100	1004

76833	Salina Cruz (Huatalco)	400		645	       300	1345**

76311	Choix Sin			?		300	       300	  600***

________

Total			$7180.00

*      Cost of transport from Leon to Guanjato is unclear

**    Cost of transport from Huatalco is unclear.  2 ½ hours drive

***  At this time it is unclear how to travel to Choix Sin

The total costs for the above :						   $13380.00

Individual GUAN Site Reports Follow

BARBADOS GUAN SITE 78954

The upper air observatory has been in the same location at Grantley Adams airport for more than 40 years.  The site performs the 12Z  observation during the entire year.  In addition they do the 00Z observation during the Hurricane Season.

Originally the station used a SCR-658 “bedsprings” system  which was replaced by the GMD1A and then the ART-1 system.  The current IMS 1500c was installed on April 27, 2003.   Two or three modifications to the basic software was been used to overcome some difficulties.  The current system configuration is   MetOS ver. 1.12 , Version 6000049-0112.

The observation is transmitted in two parts: first is the TTAA, TTBB, and PPBB or up to the 100 millibar level.  The remainder of the observation is transmitted subsequent to the termination of the flight.  The messages are emailed from the radiosonde office to the NMC where it is put on the Corabor system for insertion onto the GTS.  They system was designed for the messages to go directly from the upper air system to the Corabor but it doesn’t work.  This does not present a problem or cause any type of delays unless the telephone connection or internet is not functioning.

This observatory is very reliable and seldom misses an observation.  They are using 600 gram balloons and Sippican B-2 radiosondes.  The 12Z flights usually terminates above 50 millibars and about half of the monthly flights get higher than 10 millibars.  The GCOS goal of 5 millibars is very seldom attained.  The contractor recommends the GCOS program use larger balloons (800 gram or larger) to routinely attain heights above 5 millibars.

The IMS 1500 was designed and constructed to be maintenance free and repair effected by replacement parts.  The station technician indicated there is no need for tools, test equipment or spares for the system.  Apparently NOAA has some type of maintenance contract with the vendor for this system.  However the Barbados Met. Service has not been advised if there is a contract and what the contract includes and how long it will be in effect.  The Director of the Barbados Meteorological Service should be consulted on the terms of the contract and be fully informed on the content of the contract.  It is very difficult for the Director to plan the budget and to justify unexpected costs without knowing the full content of the maintenance contract for his equipment.  The contractor does not recommend spares, tools and test equipment for this system.

The Proton Hydrogen Generator Model 4X100 1058 is functioning reliably and is a welcome relief to the old Stuart generator.   Early on the generator caused problems which were corrected and now the generator is functioning satisfactorily.  Barbados has the same power supply problem that other Proton Generator sites are experiencing.  It appears the power supply units are sensitive to power surges and are easily damages during lightning storms or main power fluctions such as when the airports switch to the backup power generators.  There is a master and a slave power supply for the generator.  The Contractor recommends the station be provided with a spare slave power supply (180) which costs approximately $3000.00 US.  Reconditioned 180s may be available for around $1500.  If the funding exists a spare master power supply (176) should be provided at about the same cost.    The Proton maintenance manual calls for a calibration every 6 months.  However the calibration kit was not provided with the system.  The contractor recommends the calibration gas-kit ($314.00) and a clamp-on current meter D.C. amps, multimeter ($545.00) be provided at each site.

This upper air facility does have a target antenna which should used to check the elevation and azimuth angle prior to and subsequent to the flight.  These target antennas are leftovers from the old days of the GMD and ART systems and need conversion to the proper frequencies for the IMS1500C system.  The GCOS references on good observing practices and procedures indicate the elevation and azimuth angles should be checked for accuracy prior to and subsequent to every flight.  Target antennas should be provided to the other sites and routine elevation and azimuth angle checks should be performed just before each flight.

The surface observation taken at release does not use the same instrument field or instruments as the GSN observations.  The instruments used for the radiosonde observation are located across the runway from the GSN instrument field.   A Simpson Screen and 13foot wind mast is located beside the radiosonde office.  The observers get their wet and dry bulb temperatures from Casella liquid-in-glass thermometers in the Simpson Screen.  The wind system is a RM Young 3 cup anemometer Model 06201 s/n WT 09233 with dial displays in the radiosonde office. The wind system has not been calibrated in recent memory. The aneroid barometer is located in the office and is an old Weather Bureau Model CAT No.G122, s/n 7.67.   It is doubtful if this barometer has ever been calibrated or compared to a high precision barometer.  This should be done !!

The base electric power for the upper air facility is always at risk at the present location.  This is because of the distance of the observatory from the power mains at the main airport and because of the continual construction at the airport.  The Director has requested support to purchase a portable power generator to act a as backup power for the  radiosonde facility.

CLIMAT TEMP MESSAGE

Barbados does not generate or transmit a CLIMAT TEMP message.  This was discussed with the Director and remedies to this shortfall were discussed.   The Contractor believes the staffs at sites that have never done the CLIMAT TEMP report should have some training in organizing the raw data needed to construct the CLIMAT TEMP message as well as detailed instruction in actually creating the message.  Two or three meteorological technicians at each site should be trained.  A good practice would be to have the students create CIMAT TEMP messages for the past year. These messages may be of value to the GCOS and other climate centers.

CONCERNS AND CONSIDERATIONS FOR UPPER AIR OBSERVATIONS

Items 1-7 below are problems with the Imetos program that range in importance from bothersome to very important.  The Contractor has sent this information to certain other sites to see if they are experiencing the same problems.  At this time, the Contractor is reluctant to contact InterMet Systems about these problems without knowing if the other sites are having these problems.

1.  Occasionally the Imetos program shuts down (crashes) during the flight.  This is an intermittent problem and does not seem to occur with any particular set of circumstances.  This leads to an even more serious problem with the system.  The observer will have to rework the sounding to get a message for however high they got before the crash.  The results of the rework are very different than the data and message in the original calculations.  I have attached a sample of a flights original data and coded message along with a rework of the exact same flight.  The differences are obvious and extreme.  This should be brought to the attention of Intermet Systems or whoever is responsible for this program.

2.  Occasionally negative RH values will be displayed during the flight.  However the coded message doesn’t seem to be effected by this display and codes the RH correctly.  This intermittent event does concern the observer and gives rise to doubting the accuracy of the observation.  However there is no software function for the observer to check these data.  Again this should be brought to the attention of the person or company that is responsible for this program.

3.  Occasionally the tropopause level is not in the coded message but it is displayed in the standard level report.  The observer must be alert to this omission and manually insert the tropopause date into the coded message.  The solution is the same.

4.  Occasionally the maximum wind level is not in the coded message but it is displayed in the standard level report.  The observer must stay alert to this omission and manually insert the maximum wind level in the coded message.

5.  Occasionally after the balloon has burst and is falling but still transmitting data, the incoming data will replace the existing data for the maximum wind.  Once again the observer must stay alert to this problem and make the necessary corrections.

6.  The actual time and date for the 00Z observation (YYGG) can be misleading.  An example is that if the observers are performing the April 15th 0000Z observation and release the balloon on April 14th at  2315 Z then the YYGG is coded as 6523 which indicates an observation on the 15TH  at 2300z.  Hopefully the computers recognize this error.

7.  A most disturbing problem with the program is that height errors sometimes occur for certain mandatory levels.  Once again this is an intermittent problem that doesn’t seem to be related to any particular set of circumstances.  The contractor has attached the actual coded messages for the 12Z flights beginning April 15 thru April 22, 2004.  The mandatory levels that show significant departures from expected heights are the 925, 850 and 700 millibar levels.

The following table presents a summary of the heights and temperatures for those 8 consecutive days.  Note the radical differences in heights on the 19th and 20th while the temperature seems to be consistent with previous soundings.

Date 12Z 925 Millibars 850 Millibars 700 Millibars Height Temp Height Temp Height Temp

April 15, 2004 823 m 20.4c 1556m 16.8c 3182m 7.4c April 16, 2004 837 20.6 1546 16.8 3175 7.8 April 17, 2004 832 19.8 1553 16.0 3186 8.2 April 18, 2004 844 21.6 1565 16.4 3198 8.6 April 19, 2004 876 20.2 1636 15.6 3269 9.2 April 20, 2004 791 20.0 1529 16.0 3164 9.6 April 21, 2004 810 21.0 1538 16.4 3174 8.8 April 22, 2004 817 20.0 1549 17.8 3188 10.0

The contractor will attempt to determine if this same intermittent problem occurs at other sites using the same system and software.  If so, the manufacturer should be alerted to the problem and asked to make the appropriate corrections to the system and World Centers should be advised of the errors in past observations.

8.  NCEP and ECMWF have indicated that the upper air winds reported by Barbados have a wind bias of approximately 15 degrees.  The contractor will work with Barbados to determine the cause of this bias and attempt corrective measures.   Earlier this past year the Dominican Republic and St Maartens had the same type of wind bias in their reports which was corrected.  The Director advised the Contractor that a surveyor will determine true north at the upper air facility and the electronics technician will work with the Manufacturer to correct the alignment.  The Director indicated he will alert all the centers to the correction.

9.  The Vector Group vbvbvava   is not clearly understood by many of the observatories and is often not calculated or included in the messages.

BELIZE GUAN SITE  78583

When the upper air observatory at Isle de Cisne (Swan Island) was decommissioned in 1980 the NOAA/NWS moved the equipment to Belize.  The GMD1A and supporting equipment was installed at  Phllip Goldston Airport and the Belize 78583 GUAN site was established.   The GMD1A System was replaced with the ART1 system in the mid-1980s.  The IMS 1500C was installed and began operation March 19, 2003.  The Proton generator was installed in roughly the same period time.

The staff routinely records the performance of each flight using the old WS-Form B-29.  (see attachment)  This is an excellent record that GCOS should promote for use at all the GUAN sites.

The average heights attained at Belize are about the best in the CAC.  During the past year or so they have been getting higher than 10 millibars 15 to 20 times per month.  However they never exceed the 5 millibar level.

Receiver and Pedestal in Dome

The Chief Electronics Technician at the observatory reports that the IMS1500C receiver and pedestal in the dome works very well and is reliable.  He also indicates there is very little need for special tools and spares as the system works low voltage.  He reports that keeping spares on site will not be cost effective but that training in the full system operation would definitely be a positive.  Their full knowledge of the equipment is what they read in the manuals.  He has indicated that the true test of their ability to correct problems will be when they have a failure, which is not the correct time to learn.

NOTE: The above was reported before the system failure on September 2, 2004

The best description of the IMS 1500C failure as well as the shortcomings of maintenance support was written by the Belize Senior Technician and is attached to this document.  The staff is not sure if NOAA has a maintenance contract to support their IMS 1500C or the Proton Generator.

iMetOS Compaq System

The iMetOS system will not automatically transfer the Temp message to the MetLab system for transmission to the GTS.  The observer has to print out each of the Temp messages and then hand type them into the MetLab for further transmission and insertion onto the GTS.  Obviously this method promotes a lot of errors.  They have tried “cut and paste” and other ideas to make this a more automatic and less prone to errors operation.  They notified GST about this problem several months ago and were advised that a new CD was being mailed that will solve the problem.  As of this writing they have not received the CD.

Belize’s iMetOS, has from its delivery had some issues that were evident with relation to the message time generated in message report with respect to launch time. It could not recognize 00Z or 12Z transmission times if launch time was 30 minutes either side of the official flight time. The messages were manually edited to reflect the correct flight time.

Old operating procedures governing launch times gives a larger window for official flight time acceptance, and maybe with the advent of the newer automated operation of upper air observation, we need to officially adjust our launch time. It was suggested to our staff and most observers have adjusted to these smaller windows with no adverse effect. It does however make a second release tight.

Also, on occasions, the staff has observed negative RH in the raw data, but it was not reported in the message. Also there are occasional glitches that make the RH go to 1%, and then it works its way back to more accurate values.

It has happen on a few occasions where if the operator was not present, balloon burst were not detected and the balloon would be tracked to ground and the descending data used for message coding. This however can be accurately recovered.

Locking up of the operator interface has occurred on a few occasions but rework has always brought back the original data. We did eventually reload the distributed OS and solved that problem.

We have not had any reports from our forecasters that our reported heights have been anything other than acceptable, but we have had occasions where the flight clock starts before actual launch and it was noticed before launch, and a reset could have been performed. I am not sure if this may cause erroneous calculations because of timing.

As indicated in other site reports, the staff does not have complete confidence in the MetLab satellite system.  In addition to sending the Temp messages over the MetLab they also send it via AFTN to the regional center in Honduras (COCESNA) where it is sent to Washington DC.  They have confirmed the AFTN route to be the more reliable of the two systems.

Proton Hydrogen Generator

The generator works well and has been reliable for the short time it has been in operation at Belize.  The generator seems to be easily affected by lightning and they have begun the practice of isolating the system by turning it off when thunderstorms are in the area and during the night when staff is short at the office.  The system is only in use during the day hours.  This practice is bothersome to the staff but it seems to work.   The Contractor is working to find a replacement brass regulator valve Victor Equipment HG48915.  The storage tank is usually at 200psi and has enough gas for roughly 8 @600 gram balloons.

The electronic technician reports the water filter list below for both systems can keep them operating for a year with one of each item. The quality of the water supply does not warrant more frequent change than yearly.

HOGEN 40

- F210	40µ Stack Inlet Filter			Pall/Gelman PCF1Y100N04

- F506	10µ DI Water Post Filter Element	Pall/Gelman PCF1Y100N04

- GB208	Guard Bed Polishing Cartridge	MSC Liquid Filtration MR-1A-208

Aqua Solution H40

- 10” 2µ Activated Carbon Prefilter Cartridge	CC1050

- High Capacity DI Module			2613PT

These items for the Aqua Solution may be had as a purification kit part number, 2613H40.

The Proton Generator and the Aqua Solution system were delivered with replacement filters but no provisions were made for further support.  They are not sure if they will get support from NOAA Kansas City Support Center.

The staff is hesitant to say the system is very reliable until they have more operational experience.

Expendable Supplies

Belize is routinely supplied with sondes and balloons by NOAA NWS.  As of June 30th they had 151 sondes and 194 600 gram balloons in stock.  The Balloons are Kaysam 600 gram with nozzle lift of 2000 grams.  The sondes are Sippicon 1492-540.

The staff indicated a need for a supply of “severe weather balloons”.  These are fast rise balloons made of a stronger material that are easier to handle in high winds.  They burst before they reach the 100 Millibar level.

Instrument Field and Instruments Supporting the Upper Air Observation

The entire NMHS, administrative offices, upper air facilities and the instrument field are located on the airport grounds but well enough away to avoid any particular influence of airport activities.  The instrument field is grass covered and the instruments are ideally situated in the field.

The wind system is an old Casella W1400 combined wind vane and 3 cup anemometer that has never been calibrated or compared to any standard.  They use a Casella Directional dial 5231.  The staff does not know if a correction is needed but they do routinely perform normal maintenance  i.e. replacing bearings etc etc.

They us a Vaisala Digital PA11A triple barometer that seems to be reliable but it has never been compared to a standard or calibrated.  The station uses 5 meters as its height above msl.

The psychrometers and max/min thermometers are in a standard white Simpson Screen type instrument shelter.  The Casella dry/wet bulbs are model M106001 and the Max is M106202 and Min 106204.

The Official rain gage is a standard splayed base Casella 114003.

CLIMAT TEMP REPORT

Belize has not done the CIMAT TEMP report since the Fall of 2003 because they lost the person responsible for report.  They are willing to resume this report but need training.  The Contractor gave them the CLIREP CD which they decided was not adequate to train their staff.

Each month the staff sends the raw data (FTP) from each sounding to NCDC Attention Stuart Hinson.  The Contractor is trying to determine if this is how the mysterious CLIMAT TEMP reports are generated and circulated.

Additional Notes and Recommendations

The Director is requesting support to procure and install a complete automatic weather station to replace the above instruments.  Their preference is the Vaisala Milos 500 systems.

The Contractor will attempt to get GTS to supply the needed CD for the transfer of information from the iMetOS to the MetLab system.

The Contractor strongly recommends that NOAA fully informs all the GUAN sites of the extent, scope and terms of any warranties and of the terms of any maintenance contract they have covering the IMS 1500Cs and the Proton Generators.

Training in the creation of the CLIMAT TEMP message is needed.

Attachment

Belize Rawinsonde Workstation

Belize Upper Air station was upgraded with the new InterMet IMS-1500C Radiotheodolite on 19 March, 2003. This system replaced the ART-1 upgrade from the GMD system.

The IMS-1500C with its more modern Pentium 4 computing system brought a more familiar windows atmosphere to the operators of the system. From an operator’s point of view, this was a great improvement in the one man operation of the Rawinsonde flight. The training of the operators to handle the equipment and new interface were done with great ease and the transition was flawless.

From the technician end of things, it presented a completely different approach to our support of the system. There were no specific maintenance information transferred to the technicians, and it was said that basically the system would need no preventative maintenance. Keeping it clean and obstruction free were basically the extent of what local technicians were expected to do. No troubleshooting and fault recognition training were ever carried out. This I see as a big flaw in the scheme of things.

My greatest fears were realized when our system froze during the 12Z flight of 2 September, 2004. Checking the log for Rawinsonde flights, it was discovered that both on 24 and 28 August, two releases were made due to antenna locking up during the first release that did not make it to a successful altitude.

Hoping that it might be something simple in the antenna pedestal that could solve the problem, we first did a software reload of the controlling computer interface to rule out operator software problems, as I can remember that earlier in the breaking in period of the system, a reload had cleared up some glitches that were in the operator software. (The exact date of that reload was not noted.)

That did not have the desired effect of solving our problems, so it was concluded that it must be a hardware failure of some sort and since it manifested itself only in the elevation portion of the antenna, we focused on the associated connectors and did continuity test to the cables hoping for a loose connection or a break in one of the cable connecting the elevation to azimuth of the system.

We did those basics and our prayers were not answered the way we wanted, but however they were answered by a call from InterMet, noticing that we were off the air. Our conditions were relayed to them and the dialogue started with InterMet to try and resolve the problem. We needed guidance as we did not know where else to proceed.

We were instructed to initiate flight logger support software that was delivered with the system that is to be used in these circumstances. No documentation on this software was ever delivered with the system at the time of installation, so the engineer had to email the available manual. Now we had to launch that software and do logging of the system info of system initialization and forward that log to the company as we did not have a similar log to compare it to. The flight logger program did however have a display terminal that indicated that System Communication Assembly had a fault and that it failed the initialization test. This however did not seem to be cause of the fault but more likely the elevation motor that was overloading the system with some high current reading on the display. It was the consensus that the motor or the servo drivers for the motor were the likely culprits to our demise.

Now that the likely problem was identified, the solution became the next concern as when I enquired about a replacement, I was informed that the system were no longer covered by warranty and repair at the head office was the only solution. No spares were available for immediate shipment to us, and the matter of the repair cost was not ever worked out or the proper protocols to execute these actions were not set out at this time.

We were informed later after approval for the repairs were acknowledged by NWS in Washington DC, that we will be having a DHL pickup and we were looking at a possible one week turn around. This I know from my years of dealing with shippers and international transfer of equipment that it may more likely be two to three weeks.

During this time of the year, having our system down for a month or more is not readily acceptable to anyone and would not go down well being prepared to deliver timely and accurate advice if we get into any weather related emergency.

This may be an isolated event as far as system failure, but it is most likely that this will occur again here or somewhere else. Provisions must be made to cover these eventualities. Means to reduce down time must be put in place if we are to fulfill our part in this global observance of our atmosphere.

Bogota, Colombia El Dorado Airport 80222

Background

The observatory routinely performs the 1200Z observation and if funding permits they also do the 000Z run.  The NOAA/NCEP publication shows that during the years 2001 thru 2003 the site performed very well and averaged 25 to 30 observations per month.  Their heights averaged roughly 55 millibars per month with none going above 10 millibars.

Communications Problem and Recommended Solution

Bogota has been considered inoperative since January 2004.  The reason is because their daily reports are not received in WashingtonDC at all or not in time for inclusion in any model runs.  The Thanks2 report does not show any reports at all and the NCEP 3 inventories do not include Bogota because of a lack of reports.  However, they have been doing the soundings but they have not been reliably transmitted in terms of timing and method.  A solution to this problem is to arrange for Bogota to use the NOAA EDIS email system the same as others in the Region.  The Contractor has been working very hard to promote this solution but has not met with success.  Application to participate in the EDIS system was made by Colombia in late July 2004.  They followed the instructions found on their Web Page http://www.nws.noaa.gov/tg/bullguid.html.  Since that time several attempts to get approval or even a  response from NOAA have failed.  Contractor feels that senior management in NOAA should be made aware of this effort and asked to cooperate with this effort.

Equipment Problems and Possible Solutions

An additional problem at Eldorado is the lack of winds being measured during the sounding.  As you know, the Vaisala system requires two antennas.  In simple terms, one antenna received the temperature, humidity and pressure measurements (PTU).  This works very well.  The other antenna is the GPS antenna and is used to locate the instrument allowing the observers to measure the wind velocity.   This antenna (GPS)  is ineffective because of its location.  There is too much local interference which does not allow the antenna to function properly.  Vaisala representatives have visited the location twice and have indicated the large numbers of automobiles, trucks, buildings and other iron bearing structures are causing the interference. Bogota even brought the system from Rio Hache to El Dorado and did some tests and that system had the same interference

problems.  As a note: Vaisala said they are having the same problem in Chile and Brazil.  There are two obvious solutions.  The first is to relocate the antenna or entire observing site which is impractical and probably too costly.  The other is to use the RS-92 radiosondes that do not seem to be affected by the interference and have been tested successfully at El Dorado.  The Contractor was advised that Colombia still has the modification needed to convert the system to using the RS-92s.   If the communications problem can be solve, the Contractor will petition GCOS authorities to provide RS-92 sondes for the site.

Expendables

For many years Colombia has been willing and able to operate their upper air sites without much external support.  They currently operate and fund the upper air observatories at Leticia, Gaviotos, Tumaco, Rio Hache and Bogota.  They also operate the upper air observatory on San Andres Island in cooperation with NOAA.  Their funds are spread rather thinly over these sites and they do not have much financial operating room.  There are new administration changes on the horizon and they expect some rather severe budget cuts.  On June 28th 2004 they had enough  RS-80 sondes to last them until early 2005.

The RS-80s cost Colombia about $175 @ and the RS-80s 165.  They are using Totex 350 gram balloons at a cost of $22.00 apiece.

The good news is they buy bottled helium.  They rest of their observatories used hydrogen generators.  They are spending about $10K annually on helium which is about 120 cylinders.  They claim they can fill 3 balloons (350 gm) with each cylinder.

Surface Observations in Support of the GUAN Operations

The use a Sutron Modelo 8200 AWS for their observations at release of the balloon.  This is a standard Sutron that uses Radio UHL to send the observation to the upper air office.  None of the instruments have ever been calibrated or compared with any standard.  The barometer mounted on the tower is a Sutron 5600-0120-2, s/n 968424.

Monitoring

The Contractor showed the upper air staff how to monitor the real-time receipt of the Temp message at Washington DC by introducing them to the THANKS2 page.  Please note the THANKS2 page is not very clear and difficult for a Spanish speaking person to understand.  The Contractor had the page translated into Spanish and gave all the Spanish speaking countries a copy of the page. See attached.  The contractor also advised the observatory how important their observations are to the rest of the meteorological community and showed them where to find the various centers monitoring reports and products, including the NOAA Inventories.

The Contractor also suggested they release the balloon as early as allowable which is 45 minutes before 00Z or 12Z.  This will give them more time to get the first transmission to the Centers on time for the early model runs.

Attachment

# Reporte actualizado "Thanks 2" de las Observaciones de radiosondeo

[Back to QAP Home] [Previous Raob "Thanks 2" Reports]

Observaciones de radiosondeo recibidas en el Centro Nacional de Predicción del medio ambiente (NCEP, siglas en inglés), válidas para el día 30 de junio del 2004 a las 14:00:01 TMG.

Esta página es actualizada cada 15 minutos entre las 0:45 y las 02:00 y entre las 12:00 y las 14:00.

76151      76225 THKS 76256 ABc  76394 THKS 76405 THKS 76458 THKS

76526 ABc  76595 a    76612 a    76644 THKS 76654 THKS 76679 THKS

76692 THKS 76723 THKS 76805 AB   78016 THKS 78073 THKS 78355

78367      78384 THKS 78397 ABc  78486 THKS 78526 THKS 78583

78641      78720      78730      78762 THKS 78808      78866 THKS

78897 A    78954 THKS 78970      78988      80001 A    80035

80222      80241      80398      80413      80447      80462

80476

“A” significa que se recibieron los datos del grupo A (TTAA) completos. “B” que se recibieron completos los grupos TTBB y PPBB. La “C” que se recibieron completos los grupos TTCC, TTDD y PPDD. “THKS” significa que se recibió el mensaje completo, con datos en todos sus grupos, para esa hora específica.

A, B o C minúscula significan que se recibieron los datos pero están incompletos para ese grupo específico.

Espacio en blanco significa que no se recibieron los datos para esa estación en particular.

Este mensaje refleja los datos recibidos  en el NCEP y el total de los datos procesados por el sistema para esa hora.

Por favor retransmita los mensajes no recibidos. REENVÍELOS UNA SOLA VEZ, NO REALICE RETRANSMISIONES REPETITIVAS.

Revise SDM y el administrador de mensajes Gateway cuando sospeche que existen problemas de comunicación.

Para preguntas relacionadas con este resumen escriba a  raob report specialist. 
Para preguntas acerca de este sitio de Internet escriba a  Webmaster. 
NCEP Grupo de Soporte al Usuario/ última modificación: 30 de junio, 2004

San Jose Costa Rica  GUAN Site Juan Santa Maria Airpot  78762

Background

The upper air observatory at Juan Santa Maria Airport was established in 1972.  Their original equipment was the GMD using the VIZ 1680 sondes.  NOAA provided considerable VCP support to the observatory in terms of expendables and technical expertise.  NOAA had an electronics technician stationed in San Jose to support the observatory as well as other observatories in Central America.  In 1992 they changed from the GMD system to the Vaisala MW11 using RS80 GPS sondes.  The GPS sondes were too expensive for NOAA to continue support on any regular basis.  Funding and procurement problems led to long periods of time without sondes which means they did not perform the observations.  The Director has indicated that the internal bureaucratic procedures associated with procurement are the primary reasons for their inability to acquire the expendables.

WMO Volume A indicates San Jose 78762 is to perform one observation per day at 12Z.  The NCEP monitoring report shows the station was silent from October 2001 thru May 2003.  It was silent because they did not have sondes and they were having problems funding personnel to perform the observations.  A brief summary of their performance follows:

Rough Rough Average Average Soundings Heights Above Above Years Per Month Per Month 10 Mbs 5 Mbs 1995-2000 15 to 25 25-75 Mbs seldom none Mid 2000-2001 25 to 30 65 Mbs none none Late 2001-mid 2003 0 0 0 0 Mid-2003 – current mid 20s` 45 – 100 Mbs 0 0

It is obvious that their average monthly heights are terrible.  They should be averaging 10 to 15 Millibars per month.  However the use of 300 gram balloons keep the heights lower.  The contractor did check the station records which shows that nearly all the terminations were due to balloon burst. Stations that use a 600 gram balloon routinely have higher averages with more flights exceeding 10 Millibars.  Key West averages 10 to 20 millibars per sounding and nearly always exceeds 10 Millibars.  Barbados averages in the low 30 millibars and gets above 10 Millibars more than ten times per month.  GCOS should consider providing 600 gram balloons to some of the Vaisala sites.

Hydrogen Generator

They have the newer model Stuart Generator with 5 Cells.  The keep the system in A-1 condition and it operates very reliably.  The have two large hydrogen storage tanks that allow for multiple releases in case of problems.  The older electrolyser is still in the inflation shelter and can be used if the new generator fails.  They will be participating in a experiment with NASA where they will be doing 4 flights per day which will really test the endurance of the generator.  The Contractor advised them to acquire a few spares for the generator particularly spare compressors and extra water filters.  They use 300 gram balloons.

Expendable Supplies

In July 2004 they had a stock of 353 RS80-15G sondes and roughly 350 300gm balloons.  As indicated above they will begin and experiment with NASA during the Summer of 2005 where they will perform 4 soundings per day.  During this past July thru September 2004 they did practice soundings in anticipation of the next Summer’s experiment.  NASA is paying for the expendables which are RS-92s.  Therefore a modification had to be made to the Vaisalla  MW11 to accomodate the RS92 sondes.  Apparently this modification is not very difficult install or remove from the system.  San Jose will be using RS80s  from September  till the experiment begins in the Summer of 2005.  The Director did advise the Contractor that the modified MW11 can use either RS80 or RS92 sondes.

Communications

They send the Temp to Washington DC via the NOAA EDIS system.  The do not have full internet capability at the radiosonde office and are often without communications to Washington DC.  The Contractor worked with the Director and the local WMO office to provide funds for full time internet-email capability at the radiosonde office.  The Contractor provided $540 of GCOS funds for 18 months full internet capability.

CLIMAT TEMP Message

San Jose discontinued the CLIMAT TEMP program in July 2001 because the communications link with Hadley Centre failed.  The Contractor worked with both parties and resolved the issue.  Hadley Center should begin receiving CLIMAT TEMP reports from San Jose in October 2004.  The Contractor has asked the climate staff to create CLIMAT TEMP reports for all the months they missed since July 2001.

Surface Instruments Supporting the Upper Air Observation

The observatory uses a Campbell Scientific CR-10X Automatic Weather Station located approximately .6 miles from the balloon release area.  The instruments are mounted on a tower and the tower is several meters above the level of the radiosonde office where the pressure comparisons with the sonde are made.  There is always a substantial difference which may lead to a linear error in heights particularly when extrapolating down to the 900 and 1000 millibar level.  Remember the height of the observatory is considered to be 921 meters above msl.  The Contractor discussed this with the staff and with the Director and concluded they needed a precision barometer at the radiosonde office for exact comparisons with the sonde.  The Contractor purchased a Honeywell Precision HPB200F2DA-A and shipped it to San Jose for their use at the office.  The total cost of the Barometer and shipment to Costa Rica was $932.

The AWS has Vaisala CS-500 temperature and humidity sensors which are considered to be disposable and are changed periodically according to manufacturer specifications.  This is considered to be cheaper than a calibration scheme.  The barometer mounted on the tower PTD-101B.

The wind system is a RM Young 05103 which is the same as the observatory at Pto Limon.   The manufacturer recommends a service interval of every two years which includes basic inspection and verification of signal outputs.  Other organizations like NRC Canada require these instruments to be calibrated every six months.  RM Young has quoted a price of $1134.00 for the equipment and instruments needed for calibration.

.

Kingston, Jamaica GUAN Site 78397

Background

During the IGY years 1957-58, a GMD system was installed at Norman Manley Airport and the USA began providing support for this observatory.  During the late 1990s the GMD was replaced by a Loran C system which was replaced on May 14, 2003 by an IMS 1500C along with the new Proton Hydrogen generator.   The facility has been in approximately this same location for nearly 40 years.

Kingston performs the 12Z sounding year around and the 00Z observation during the Hurricane Season.   Prior to the installation of the IMS 1500C their average monthly  heights were in the 20 to 30 millibar range with 10 to 20 above 10 Millibars but seldom above 5 Millibars.  During the 12 months following the installation of the IMS 1500C their average monthly heights dropped significantly to between  220 and 120 Millibars and only exceeded 10 Millibars  twice during that period.  The Contractor believes the lack of experience and training in the system was responsible for the poor performance.  Since April 2004 the heights have increased to averaging 45 Millibars per month.

CLIMAT TEMP Reports

They do not prepare or submit a CLIMAT TEMP report.  The Director has indicated a willingness to begin this program if her staff can be properly trained in the assimilation of the data and creating the CLIMAT TEMP code.  Hadley Center in the UK advised the Contractor they received CLIMAT TEMP reports from Jamaica 78397 over the past 4 years as follows: 11 in 2001, 10 in 2002, 5 in 2003, and 4 so far in 2004.  The Director is completely unaware of these reports and would like to know who is doing them and why she hasn’t been informed.

IMS 1500C

Receiver and Pedestal in Dome

The station has experienced a lot of problems with loss of signal which ususlly causes an early termination of the sounding.  If the signal gets weak or diffused the antenna does not search for the signal in the approximate direction of the sonde.   Instead the antenna drives downward to nearly horizontal or zero degrees elevation.  The observer has to move the antenna back to the approximate position of the airborne sonde and try to lock on to the signal.  This usually fails and the flight is terminated.  Often the observers is busy at some task that keeps them from giving the flight 100 percent of their attention.  If the signal is lost for a few minutes during these times it is almost a sure early termination.

The local winds aloft at Kingston tend to move the balloon directly over the antenna during the flight.  This is the primary cause for the loss of the signal at Kingston.  An example would be as follows:  As the balloon drifts from due west to directly overhead and then to due east, the antenna has to rotate 180 degrees on the pedestal and intercept the balloon in its easterly position. The antenna does not react properly and instead drives to zero elevation.  The Contractor has not found this problem at other sites.

The electronics technicians are satisfied with the performance of the IMS 1500C system.   They know that NOAA has a maintenance contract to support the system but they do not know any details of the contract or what they can expect in the future.  As with the other sites, the technicians do not feel they know enough about the system to perform any simple repairs or corrections or to monitor the systems performance to identify any early signs of failures or problems.  They want training for 3 technicians in the total IMS1500 C system.

They do not have a target antenna to check the elevation and azimuth angles prior to and after the flights.  The GCOS Best Practices calls for these checks for each flight.

iMetOS Compaq

The iMetOS is connected to the Corobor system via a dedicated telephone line.  If the telephone connections fail then they copy the observation to a floppy and carry it to the NMC for insertion onto the GTS.  The Contractor feels the NOAA internet/email system EDIS would be easier to use and more reliable.

Expendables

Kingston is using the Sippican 1492-540 sondes and 600 gram balloons.  These are routinely provided to  Kingston by NOAA/NWS.

Hydrogen Generator

The main and slave power supplies for the Proton Generator often fail because of power surges caused by nearby lightning or by the airport switching from main to emergency power sources or just an unstable power supply.  The Kingston Proton Generator was not in operation in July/August because of power supply failures.  They received new power supplies just before the passage of Ivan and were able to resume operations.  However, during Ivan the airport had to switch to it’s own back-up generators which caused a power spike which dmamged the power supplies once again.  They have recently ordered new power supplies from the Proton.  When the generator is inoperative the NMHS has to purchase very expensive bottled hydrogen which the Director indicated cannot continiue.  The technicians indicated they need additional micron and carbon filters for their water filtration system including the Aqua Solutions equipment.  They cannot purchase these filters locally.  The Contractor recommends the following for the Proton Generator:  Spare Master and Slave Power Supply  units, additional water filters ( 3 types),  and spare internal water pumps.  The technicians do not feel comfortable in recommending spares or judging the performance of the generator until they have more experience with the system.

Surface Observing System Supporting the GUAN Site

The surface observing equipment that supports the upper air observation is an automatic station on 10 meter tower located on the grassy area between the paved runways.  The automatic station is a Casella Ltd.,  Model DAU16, Eprom version 80-120836A-09, serial number 031093.  The configuration is M103107B.

None of the sensors including air pressure, wind velocity, temperature and humidity have been calibrated for several years.  The observation is provided to the upper air facility via telephone from the NMC.  These are the same instruments as the Montego Bay installation.  This automatic system is subject to the ocean breezes and the salt content.  This does cause considerable corrosion causing the instruments to be replaced every few years.

Manzanillo, Mexico GUAN/GSN Site 76654

The Contractor has not performed the initial mission to any of the Mexican GSN or Guan sites.  The Contractor met with the MMS radiosonde staff in March 2000 and reviewed the status of all 17 upper air observatories in Mexico.   With regard to the Manzanillo 76654 observatory he was advised that MMS replaced the GMD1A with a Digicora on January 10, 1998.  They have an older Stuart Hydrogen generator but it was performing well at that time.  The staff indicated they needed some spares and replacement parts. They were routinely performing the 12Z observation using 350 gram balloons.   The Vaisala Digicora provides a print-out of the upper air message which the observers re-type on a separate computer which sends the message to the STAR4 system at MMS Hdq.

The NCEP report indicates Manzanillo has been averaging 25 to 30 soundings per month with a very low average height ranging from 50 to 180 Millibars.  They never get above 10 millibars.

Hadley Center reports that Manzanillo submitted some CLIMAT TEMP reports in 1999 and early 2001 and none since 2001