ARM TR-021

Pyranometer Handbook

November 2004

T. Stoffel

Work supported by the U.S.  Department of Energy,

Office of Science, Office of Biological and Environmental Research

Contents

1.	General Overview	1

2.	Contacts	1

3.	Deployment Locations and History	1

4.	Near-Real-Time Data Plots	1

5.	Data Description and Examples	1

6.	Data Quality	7

7.	Instrument Details	7

Tables

1.		6

# General Overview

A pyranometer is a type of radiometer used to measure broadband solar irradiance on a planar surface. Accurate measurements of solar irradiance are needed to understand the primary source of energy input to the earth-atmosphere-ocean system. The ARM Program uses pyranometers for ground-based, in situ shortwave radiation measurements collected by the ARCS/GNDRAD, ARCS/SKYRAD, SIROS, and SIRS instrument platforms. (Note: for additional information on the pyranometer, see the SIRS page.)

# Contacts

## Mentor

Tom Stoffel
National Renewable Energy Laboratory
1617 Cole Blvd.
Golden, CO 80401-3393
Voice:  303-384-6395
Fax:  303-384-6391
E-mail: Tom_Stoffel@nrel.gov

## Instrument Developer

The Eppley Laboratory, Inc.
12 Sheffield Ave.
Newport, RI 02840
Phone:  401-847-1020
http://www.eppleylab.com

# Deployment Locations and History

This section is not applicable to this instrument.

# Near-Real-Time Data Plots

This section is not applicable to this instrument.

# Data Description and Examples

## Data File Contents

### Primary Variables and Expected Uncertainty

Depending on orientation and mounting system hardware, a pyranometer can be used to measure various shortwave, or solar, irradiances (Watts per square meter):

- Downwelling total hemispheric:  An unshaded pyranometer mounted on a stable horizontal platform will provide measurements of the combined direct "beam" and diffuse "sky" irradiance components.

- Downwelling diffuse:  A pyranometer mounted on an automatic solar tracker, equipped with a properly designed shade mechanism to block the direct solar rays, will provide measurements of diffuse irradiance.

- Upwelling diffuse:  A pyranometer installed in an inverted horizontal position, i.e., mounted with the detector pointed at the ground, will measure the amount of reflected shortwave irradiance.\

#### Definition of Uncertainty

Model PSP pyranometers are capable of providing 1-minute averages of down welling global (total hemispheric), diffuse (sky), and upwelling (reflected) shortwave irradiance measurements within +/- 5% or +/- 10 Watts per square meter, whichever is larger, at the 95% confidence level. This estimate assumes the instrument has been calibrated within the last year to the World Radiometric Reference (WRR), properly installed, and is adequately maintained. A detailed analysis of the response characteristics contributing to the measurement uncertainty estimate is provided in the "Theory of Operation" section.

There are a number of conditions under which pyranometer measurements may be incorrect. The data user should examine the data quality flags (described in "Explanation of Flags Applied During Data Ingest) and data quality reports (DQR) to determine whether significant malfunctions have occurred due to problems with a pyranometer. The more frequent sources of error are described briefly here.

Contaminated optical surfaces can reduce or enhance the apparent solar flux at the detector. Accumulations of dew, frost, dust, snow or ice on the outer dome are generally reduced or eliminated by the use of a ventilator or when the pyranometer is mounted in the inverted position. Reduction of the dome transmittance due to contamination will decrease the signal from a pyranometer. Under some circumstances, the presence of dew, frost, snow or ice can enhance the amount of solar radiation reaching the pyranometer detector. In small quantities and in locations on the dome that can reflect more photons onto the detector than available in the direct beam, the signal from a pyranometer may increase from the expected value.

On time-scales of a few minutes, the presence of scattered clouds, especially towering cumulus, can reflect additional solar radiation onto the pyranometer to cause irradiance measurements in excess of the corresponding extraterrestrial value (a physical limit).

Improper alignment of the solar tracking disk can leave all or part of the pyranometer detector surface unshaded from the direct radiance, resulting in high values for down welling diffuse. Interruptions of electrical power to the solar tracker can cause such misalignments. The Model PSP is a passive instrument, providing a signal even under this condition.

The down welling hemispheric shortwave irradiance will be in error if the pyranometer detector is not horizontal. Tilted toward the sun by as little as one degree, the pyranometer output can be noticeably biased.

Local obstructions to sun such as trees, buildings, power lines, or towers can shade or reflect sunlight naturally available to the pyranometer detector. The location of a SIRS measurement station has been selected to eliminate or greatly reduce the effects of such obstructions to accurate irradiance measurements. Generally, deviations of the local horizon within an elevation angle of 5? will not result in measurable differences from a pyranometer.

### Secondary/Underlying Variables

This section is not applicable to this instrument.

### Diagnostic Variables

This section is not applicable to this instrument.

### Data Quality Flags

See SIRS Data Object Design Changes for ARM netCDF file header descriptions.

### Dimension Variables

This section is not applicable to this instrument.

## Annotated Examples

This section is not applicable to this instrument.

## User Notes and Known Problems

This section is not applicable to this instrument.

## Frequently Asked Questions

How do you pronounce "pyranometer"?

Webster's dictionary gives the first pronunciation as "peer`-o-nom-et-ur" and "pie`-rah-nom-et-ur" as an alternate.

Why are nighttime data typically negative?

Shortwave radiometers should produce a null signal in the absence of solar radiation. However, all commercially available pyranometers based on thermoelectric transducers, can exhibit a non-zero output signal in the absence of solar radiation. These signals are a result of thermal gradients within the pyranometer.

The Eppley Laboratory, Inc. Model PSP is a thermoelectric, single-black detector design with well-established performance characteristics [Hulstrom, 1989]. Two Schott Glass hemispheres protect the thermopile detector from the weather. These inner and outer domes also reduce the thermal exchange between the detector, or thermopile "hot" junction, and the environment. The body of the instrument is a relatively massive bronze casting that is used to control the "cold" or reference junction of the thermopile. Recent studies of this and other models of pyranometers suggest thermal offsets are producing as much as 30 Wm-2 during clear-sky, nighttime conditions. Similarly, daytime measurements of down welling diffuse shortwave irradiance must be corrected for thermal offsets. ARM will use a correction method based on correlations with the net infrared and observed diffuse irradiances [Dutton, et al, 2001]. Beginning in 2001, all SIRS diffuse irradiance measurements will be made with an Eppley Model 8-48, a so-called "black and white" detector designed to almost eliminate the thermal offset.

How often should the outer dome be cleaned?

Research-quality irradiance measurements require frequent cleaning of all optical surfaces. Ventilators reduce the need for maintenance by producing a steady flow of air over the pyranometer dome preventing dew and frost formation and, to a lesser extent, accumulations of dust, snow and ice. The frequency and amounts of contamination are site dependent. Some studies suggest the effect can be about 5% for daily total global irradiation [Marion, 19XX]. An analysis of historical SIRS data to determine the effects of cleaning will begin in 2001. Additionally, a measurement experiment comparing duplicate measurement systems with varying cleaning intervals is planned for later the same year.

What is the maximum signal cable length?

Daytime pyranometer signals typically range from 10 microvolts to 15 millivolts, depending on the amount and distribution of clouds. The measurement accuracies of all-weather data acquisition systems is typically ?3 microvolts. Experience has shown a "good" signal cable made of 18 AWG to 22 AWG copper, twisted pair, shielded, and properly grounded can be used at lengths approaching 50 m before effecting the pyranometer signals.

When should the PSP desiccant be changed?

When the indicating silica gel turns pink, the desiccant needs to be recharged or replaced. White or clear silica granules indicate complete saturation and the failure of a seal within the radiometer.

Why is the direct irradiance sometimes greater than the unshaded pyranometer?

Direct normal (beam) irradiance varies with solar zenith angle and the amounts of absorbing and scattering elements in the atmospheric path to the receiver. Global (total horizontal) irradiance is the sum of the direct normal and diffuse (sky) components: Global = Direct Normal x Cos(Z) + Diffuse. Depending on the site location (latitude) and time of year (solar declination), the direct normal can exceed the global irradiance or vice versa.

What is the maximum irradiance expected from a pyranometer measurement?

>Unshaded = Global = 1400 Wm-2 (greater than exoatmospheric) in the presence of towering cumulous or when the solar disk is near the edge of a cloud. 
>Shaded = Diffuse = 600 Wm-2 for mid-latitude, continental locations.

Where can I send a pyranometer for recalibration?

All radiometer manufacturers provide recalibration services, typically using artificial sources. The ARM Program has developed the Radiometer Calibration Facility at the SGP to provide annual recalibration of all SIRS shortwave radiometers (pyranometers and pyrheliometers) using natural sunlight.

How often should a pyranometer be recalibrated?

Depending upon exposure, a properly maintained pyranometer can maintain a responsivity (calibration factor is proportional to the inverse of responsivity) within ?0.5% for up to one year. Annual recalibration of field radiometers is an accepted practice for maintaining research-quality measurements.

Is it necessary to remove a pyranometer from service for calibration?

No. Traveling standards can be used to compare with field instruments, but the resulting calibration uncertainties are greater than using reference standards and the ?summation? technique. Field calibrations require stable, clear-sky conditions during the comparisons. ARM is evaluating the Alternate Calibration Method in which the unshaded (global) and shaded (diffuse) instruments are periodically (seasonally) exchanged at the site. Comparisons of the irradiances at similar solar zenith angles and the coincident direct normal irradiance can produce calibration factors for the two pyranometers.

What is "cosine response"?

The angular response of a perfect pyranometer detector would vary exactly with the cosine of the incoming direct irradiance. Manufacturers typically specify the departure from perfect cosine response at incidence angles less than 85? on the order of ?5%.

What are the effects of ventilators on pyranometer performance?

The forced air over the pyranometer dome helps reduce the thermal gradient between the hot and cold thermopile junctions, thereby reducing the pyranometer zero-offsets.

Does it matter how I install a pyranometer?

The pyranometer must be level (parallel to horizon) within 0.5?

An air gap between the pyranometer base and mounting surface should be at least 5 mm to insulate the instrument from unwanted solar heating and reduce the chances for water to penetrate the gasketed inspection plate in the base of the instrument.

Mounting material should reduce the heat transfer to the pyranometer (gratings work well).

Does it matter where I install a pyranometer?

Provide a free horizon for full solar access all year. Obstructions should be no higher than 5? above the local horizon, far from any highly reflective surfaces (window glass or metal walls). Avoid possibilities of shading by trees, buildings, power poles, roof vents, etc.

Does the amount of circumsolar (forward scattering around the solar disk) effect pyranometer measurements?

Unshaded pyranometers are not affected by the redistribution of direct or diffuse irradiance. Shaded pyranometers will detect the amount of circumsolar irradiance ?leaking? past the shade device, typically occulting a 5? field of view.

What effects do rapidly changing cloud scenes have on pyranometer measurements?

The time constant of the Model PSP is about 1 second for 1/e. The Model 8-48 is 5 seconds for 1/e. Changes in irradiance are detectable if the radiometer time response is fast enough. Radiometers with detectors based on the photoelectric effect are very fast (~10 microseconds), but have restricted spectral response. Thermoelectric (thermopile) detectors have broadband spectral response, but are slower (2 to 5 seconds). When comparing solar irradiance data from different sites or solar irradiance components from the same site, it is very important to know the measurement characteristics of the radiometers, especially time response. SIRS data, based on PSP pyranometers, a NIP pyrheliometer, and PIR pyrgeometers, include 1-minute averages, standard deviations, minima, and maxima for the 2-second irradiance samples. These data attempt to capture the changing radiation fields due to clouds, solar position, and other factors.

What is the difference between measurements made with a silicon detector vs. thermopile detector pyranometer?

At several locations, ARM has collocated radiometers based on these two detector types. Their response characteristics are generally quite different:

Table 1.

| Response Characteristic | Thermopile | Photodiode |
| --- | --- | --- |
| Spectral | 280 nm ? 2,950 nm | 400 nm ? 1100 nm |
| Time Constant | 1 sec (1/e) | ~10 microseconds (10-90%) |
| Angular | ?1% up to 70?, ?3% 70?-80? | 3% up to 80? incidence angle |
| Temperature | ?1% from ?20? to +40?C | ?0.15% per degree C (max) |
| Linearity | ?0.5% up to 2800 Wm-2 | 1% up to 3000 Wm-2 |
| Stability | < 1% per year | < 2% per year |

# Data Quality

## Data Quality Health and Status

The following links go to current data quality health and status results.

- DQ HandS (Data Quality Health and Status)

- NCVweb for interactive data plotting using.

The tables and graphs shown contain the techniques used by ARM's data quality analysts, instrument mentors, and site scientists to monitor and diagnose data quality.

## Data Reviews by Instrument Mentor

This section is not applicable to this instrument.

## Data Assessments by Site Scientist/Data Quality Office

All DQ Office and most Site Scientist techniques for checking have been incorporated within DQ HandS and can be viewed there.

## Value-Added Procedures and Quality Measurement Experiments

Many of the scientific needs of the ARM Program are met through the analysis and processing of existing data products into "value-added" products or VAPs. Despite extensive instrumentation deployed at the ARM CART sites, there will always be quantities of interest that are either impractical or impossible to measure directly or routinely. Physical models using ARM instrument data as inputs are implemented as VAPs and can help fill some of the unmet measurement needs of the program. Conversely, ARM produces some VAPs not in order to fill unmet measurement needs, but instead to improve the quality of existing measurements. In addition, when more than one measurement is available, ARM also produces "best estimate" VAPs. A special class of VAP called a Quality Measurement Experiment (QME) does not output geophysical parameters of scientific interest. Rather, a QME adds value to the input datastreams by providing for continuous assessment of the quality of the input data based on internal consistency checks, comparisons between independent similar measurements, or comparisons between measurement with modeled results, and so forth. For more information, see the VAPs and QMEs web page.

# Instrument Details

## Detailed Description

### List of Components

A pyranometer is the lowest component-level device used by an ARM instrument platform. Please refer to ARCS/GNDRAD, ARCS/SKYRAD, SIROS, or SIRS for a description of how pyranometers are used in these systems.

### System Configuration and Measurement Methods

The ARM Program uses The Eppley Laboratory, Inc. Model PSP pyranometer for shortwave hemispheric irradiance measurements. The PSP comprises a circular multi-junction wire-wound Eppley thermopile which can withstand severe mechanical vibration and shock. Its receiver is coated with Parson's black lacquer providing non-wavelength selective absorption. The instrument is supplied with a pair of removable precision ground and polished hemispheres of Schott optical glass. Both hemispheres are made of clear WG295 glass which is uniformly transparent to energy between 0.285 to 2.8?m. The instrument has a cast bronze body with a white enameled guard disk (shield), a spirit level for accurately determining the horizontal orientation, adjustable leveling screws and a desiccator which can be readily inspected.

This thermoelectric device (see "Theory of Operations" for details) produces electrical current proportional to the broadband shortwave irradiance reaching the detector. A shielded, twisted pair, copper wire signal cable connects the PSP to a data logger. The data logger is generally programmable and provides 1-minute statistics of analog to digital conversions made every few seconds (see other instrument platforms for details, e.g., SIRS).

Pyranometers are located to provide maximum solar access (i.e., the local horizon is unobstructed to the extent possible). The unshaded pyranometer is mounted in a ventilator which is in turn secured to a stable horizontal surface about 1.5 meters above the ground. The pyranometer/ventilator system is carefully leveled to provide accurate down welling shortwave irradiance. The ventilator provides a continuous stream of air over the outer dome of the pyranometer to prevent dew, frost, or dust accumulation that would effect the radiation reaching the detector.

A solar tracker is used to position a shading disk or ball above a pyranometer for down welling diffuse irradiance data. The shading device blocks the direct (beam) radiation from reaching the pyranometer detector. Continuously accurate tracker alignment with the sun's position is required for reliable measurements of diffuse (sky) radiation. In this configuration, the pyranometer is also mounted in a ventilator for reduced maintenance.

The inverted pyranometer is mounted several meters above the ground, typically on a meteorological tower mast, over a surface representative of local ground cover. An inverted pyranometer is generally not mounted in a ventilator as this orientation naturally protects the instrument's optical surfaces from contamination.

### Specifications

This section is not applicable to this instrument.

## Theory of Operation

This section is not applicable to this instrument.

## Calibration

### Theory

SKYRAD Shaded PSP

On a clear day at Momote at noon it should measure about 100 watts per square meter. The diffuse component will increase when clouds are present.

SKYRAD Unshaded PSP

On a clear day at Momote at noon it should measure about 1100 watts per square meter. The value will be reduced when clouds are present.

GNDRAD PSP

On a clear day at Momote at noon it should measure about 200 watts per square meter reflected from the grass below.

### Procedures

This section is not applicable to this instrument.

### History

This section is not applicable to this instrument.

## Operation and Maintenance

### User Manual

This section is not applicable to this instrument.

### Routine and Corrective Maintenance Documentation

This section is not applicable to this instrument.

### Software Documentation

ARM netCDF file header descriptions may be found at SIRS Data Object Design Changes.

### Additional Documentation

This section is not applicable to this instrument.

## Glossary

Absolute Cavity Radiometer
Broadband
Diffuse Irradiance
Direct Normal Irradiance
Extraterrestrial Irradiance
Global Irradiance
Irradiance
Radiance
Shortwave Irradiance
Solar Constant
World Radiometric Reference

Also see the ARM Glossary.

## Acronyms

AHF:  Automatic Hickey-Frieden.  The model name for an absolute cavity radiometer manufactured by The Eppley Laboratory, Inc.

ARCS:  Atmospheric Radiation and Cloud Station

BORCAL:  Broadband Outdoor Radiometer Calibration.  The summation method of pyranometer calibration as performed at NREL and the SGP/RCF.

GNDRAD:  Ground Radiation Instrument Systems.  Radiometers positioned to measure irradiance reflected by and emitted by the natural surface.

NREL:  National Renewable Energy Laboratory.   One of the eight national laboratories operated for the US Department of Energy.

PSP:  Precision Spectral Pyranometer.  The model name for the instrument manufactured by The Eppley Laboratory, Inc.

RCC:  Radiometer Calibration and Characterization.  The hardware and software implementation developed by NREL for automating a BORCAL.

RCF:  Radiometer Calibration Facility.  Located at the SGP Central Facility, the RCF is designed to perform calibrations of all broadband longwave and shortwave radiometers used by the ARM Program. Calibrations are traceable to international standards through the Metrology Laboratory and the Solar Radiation Research Laboratory at NREL.

SGP:  Southern Great Plains.  One of three climate regions under study by the ARM Program.  SGP encompasses southern Kansas and northern Oklahoma.

SIROS:  Solar and Infrared Radiation Observing System.  Predecessor to the SIRS radiometer platform for measuring broadband shortwave and longwave surface irradiance budgets.

SIRS:  Solar Infrared Radiation Station.  Current radiometer platform used in the SGP for measuring broadband shortwave and longwave surface irradiance budgets.

SKYRAD:  Sky Radiation Instrument Systems.  Radiometers positioned to measure shortwave (solar) and longwave (atmospheric or infrared) from the sky dome.

WG295:  A type of optical glass with uniformly high transmittance between wavelengths of 0.285 and 2.8 microns.

WRR:  World Radiometric Reference.  The international calibration reference for broadband shortwave irradiance radiometers maintained by the World Radiation Center in Davos Switzerland.

Also see the ARM Acronyms and Abbreviations.

## Citable References

Coulson, K.L.  1975.  Solar and Terrestrial Radiation Methods and Measurements, Academic Press, New York, New York.  ISBN 0-12-192950-7, pp. 322.

Dutton, E.G., J.J. Michalsky, T. Stoffel, B.W. Forgan, J. Hickey, D.W. Nelson, T.L. Alberta, and I. Reda. Measurement of Broadband Diffuse Solar Irradiance Using Current Commercial Instrumentation with a Correction for Thermal Offset Errors.  J. Atmospheric & Oceanic Tech.  (Accepted for publication July 2000).

Forgan, B.W.  1996.  A New Method for Calibrating Reference and Field Pyranometers.  Journal of Atmospheric and Oceanic Technology, Vol. 13, p.638-645.

Hulstrom, R.L.  1989.  Solar Resources, The MIT Press, Cambridge, Massachusetts.  ISBN 0-262-08184-9, pp. 408.

Iqbal, M.  1983.  An Introduction to Solar Radiation, Academic Press, New York, New York.  ISBN 0-12-373750-8 (ISBN 0-12-373752-4 for paperback), pp. 390.

Marion, ****ASES paper on MPA and cleaning study****

Maxwell, E.L, and D.R. Myers, "Daily Estimates of Aerosol Optical Depth for Solar Radiation Models" Solar '92, Proceedings of 1992 Annual Conference of the American Solar Energy Society, Cocoa Beach, Florida, June 15- 18, 1992, p.323.

McCluney, W.R.  1994.  Introduction to Radiometry and Photometry, Artech House, Inc., 685 Canton Street, Norweed, Massachusetts.  ISBN 0-89006-678-7, pp. 402.

Reda, I.   1999.  Improving the Shade/Unshade Method to Calculate the Responsivities of Solar Pyranometers.  NREL/TR-26483.

World Meteorological Organization.  1971.  Guide to Meteorological Instrument and Observing Practices.   4th edition, WMO-No. 8, TP. 3, Geneva, Switzerland.  (Note: Most WMO publications are available in the U.S. from the American Meteorological Society, 45 Beacon St., Boston, Massachusetts 02108-3693, Telephone:  617-227-2425.)