July 22, 1994

Contract # TW0592

SPECIAL SPECIFICATION

SECTION 15492

LABORATORY FUME HOODS

Page

PART 1 - GENERAL

1.01 Description.	2

1.02 Related Work.	2

1.03 Quality Assurance.	2

1.04 References..	2

1.05 Submittals.	3

1.06 Delivery, Storage, and Handling.	3

1.07 Project/Site Conditions.	4

PART 2 - PRODUCTS

2.01 Acceptable Manufacturers.	5

2.02 Equipment	5

PART 3 - EXECUTION

3.01 Inspection..	8

3.02 Installation/Application/Erection.	8

3.03 Field Quality Control.	8

3.04 Adjusting and Cleaning.	9

3.05 Commissioning..	9

3.06 Demonstration..	9

SPECIAL SPECIFICATION

SECTION 15492

LABORATORY FUME HOODS

PART 1 - GENERAL

1.01	DESCRIPTION

A.  	This Section covers requirements for providing bypass type fume hoods with pre-piped and pre-wired connections.

1.02	RELATED WORK

A.  	Section 01300

B.  	Section 15010

C. 	Section 15892

D.  	Section 15994

E.  	Section 15995

1.03	QUALITY ASSURANCE

A.  	Certify the fume hood performance according to American National Standards Institute (ANSI)/American Society of Heating, Refrigerating, and Air-Conditioning Engineers (ASHRAE) 110 criteria.

1.04	REFERENCES

A.  	ANSI/ASHRAE 110

B.  	National Fire Protection Association (NFPA) 70, National Electrical Code (NEC).

1.05	SUBMITTALS

A.  	Make submittals according to this Section, Section 15010, and Section 1300.

B. 	Product data for each fume hood:

1.  	Unit design requirements.

2.  	Unit features.

3.  	Weights (shipping, installed, and operating).

4. 	Furnished accessories.

5.  	Installation and start-up instructions.

C.  	Shop drawings detailing fabrication and installation of fume hood including plans, elevations, sections, component details, attachments, color finishes, and other construction elements.  Include the following:

1.  	Dimensions.

2.  	Weight loadings and distribution.

3.  	Clearances for maintenance and operation.

4.  	Size and location of field connections.

D.  	Wiring diagrams detailing power and control wiring.  Differentiate clearly between manufacturer-installed wiring and field-installed wiring.

E.  	Maintenance data for each fume hood.

F.  	Certification of performance from the factory.

G.  	Certification letter from manufacturer stating that the fume hood has been through start-up procedures and that it is functioning properly.

1.06	DELIVERY, STORAGE, AND HANDLING

A.  	See Section 15010.

B.  	Protect fume hoods from damage by factory packing.

C. 	Reject any damaged fume hoods upon arrival.

D. 	Store fume hoods to prevent damage, and protect from weather, dirt, fumes, water, and construction debris.  Provide a clean dry space if available.

E.  	Handle fume hoods according to the manufacturer's rigging and installation instructions for unloading and transporting into the final location.

1.07	PROJECT/SITE CONDITIONS

A.  	Coordinate layout and installation of the fume hoods with the piping system and exhaust ductwork.  Revise locations to suit field conditions and as approved by the Architect/Engineer (A/E).

B.  	Coordinate installation with laboratory furniture installation.

PART 2 - PRODUCTS

2.01	ACCEPTABLE MANUFACTURERS

A.  	Hamilton Scientific, Inc. or approved substitute

2.02	EQUIPMENT

A.  	General

1.  	Provide bypass type fume hood with a cabinet or table and structure as indicated in Tables 1 and 2.

2.  	Provide features as indicated in Tables 1 and 2.

B.  	Physical Characteristics

1.  	Provide double end walls for the superstructure with the front of the panel at the hood opening chamfered to provide a streamlined hood air entrance.  Provide hood interior end panels flush with the entrance shape.  Enclose the area between the double wall ends to house the required sash counter-balance weight, duplex receptacles, wiring, and remote service valves.

2.  	Provide an automatic air bypass sash opening which is dependent on mechanical and electrical linkage.  Design the air bypass to pass 20 percent or more of the flow rate when the hood sash is closed.

3.  	Provide a removable baffle with a fixed open center slot and adjustable openings at top and bottom.  Locate the baffle at the rear of the hood for adjustment of air flow due to types of gases.

4.  	Provide minimum 16-gage (ga.) metal; use 304 or 316 stainless steel where stainless steel is identified.

5.  	Provide pre-piped and labeled piping connections, so all connections are made exterior to the fume hood.  Provide pre-wired power connections with a pre-wired, 3-prong, 8-ft. long, power cord.

6.  	Construct the fume hood exterior of cold rolled steel component parts screwed together to allow for removal of the end panels, front and facia pieces, top facia and air foil strips for maintenance.  Provide eight foot wide fume hoods that have a vertical sash movement with either a double sash and a center post or a single sash and a reinforced opening frame.  Provide all component parts with an acid, alkali, and solvent resistant finish on both exterior and interior surfaces after fabrication of all cold-rolled steel parts and prior to final assembly.  Provide approved finish color.

7.  	Provide at the top of the hood a fluorescent or incandescent light fixture as indicated in Table 2. Provide a fixture shielded from the hood interior by a tempered glass panel sealed to the hood by an extruded vinyl channel.  Provide electronic ballasts and T-8 lamps for fluorescent fixtures.

8.  	Provide a sash for the hood.  Provide 7/32 to 9/32-inch (in.) laminated safety glass or 1/4-in. transparent lexan for the sash as indicated in Table 1.  Construct the sash of an 18 gauge painted steel rolled shape, 3/4-in. thick by  2 1/2-in. wide, which is mitered, welded, and ground smooth at the corners to provide a complete frame with no visible joints.  Use internal interlocked, rigid, vinyl extrusions with the outer member, sealing and retaining the glass.

a. 		Counterbalance up/down moving sashes with a weight and sash cable system to prevent any tilting of the sash during operation.  Do not use spring type counterbalances.  Use stainless steel sash cables operating on ball bearing sheaves.  Provide replaceable plastic guides for the sash frames, which operate in stainless steel sash guides to prevent metal-to-metal contact.

b		Provide ball-bearing rollers and guides for the left/right moving sashes.  Use stainless steel for the roller construction.  Use replaceable plastic inserts in the stainless steel guides to prevent metal-to-metal contact.

9.  	Construct the fume hoods to permit passage through a 3-foot by 7-foot personnel door.

10. 	Design the fume hood for exhaust duct connection at the top.

11. 	Interior Surface Construction

a. 		Provide fume hood interior materials as indicated in Table 1.  Weld stainless steel materials together.  Triple weld polypropylene materials together.  Provide gaskets to join dissimilar materials.

b. 		Provide a continuous curb for the working surface.  Provide reinforced working surface capable of supporting 100 pounds per square foot without damage or warping.

c. 		For areas exposed to the fume hood interior, use stainless steel fasteners on removable parts.

d. 		Provide 14-ga. reinforced stainless steel for the work surface as indicated in Table 1.

12. 	Features

a. 	Provide features as indicated in Table 2.

b. 		Provide oval cup sinks with a size of approximately 3 in. by 6 in., unless notes otherwise.  See Table 2.

c. 		Provide a neoprene gasket between the working surface and the sinks.

13. 	Service Locations

a. 		See Table 1.

b. 		Locate the duplexes, remote baffle operator, remote valve operators, and light switches on the end walls (exterior to the fume hood sash).  Replace duplexes with a power strip having equal number of outlets and same protection, if so desired.

c. 		Provide 3/8-in. service valves with remote serrated hose connection located in the fume hood interior side panels.

14. 	Use only colors approved by SNL.

PART 3 - EXECUTION

3.01	INSPECTION

A.  	Inspect areas and conditions, with Installer present, for compliance with requirements for installation tolerances, and other conditions affecting performance of the fume hoods.

B.  	Inspect rough-in for utilities to verify actual locations of connections prior to installation.

C.  	Do not proceed until unsatisfactory conditions have been corrected.

3.02	INSTALLATION/APPLICATION/ERECTION

A.  	General

1.  	Install fume hoods level and plumb, in accordance with manufacturer's written instructions.

2.  	Arrange installation of units to provide access space for service and maintenance.

B. 	Connections

1.  	Piping installation requirements are specified in other Division 15 sections.  The Drawings indicate the general arrangement of piping, valves, fittings, and specialties.

2.  	Duct installations and connections are specified in other Division 15 sections.

3.  	Electrical Connections:  The following requirements apply:

a. 		Electrical power wiring is specified in Division 16.

b. 		Connect unit components to ground in accordance with the National Electrical Code.

3.03	FIELD QUALITY CONTROL

A.  	Arrange and pay for a factory-authorized service representative to perform the following:

1.  	Inspect the field assembly of components and installation of the fume hoods including piping, ductwork, and electrical connections.

2.  	Prepare a written report on findings and recommended corrective actions.

3.04	ADJUSTING AND CLEANING

A.  	Adjust sashes for proper operation.

B.  	Adjust baffle linkages for proper baffle operation.

C.  	Clean unit fume hood interiors to remove foreign material and construction dirt and dust.

3.05	COMMISSIONING

A.  	Provide labor, material, equipment, etc., required to facilitate the commissioning process.  Perform form tests and verification procedures required by the commissioning process when requested by the Commissioning Authority and directed by the General Contractor.  See Section 15995.

3.06	DEMONSTRATION

A.  	Provide a factory-authorized service representative supervise start-up of equipment, to demonstrate operation, and to train personnel.

B.  	See Section 15995.

END OF SECTION

EXCEPT FOR ATTACHMENTS

FUME HOOD SPECIFICATION TABLES 1 & 2

FUME

HOOD

SPECIFICATION

TABLE

1

Dimensions                                  Construction             Envelope              Basis of Design

Equip.Room                                                             Material                 Seal

No.

FH-

ER =

epoxy

(chemical

resistant)

resin,

G = glass,

L = lexan,

PP =

polypropylene,

PVC = polyvinyl

chloride,

S = slate,

SS = stainless

steel

Hood Maximum  Working Working Opening  InteriorSashWorkingEnvelope/ Frame/Drain     W =

Width  Height  Surface Surface  Height  Working     Surface  Plenum Chassis        Welded,

Feet  Inches   Height  Depth   Inches   Height                                  C = Calked,

Inches Inches            Inches                                  G = Gasketed

1    1405   4      90       36      30      30       44    L     PP      PP      SS   PVC      W/G

2    1405   4      90       36      30      30       44    L     PP      PP      SS   PVC      W/G

3    1206   6      90       36      30      30       44    G     ER      SS      SS   PVC       W

4    1206   4      90       36      30      30       44    G     ER      SS      SS   PVC       W

5    1104   4      90       36      30      30       44    G     ER      SS      SS   PVC      W/G

6    1104   4      90       36      30      30       44    G     ER      SS      SS   PVC      W/G

7    1104   4      90       36      30      30       44    G     ER      SS      SS    -       W/G

8    1104   4      90       36      30      30       44    G     ER      SS      SS    -       W/G

9    1109   6      90       36      30      30       44    G     ER      SS      SS   PVC      W/G

10   1109   6      90       36      30      30       44    G     ER      SS      SS   PVC      W/G

12   1706   5      90       36      30      30       44    G     SS      SS      SS    -        W

15   2213   4      90       36      30      30       44    L     ER      PP      SS    -       W/G

FUME

HOOD

SPECIFICATION

TABLE 2

Features               Service

Equip.    Room                                                     Locations

No.

FH-

Sash Slides    Base Cabinet             Lighting  Ar   FP = foot pedal operator,1 on

=    LS = lightthe

(125    argon,switch,left,

footcandles DI = deionizedN22L =

at the   water,=  2 on

work    DPLX nitrogen,the

surface)  =    NG  left,

120  =   etc.,

V,   naturalR =

GFCI,gas,1 on

Duplex,NPW = non-the

NDPLXpotableright,

=    water,2R =

120  PCW = process2 on

V,   coolingthe

Non- water, right,

GFCI,L   etc.

Duplex,=

U/D  Y   D =  DoorsF =FluorescentSinkDrainN2  Ar

=    =   Dr = Drawers

Up   Yes Ds = DoorsI =IncandescentY =

&              Yes

or   N        Shelves

=   S =  Shelves        N =

Down No                       No

L/R

=

Left

or

Right

NG       DI      NPW       NPW         PCW            DPLX           LS       1     1405 U/D   Y     D      F      R      Y        L

FP

-        -        R         -           -              -             2L       L      2   1405 U/D    Y      D      F      R        Y

L        -        -         R           -              -             -        2L     L    3   1206  U/D     Y      D      F        R

Y        L        -         -           -              R             R        -      2L   L    4    1206   U/D     Y      D        F

R        Y        L         -           -              -             R        R      -    2L   L     5     1104   U/D     Y        D

F      Note 1     Y         L           -              R             R        R      -    -    L     L      6     1104   U/D       Y

D        F     R,Note 2     Y           L              -             R        R      R    -    -     L      L      7     1104     U/D

Y        D        F         -           N              L             -        R      R    R    -     -      L      L      8       1104

U/D       Y        D         F           -              N             L        -      R    R    R     -      -      L      L        9

1109     U/D       Y         D           I              R             Y        L      -    R    R     R      -      -      2L       L

10      1109     U/D        Y           D              I             R        Y      L    -    R     R      R      -      -        2L

L        12      1706      L/R          Y              D             F        -      N    -    -     -      -      -      -        -

6L       L        15       2213        U/D             Y             Ds       F      -    N    L     -      -      R      -        -

-        3L       L

Notes:

1.      Provide a sink for the entire work surface.

2.      Provide a 12-in. by 12-in. sink.