Information on C4F10 systems………..

Forward RICH of DELPHI

- Evolution of the transparency of C4F10 in the Forward RICH of DELPHI  (5.22.00)

- Original gas was cleaned by molecular sieves to get rid of water

  - Transmittance spectra fell for wavelengths lower than 1950 A

  - Used 13X molecular sieve

- Massive use of activated carbon in addition to molecular sieves raised transmittance to good levels

  - Took 1.5 kg of carbon to clean 8 liquid liters of C4F10

- Fabry-Perot interferometer used to measure refractive index

- Operated at ambient temperature

- Sound velocity measurement allowed determination of C4F10 amount in mixture

  - Claimed to be cheap and reliable

- Used massive amounts of molecular sieves and activated carbon

COMPASS RICH

- The COMPASS RICH project (10.23.99)

- Gas RICH with 3 m long C4F10 radiator, atm. Pressure, 25 oC

- Goal was water vapor and oxygen traces below 5 ppm

- H2O and O2 removed using molecular sieves and C4F10 and N2 separation is obtained with a liquifier

- Gas flow was 1 volume every 8 hrs.

- Temperature stability at +/- 1oC.

- COMPASS RICH-1 (3.13.01)

- Filling and emptying show gas replacement at 97% level.

LHCb RICH

- LHCb RICH Gas System Proposal (8.6.00)

- RICH 1 filled with 4 m3 of C4F10.

- Gas systems will run as closed loop circulation with a gas recovery system.

  - Inline liquefaction of return gas to separate volatile impurities

  - C4F10 enters the separation system with N2/O2 and will be pre-cooled while passing the return gas thru 2 heat exchangers.

  - C4F10 cooled down to –50 oC

  - Pure C4F10 is then collected and transferred to a reservoir which serves as a supply for the RICH

  - N2/O2 leave thru exhaust

  - Pump in return line compresses gas before entering the separation plant

  - C4F10 will be liquid behind the compressor

- Oxygen and water not too exceed 100-200 ppm

  - Molecular sieve used to reduce water impurities along with activated carbon

  - 3X and 4X sieves also used

  - Gas analysis before and after purification

- Residual nitrogen has an effect on refractive index and should be kept constant and below 1%

- The gases are circulated in a closed loop with an expected flow rate of 10% of the total gas volume.

- Two gas inlets/two gas outlets

  - C4F10 from bottom and N2 from top

- Monitoring, Alignment and Control of the RICH Detectors (9.6.00)

- Refractive index depends upon temperature and pressure

- Gases are clean from trace impurities after 1 week of recirculation

- Ultrasonic sonar measures variations in the refractive index due to increased concentrations of nitrogen in the system

- Monochromator measures transparency

- http://dustbunny.physics.indiana.edu/~dzierba/pidtemp/

- Endcup CRID group at SLD

- Can come with oil contamination

- Must recirculate and have a reasonably tight vessel

- Use simple distillation

- HERA-B RICH

- Filled with C4F10 and operates at room temperature

- HERMES @ DESY

- C4F10 RICH for one radiator and silica aerogel tiles for the other

- Gas control system recirculates the radiator gas through the main volume at a slight over pressure.