Review of

ITER Toroidal Interferometer/Polarimeter Design

October 21, 2005

C.W. Domier, K.C. Lee, N.C. Luhmann, Jr.

UNIVERSITY OF CALIFORNIA, DAVIS

Department of Applied Science

PRELIMINARY REPORT ON THE ITER MULTICHANNEL TOROIDAL INTERFEROMETER/POLARIMETER DESIGN

C.W. Domier, K.C. Lee, N.C. Luhmann, Jr.

University of California at Davis
Dept. of Applied Science
Davis, CA 95616

# ABSTRACT

Designs for a multi-channel toroidal interferometer/polarimeter (TIP) system to measure the electron density profiles on ITER are reviewed. It is concluded that a two wavelength system design for vibration compensation is still required; however, the choice of wavelengths is not yet fixed. Newly reported studies into environmental effects on retroreflectors and in-vessel mirrors reveal serious problems with the short wavelength 10.6 µm/5.3 µm choice for the TIP system. These problems are alleviated somewhat by going to a slightly longer wavelength choice of 10.6 µm/9.27 µm or 12.1 µm/9.0 µm. Consideration of a longer 57.2 µm/47.6 µm wavelength option is suggested for consideration, with the pros and cons of each approach identified. While the longer wavelength option has some partially resolved technology issues which need to be addressed and requires increased diameter retroreflectors and input/exit mirror, success of the shorter wavelength option may require a major redesign of the input/exit mirror arrangement.

The need for a real-time alignment system is reaffirmed, with the preferred choice being a self-scanning IR/FIR-based system over that of a visible light system due to potential mirror reflectivity problems associated with mirror erosion and carbon deposition. The dispersion interferometer concept, which employs nonlinear optics to frequency double an input laser beam which then co-propagates with the original un-doubled laser beam, is examined for applicability to ITER. Finally, a number of potential research and development areas are identified that could greatly impact the TIP system design, as well as potential physics issues that could be addressed using the TIP system. One key area here is the possible application of hollow waveguides for transporting signals to/from the ITER plasma.

# A. INTRODUCTION

The development of a multi-channel toroidal interferometer/polarimeter (TIP) system is desired to measure the electron density profile on ITER with high reliability and sufficient time resolution (1 msec) and accuracy (1%) for plasma feedback control. Supplementary data are to be provided by a poloidal interferometer/polarimeter system.

A short wavelength probing laser beam is used in TIP both as an interferometer to measure line average electron density and as a polarimeter to measure the plasma-induced Faraday rota-tion. A plan view of ITER showing five tangential probe beam paths through the ITER plasma is shown in Fig. 1, as taken from Sec. 2.6.3.3 of the ITER Technical Basis – Plant Description Document [1]. The beams enter and exit the plasma from a single port and are reflected back along the same beam path by plasma facing corner cube retroreflectors (CCRs) embedded within the ITER blanket shield modules (BSMs). By requiring only a single port with no large vibration damping structures, the total penetration area through the BSMs is minimized and the neutron leakage is thus reduced. Reflective optics are employed throughout to avoid neutron damage issues associated with transmission optics, with the exception of vacuum windows which are located well behind the port plug in a relatively benign radiation environment.

Fig. 1. Schematic layout of the ITER toroidal interferometer/polarimeter system showing the five lines of sight through an equatorial port (from [1]).

This diagnostic will provide the primary line integral density measurement input to the ITER basic control algorithms. The system must therefore be designed for near ~100% availability, as any signal drop-out will most likely result in a controlled pulse termination. Extreme care must therefore be taken in the multichannel TIP design. The purpose of this report is to evaluate the current ITER designs, and to identify critical design issues that are still unresolved.

# B. OVERVIEW OF EXISTING ITER TIP DESIGNS

An initial TIP design was developed for the ITER-98 device by Carlstrom et al. in the late 1990s [2, 3]. This design was based on the use of collinear 10.6 µm (CO2) and 5.3 µm (CO) laser probe beams, with the shorter wavelength beam employed for vibration compensation. A pair of acousto-optic modulators are employed to generate intermediate frequency (IF) signals for interferometry and polarimetry. In the example listed in Ref. 3, right and left circularly polarized beams are generated with 42 MHz and 38 MHz frequency shifts, respectively, from the unshifted laser beam. The two beams are combined to form an elliptically polarized beam which rotates at half the difference frequency (4 MHz) between the two beams. The rotating elliptical beam is passed through the plasma and interferes with a reference (unshifted) laser beam at a single detector, from which both interferometry and polarimetry signals are derived. This general approach (ignoring wavelength) follows the procedure employed earlier on TEXTOR [4], MTX [5] and RTP [6, 7], and more recently on a number of other plasma devices including NSTX [8].

The Carlstrom et al. design was revisited and revised by Kondoh et al. [9] for the revised ITER device. It shares many of the same design aspects, but with some key differences. Firstly, the extremely short 5.3 µm wavelength of the vibration compensation laser was rejected due to problems experienced on JT-60 with a two color interferometer system employing 10.6 µm (CO2) and 3.39 µm (HeNe) lasers [10]. This JT-60 interferometer implementation experienced significant stability and reliability problems arising from mirror damage, window coating, and mechanical vibrations at the shorter wavelength. Significantly improved operation was achieved on JT-60 after the 3.39 µm laser was replaced with a longer wavelength 9.27 µm (CO2) laser [10-12], as they now propose for use on ITER. Alternatively, they suggest the use of a 12.1 µm (14C16O2) / 9.0 µm (12C18O2) arrangement should the 10.6 µm/9.27 µm combination have insufficient wavelength separation for good vibration compensation.

CO2 beam position sensors are proposed in the Carlstrom et al. design to control alignment mirrors and to keep the beams focused on the detectors, whereas a visible light laser and a camera are employed in the Kondoh design. This visible light alignment system was developed for the ITER diverter impurity monitor, and would monitor light scattered from a 3 mm × 3 mm diffuser mounted at the side of each corner cube retroreflector (CCR). Another area in which the two designs differ is the size and operation of the CCRs. In the Carlstrom et al. design, the return beam reflected from the CCRs (35 mm × 70 mm clear aperture) is horizontally offset by 35 mm to prevent the return beam from re-entering the laser and thereby adversely affecting its stability. The Kondoh design minimizes the retroreflector design requirement (essentially doubled in clear aperture in the Carlstrom et al. design) by eliminating the offset in the reflected beam. Here, the Kondoh retroreflectors have a clear aperture of 30 mm × 30 mm.

Vacuum windows in the Carlstrom et al. design were to be made of ZnSe. Estimates of the Faraday rotation expected in these windows at 10.6 µm was considered acceptably low (0.25º) assuming a radial field of 0.3 T at the window. The Kondoh team has proposed replacing these windows with chemical vapor deposition (CVD) diamond windows since the Verdét constant of the diamond is 1/17 that of ZnSe, thus significantly reducing the Faraday rotation. Additional benefits of CVD diamond include increased hardness, reduced thickness, excellent heat transfer and ease of vacuum sealing, all of which trade off against the higher cost of CVD diamond.

# C. CRITICAL ISSUES

## 1) WAVELENGTH

The problems of window damage and coating experienced on JT-60 at extremely short wavelengths raise serious concerns. An equally serious concern has to do with the cumulative effects of erosion on the retroreflectors. A study by Dr. Voitsenya of the Kharkov Institute for Physics and Technology, in which CCRs were bombarded with deuterium ions from a plasma source [13], has shown that long term sputtering by charge exchange atoms (CXA) and/or deposition of carbon-based (or beryllium-based) contaminant layers can affect the optical properties of the CCR.

Qualitatively, Bennett’s formula on surface roughness, R=R0exp[–(4πd)2/λ2], is frequently used to express the specular reflectance R of a rough surface whose ideal reflectivity is given by R0 where d is the mean surface roughness and λ is the wavelength of the reflected light [14]. This is plotted in Fig. 2 for both a single mirror surface and a CCR as a function of d/λ. As is pointed out by Voitsenya, the specular reflectance of a CCR at ~10 µm has decreased by 20% for a surface roughness d ≈ 200 nm.

Fig. 2. The specular reflectance of a metal mirror as a function of the mean roughness according to the Bennett formula (from [13]).

Although CXA sputtering was shown to seriously degrade the CCR reflectivity at 10 µm, mounting the CCRs in narrow and long grooves should eliminate CXA sputtering concerns as long as the CXA flux can be reduced by at least an order of magnitude as compared to that at the first wall. The primary concern instead is the role of carbon film deposition. The amount of carbon film buildup expected in ITER is at present unknown, but studies have been performed on TEXTOR [15] using a “periscope” with a channel length three times that of the diameter (35 mm) placed in the scrapeoff layer and oriented along a tangent to the toroidal direction. Extrapolations from this study (see Fig. 3) would predict a carbon film layer of 250 nm (such as was built up on TEXTOR after a total of 1047 sec of plasma discharge) to be reached within one to two ITER discharges. It should be noted, however, that the viewing lines on the TEXTOR study were not fully tangential and that the deposition on ITER may be substantially worse.

Fig. 3. The thickness of carbon-based film deposited on a test mirror in the TEXTOR periscope system (from [15]).

Increased surface roughness, as a result of carbon film deposition, has the effect not only of decreasing the mirror reflectivity but can also alter the polarization angle of the reflected beam. The polarization effect depends strongly on the initial polarization of the probing beam and is maximum when the electric field is parallel to the mirror surface. Plotted in Fig. 4 are the calculated changes in polarization angle after a single reflection from a carbon coated rhodium mirror (rhodium film electroplated onto a metal substrate) at 10 µm. The figure shows that the polarization rotation effect (due to surface roughness) depends on both film thickness and angle of incidence, and that the desired ITER polarization accuracy of 0.01º is achieved with a relatively thin carbon film thickness of 100 nm only with incidence angles < 40º. No data are available on the polarization effect for triple reflections as would be the case from a retroreflector; further calculations are therefore required. Of perhaps greater concern to the short wavelength operation design is the polarization effect of carbon film buildup on the input/exit mirror. The angle of incidence upon this mirror will have to be chosen to be as small as possible so as to have the smallest polarization changes, which implies an increase in the number of turning mirrors required to provide the five probing beams. This in turn implies significant changes to the reflective optics which carry the probing signals to and from the input/exit mirror.

Fig. 4. Polarization angle rotation at λ=10 µm calculated for a single reflection from a rhodium mirror as a function of carbon (C) film thickness and incidence angle (from [13]).

ITER also has a high level of potential synchrotron emission in the 3-5 µm range due to suprathermals. This, together with the earlier expressed erosion and deposition concerns regarding the input/exit mirror and CCRs, would make it prudent to discard the 10.6 µm/5.3 µm option in order to maintain the robustness of the diagnostic. While the 10.6 µm/9.27 µm and 12.1 µm/9.0 µm options could prove viable (the use of heated mirrors may result in reduced deposition, while research into mirror mitigating/coping methods via in-situ laser cleaning and/or heating of the vacuum vessel is ongoing), the former may have insufficient wavelength separation for good vibration compensation. On the other hand, the 10.6 µm/9.27 µm CO2 laser technology is well proven and highly dependable while the 12.1 µm (14C16O2) / 9.0 µm (12C18O2) laser technology is far less advanced and may not be sufficiently reliable for use on ITER.

As an alternative, it is worth considering longer wavelength operation since both mirror reflectivity and polarization effects are significantly ameliorated. Additional benefits include an increased interferometer phase shift (scales linearly with wavelength, an increased Faraday rotation (scales with the wavelength squared), and a decreased phase shift generated by mechanical vibrations (scales inversely proportional to wavelength). The main disadvantages are increased diffraction and refraction effects, and a significantly increased Cotton-Mouton effect (scales with the wavelength cubed).

Diffraction effects, which dictate how rapidly a collimated beam spreads out as it propagates, result in larger diameter retroreflectors and/or steering mirrors being required at longer wavelengths. Plotted in Fig. 5 are Gaussian beam simulations of the five TIP beams, calculated at 10.6 µm. Beams that travel the longest distances to their corresponding CCRs are less tightly focused, and therefore are of the most concern.

Fig. 5. Schematic of the beam layout and designed beam diameter of a Gaussian CO2 laser from the laser output to the retroreflectors (from [9]).

A general rule of thumb is that the diameter of each optical element must be at least 2.2 times the beam diameter (the 1/e width of the intensity distribution of the laser beam), thereby allowing a 99.2% theoretical transmission through the optic. In the simulations of Fig. 5, the Ch. 1 beam diameter is approximately 11 mm at the CCR. This would lead to a minimum retroreflector diameter of 24 mm to avoid vignetting the reflected beam. Applying a 25% safety factor, this yields a CCR diameter of 30 mm as chosen by Kondoh et al. [9].

CCRs with a 37 mm diameter (37 mm wide by 60 mm tall) are envisaged for use in the ITER poloidal polarimeter system, as described by Donné et al. [16]. A CCR of diameter 37 mm employed in the TIP system would allow a wavelength increase to (37/30)×10.6 µm = 13.1 µm using the same diameter input/exit mirror and the same 25% safety factor. The 10.6 µm/9.27 µm and 12.1 µm/9.0 µm options are thus certainly feasible with CCRs of this size.

The next step up in wavelength with sufficient optical power would be a 57.2 µm/47.6 µm CH3OD laser system as is under investigation by Nakayama et al., with output powers of 1.6 W and 0.8 W, respectively [17]. Assuming a 13 mm beam waist diameter at the CCR corresponding to a CCR with a 37 mm diameter, this would translate to an input/exit mirror beam diameter of ~90 mm (for λ=57.2 µm) as compared to ~23 mm (for λ=10.6 µm) in the Kondoh design. A CCR of diameter 50 mm would couple to an 18 mm beam waist diameter and an input/exit mirror beam diameter of 66 mm which is somewhat smaller than the ~75 mm diameter mirror envisaged for the poloidal interferometer/polarimeter system.

Refraction effects will bend the probing beam, resulting in translation of the beam at the location of the retroreflector. Plotted in Fig. 6 are the simulated radial refraction and interferometer phase shift at 10.6, 12.1, 47.6 and 57.2 µm for ITER plasmas with two different density profiles given by ne = 1019+a×1020×(1–ρb) m–3 where ρ is the normalized radial coordinate. As was done by Donne et al. [16], both a flat (ITER reference) plasma with a=1, b=10 and a peaked plasma with a=2, b=2 have been studied. While refraction is certainly increased at these longer wavelengths, the worst case level of radial refraction expected at 57 µm is still small (~6 mm) as compared to the beam diameter and will be larger than the corresponding levels of vertical refraction arising from the shaped ITER plasmas. These small displacements can be easily corrected by (a) appropriate beam steering, or (b) extending the size of the retroreflector along the plasma midplane in a manner similar to that proposed by Donné for extending the retroreflector vertically in the poloidal system. One item of concern which has not been evaluated here is the refractive effects that ELMs might have at these longer wavelengths, which should be an item of future study.

Based on these calculations, neither diffraction nor refraction are, in themselves, limiting factors to 57.2 µm/47.6 µm operation. We now turn our attention to the Cotton-Mouton effect which arises from plasma birefringence [18-20]. For an elliptically polarized beam, the Cotton-Mouton effect causes the ellipticity of the polarization ellipse to change, while the Faraday rotation effect causes the ellipse to rotate with constant ellipticity. If the Cotton-Mouton effect is large, this can considerably complicate the optical layout needed to acquire the desired Faraday rotation measurements [20]. We note that one of these techniques, which involves the simultaneous phase measurements of both Faraday and Cotton-Mouton effects using polarization modulation, has been investigated numerically [21] and tested experimentally at 119 µm [22] for possible future use in the ITER poloidal interferometer/polarimeter system.

(a) 
(b)

Fig. 6. (a) Refraction effects, and (b) interferometer phase shift, as a function of wavelength and beam position for simulated ITER plasmas.

Plotted in Fig. 7 are simulated Faraday rotation and Cotton-Mouton effect predictions for the same conditions as assumed in Fig. 6. The Cotton-Mouton effect, which is extremely small at ~10 µm, yields a shift as large as 7º at 57.2 µm. While not insignificant, careful optical design can limit the amount of Faraday rotation measurement inaccuracy that this relatively low change in ellipticity produces. Consequently, there should be no need for a complicated optical layout and detection scheme to measure the Cotton-Mouton at this level, particularly when compared to the large Faraday rotation of ~1600º predicted on this channel for the same plasma parameters.

At 10.6 µm, the role of the Faraday rotation measurement is to improve the reliability of the system, particularly during long pulse operation, since the phase rotation is less than one fringe and does not require a fringe counter based system to keep track of multiple fringe shifts. In the case of a 57.2 µm/47.6 µm operation, the single fringe condition (corresponding to 180º Faraday rotation, as most phase detection schemes detect twice the Faraday rotation) is easily exceeded on all channels. Nonetheless, an unambiguous determination of the absolute Faraday rotation level is still possible without having to track multiple fringe shifts by using Faraday rotation data from both wavelengths. Faraday rotation scales as λ2, which at these wavelengths translates to the 57.2 µm signal undergoing 1.44 times the rotation experienced at 47.6 µm. Another way of putting this is that the 57.2 µm signal undergoes 13 fringes of rotation in the time that the 47.6 µm undergoes 9 fringes (13/9 = 1.44). Recording the relative rotation experienced at both 57.2 µm and 47.6 µm allows the absolute rotation to be determined within a 13 fringe (at 57.2 µm) window. This translates to a maximum allowable rotation angle (for unambiguous determination without tracking) of 2340º, which is less than the ITER simulations in Fig. 7 predict. Hence, robustness is still ensured. The relatively low amount of Faraday rotation expected at ~10 µm (see Fig. 7) led in turn to the requirement of a Faraday rotation resolution of 0.1º to keep the phase detection error less than 1% [3]. The Faraday rotation levels at ~50 µm are increased by a factor of >20 times; hence, a minimum Faraday rotation resolution of 2.0º would be appropriate here and relatively easily achieved.

(a)
(b)

Fig. 7. (a) Faraday rotation, and (b) Cotton-Mouton effect, as a function of wavelength and beam position for simulated ITER plasmas.

The next question that arises is whether or not there are any technology issues that could hamper successful operation on ITER at these wavelengths. The issues here are laser modulation, detection, beam splitters, and reliability.

Heterodyne detection techniques require that there be a difference frequency between the probing beam and the reference beam which interfere on the detector. Techniques to generate this difference frequency are: (1) acousto-optic (Bragg cell) modulation, (2) Doppler shift by reflecting radiation off a rotating cylindrical grating or “Veron” wheel, and (3) beating the output of two lasers at slightly offset frequencies.

Acousto-optic modulators (AOMs) are commercially available at wavelengths up to 12 µm. The diffraction efficiency depends strongly on optical wavelength (scales as λ–2), however, making it difficult to apply to the FIR regime with high efficiency. A study made in the 1980s identified a number of liquids with potential for long wavelength FIR acousto-optic modulators [23]; however, this work did not progess beyond a very preliminary study. Rotating gratings have been successfully applied at wavelengths of 119 µm and above, with frequency shifts as high as 100 kHz [22]. Although more difficult and more expensive to fabricate at ~50 µm, this is certainly a feasible option although one whose low bandwidth will limit the temporal response. Although a fast time response is not a requirement of the ITER TIP system for density profile determination, a high IF frequency permits the system to follow rapid density changes without loss of signals. In addition, recent experiments with the FIReTIP system on NSTX have shown the usefulness of a high bandwidth TIP system for the study of edge density fluctuations. One option here for high bandwidth is the so-called super rotating grating, in which a rotating grating is fabricated on the top of the rotor of a turbomolecular pump with a high rotation speed of 43,200 rpm [24]. This has been employed on CHS to achieve a 1 MHz frequency shift at 337 µm with a frequency stability Δf/f = 4×10–4 [25], althouth the robustness and reliability of the super rotating grating concept for use on ITER needs to be examined in more detail.

Optically-pumped FIR lasers have a natural linewidth of a few MHz. A three laser implementation, such as on the FIReTIP system on NSTX, is another option for simultaneous interferometry/polarimetry operation. The natural linewidth of the CH3OD laser lines is sufficient to obtain a 2-3 MHz frequency shift for polarimetry, while the Stark effect technique successfully employed on NSTX [8] can be used to provide a 5-10 MHz frequency shift for interferometry. It should be noted here that a Stark-effect linewidth broadening technique has not yet been attempted on a CH3OD laser, although the symmetry of the CH3OD molecule should ensure its success similar to that achieved on NSTX with a CH3OH laser at 119 µm [26].

Another area of concern for ~50 µm operation is identification of suitable materials for use as vacuum windows and beam splitters. Here, Chubu University and NIFS in Japan have been at the forefront of this research area [27]. CVD-diamond and silicon are both shown to be suitable window materials with high transmittance (>80%) at ~50 µm. Silicon etalons and wire meshes (similar to that employed in the FIReTIP system on NSTX [8]) have both been studied for use as beam splitters. Silicon etalons have shown promise as beam splitters; however, they require a tight tolerance on etalon thickness (Δt < 5%). The wire mesh beam splitters employing simple cross patterns appear to be relatively easy to fabricate, but it has proven difficult to obtain the same transmittance/reflectance ratio at 57.2 µm and at 47.6 µm.

Detector technology choices at ~50 µm are (a) Schottky barrier diode (SBD) corner cube mixers which have been successfully used at NIFS [28,29], (b) Ge:Ga photodetectors with a noise equivalent power (N.E.P.) < 8×10–13 W Hz–1/2 at 1 kHz which have also been successfully used at NIFS [28], and (c) InSb detectors with an N.E.P. < 7.5×10–13 W Hz–1/2 at 1 kHz [29]. Ge:Ga photodetectors have been experimented with at NIFS where they have been used with good results in conjunction with their 57.2 µm /47.6 µm lasers. Of these, only the SBD mixers do not require cryogenic cooling and are thus an attractive option for ITER.

Continued advances in SBD technology have taken place in the past decade, not yet reflected in the NIFS studies. Whisker-contacted Schottky diodes with 16 THz cutoff frequencies are commercially available from NMRC, Ireland [30]. Because whisker-contacted diodes are very sensitive to mechanical influences, and the contacting procedure is very difficult and time consuming, much of the research in this field involves the development of planar devices suitable for monolithic integration. Planar Schottky diodes have been developed at Virginia Diodes in the U.S.A. with cutoff frequencies of up to 7 THz [31], and by the Institut für Hochfrequenztechnik in Germany with cutoff frequencies up to 9 THz [32]. A receiver noise temperature of 8,250 K has been reported at 2.5 THz using GaAs Schottky monolithic membrane diodes developed at the Jet Propulsion Laboratory (JPL), with an estimated mixer noise temperature of 3,500 K [33].

To summarize, newly reported studies into environmental effects on retroreflectors and in-vessel mirrors reveal serious problems with short wavelength operation. Of the short wavelength options previously considered for the ITER TIP system, the 10.6 µm/5.3 µm option would suffer the greatest from these problems and it is recommended to discard it. The wavelength choices are thus between a 10.6 µm/9.27 µm or a 12.1 µm/9.0 µm system as proposed by Kondoh et al., and a 57.2 µm/47.6 µm system being pursued by NIFS. The former has unresolved issues regarding mirror and retroreflector damage and coating that do not yet have a solution. The latter requires a significant increase in the size of the retroreflectors and input/exit mirror, and some technological issues require further development; however, it appears to be a viable alternative.

## 2) BEAM PROPAGATION TO/FROM THE ITER PLASMA

One technology not discussed in either of the ITER TIP design papers is the use of hollow glass waveguides (HGWs). This type of waveguide typically consists of a reflective metallic layer and a reflection-enhancing dielectric layer deposited on the inside of a small bore silica or glass tubing. These waveguides have been developed to transport high power CO2 (10.6 µm) radiation with low loss. Early papers in this field employed constant-bore guides, with reported losses of 0.17 dB/m for 1000-µm-bore guides and 0.46 dB/m for 530-µm-bore guides propagating the fundamental HE11 mode [34]. The loss increases to 0.93 and 1.36 dB/m, respectively, when guides are bent to a radius of 0.25 m.

The use of tapered-bore hollow glass waveguides is now under investigation. These can be divided into two types: short and long tapers. Short tapers are tapered at a steep angle 1° over a short distance, typically <10 cm, and are often used in conjunction with nontapered HGWs at the input end to reduce coupling losses. Long, gradually tapered HGWs are tapered at a very small angle (typically 0.1°) and are intended to be used in place of traditional, nontapered guides. The large bore at the input end improves coupling, and the small bore at the output end allows for the enhanced flexibility. Reported losses are higher than in nontapered HGWs, but may have applicability to ITER in that the attenuation constants are less sensitive to bending [35].

HGWs could be employed in ITER TIP for either short or long wavelength implementations. In addition to the obvious advantage of reducing the beam penetration area in the BSMs, the use of bent HGWs can reduce neutron streaming while reducing the number of reflective optics required to bring signals to/from the ITER plasma. The use of HGWs may also circumvent much of the mechanical displacement that would otherwise have to be compensated for.

## 3) ALIGNMENT METHOD

A real-time alignment system is required in both short and long wavelength implementations. The choices here are between a visible light system using a CCD camera as proposed by Kondoh et al. [9] or an IR/FIR system that uses one of the two IF/FIR beams themselves for alignment. The latter has been chosen by Donne et al. for the poloidal system, who propose “that as soon as the reflected signal for an individual chord becomes below a certain detection threshold, the chord is automatically switched from measuring mode to a scanning mode” [15]. In the poloidal system design, one of the turning mirrors is tilted in two directions by ±0.32º, making it possible to scan the beam over an area of 58×58 mm2 at the retroreflector position. This technique, applied to the toroidal system, is highly suited to real-time operation without operator intervention and is the preferred option independent of the choice of operating wavelength. In addition, the self-scanning option has the benefit of being able to directly handle low levels of beam refraction that the ITER simulations predict at ~50 µm.

Previously discussed concerns regarding mirror erosion and redeposition on the input/exit mirror are also valid at visible wavelengths. Mirror reflectivity data collected on Tore Supra and TEXTOR, detailed in Ref. 15, are reproduced in Fig. 8. It is shown that reflectivity changes are most pronounced at the shorter wavelengths, with almost no change observed in the IR region. This would support the use of an IR- or FIR-based system rather than a visible light system for alignment.

Fig. 8. Plots of (left) mirror sample reflectance before (dotted lines) and after (solid lines) exposure inside Tore Supra, and (right) molybdenum mirrors reflectance after exposure to 312 sec of TEXTOR plasmas in a deposition dominated zone (from [15]).

One uncertainty is the role that HGWs may plan in reducing the level of mechanical displacement that would have to be compensated for. This is an area for future study, as this could significantly impact the required scanning area of the real-time alignment system.

## 4) DISPERSION INTERFEROMETER CONCEPT

The dispersion interferometer (DI) concept [36-38] is a two wavelength interferometer scheme in which nonlinear optics are used to frequency double an input laser beam, which then co-propagates with the original un-doubled laser beam. Its primary advantage is low sensitivity to mechanical vibrations. Originally implemented using a Nd:YAG laser (λ=1.06 µm) [36,37], the concept has more recently been implemented on TEXTOR using a CO2 laser (λ=9.57 µm) with good results [38].

A CO2 laser-based DI system for ITER, i.e. a λ=9.6 µm/4.8 µm or a λ=10.6 µm/5.3 µm configuration, would suffer from the same mirror and CCR problems that the Carlstrom et al. design had, and should be rejected for the same reasons. Although the technique could, in principle, be applied to a longer wavelength laser, the doubling efficiency decreases rapidly with increasing wavelength, adversely affecting signal-to-noise ratios. At wavelengths longer than ~10 µm, doublers are not efficient enough for practical application.

# D. RECOMMENDED RESEARCH AND DEVELOPMENT TASKS

The list of recommended research and development tasks to be undertaken depend strongly on the choice of operating wavelength, and are therefore broken down accordingly.

## 1) SHORT WAVELENGTH (~10 µm) TASKS

### OPTICAL DESIGN

As discussed earlier, one of the greatest concerns to the short wavelength design is the polarization effect of carbon film buildup on the input/exit mirror. The TIP optical design should be examined carefully to estimate the polarization changes expected for each of the five probing beams as a function of carbon film deposition depth. Alternatives to the shared, single input/exit mirror configuration should be explored to minimize these polarization changes.

## 2) LONG WAVELENGTH (~50 µm) TASKS

### CH3OD LASER DEVELOPMENT

There needs to be a successful demonstration of a Stark-tuned CH3OD laser to ensure sufficient bandwidth and stability can be achieved at ~50 µm. This could be done by constructing a modified version of a CO2-pumped FIR laser (employing CH3OD rather than CH3OH) for 57.2 µm operation, for comprehensive laboratory testing.

## 3) WAVELENGTH INDEPENDENT TASKS

### INVESTIGATION OF WAVEGUIDE INTERFACE SYSTEM

A study should be made of the selective use of HGWs in the ITER TIP system. Possibilities here include running long lengths or constant-bore or long tapered HGWs throughout the ITER blanket shield modules, or solely using short tapered HGWs in regions such as interfacing joints where alignment can be critical. One issue which would need to be investigated is the amount of polarization and phase changes due to variations in bending curvature due to machine motions.

### SELF-SCANNING ALIGNMENT SYSTEM

A test of the self-scanning laser alignment concept should be undertaken. This could be undertaken at any laser wavelength desired, as the concept is in itself not wavelength dependent.

### FIRST MIRROR DESIGN

Continued research into plasma first mirror and retro-reflector design are essential, as well as research into mirror mitigation/coping methods. Examples of the latter include baffles, shutters, the use of in-situ laser cleaning or laser ablation, and heating of the vacuum vessel (via, for example the use of RF cleaning discharges). These mirrors must also maintain a good optical quality in the presence of plasma radiation, neutral particle bombardment, and nuclear heating.

### IN-SITU CALIBRATION

A plan for in-situ calibration and alignment maintenance/verification should be developed.

# E. PHYSICS ISSUES WHICH COULD BE ADDRESSED

There have been many physics studies conducted using FIR interferometer/polarimeter systems, and many more proposed for future study. The wide range of MHD instabilities that can be monitored using the ITER TIP system range from relatively low frequency energetic particle modes (EPMs) including alpha particles and “fishbones” (a type of EPM), through intermediate frequency toroidal Alfvén eigenmodes (TAEs) and resonant TAEs, to high frequency compressional Alfvén and global Alfvén eigenmodes. Note that these studies require a considerably faster temporal response than that required for ITER density control (~ 1 msec), but such system bandwidths are easily achieved in both short and long wavelength configurations. The TIP electronics should therefore be designed with a wide bandwidth, to the extent that this is possible without sacrificing stability and robustness of the line integral density measurements.

A Fizeau interferometer scheme to measure the toroidal plasma current has been proposed for ITER, based on work being done on the Madison symmetric torus (MST) device [39]. The Fizeau effect provides a measure of the electron current along the line of sight of the beam, and involves measuring the relativistic phase shift of an electromagnetic wave associated with movement of a dielectric medium. This work is interesting, but any changes to the detection geometry that detracts from the basic robustness of the TIP system should be avoided.

# F. LITERATURE CITED

[1]	ITER Technical Basis – Plant Description Document, ITER Technical Document: G A0 FDR 1 01-07-13 R1.0, Sec. 2.6.3.3 - Optical/Infrared Systems.

[2]	R.T. Snider, T.N. Carlstrom, T.D. Hodapp, F.C. Jobes, W.A. Peebles, “Application of Interfer-ometry and Faraday Rotation Techniques for Density Measurements on the Next Generation of Tokamaks,” Review of Scientific Instruments 68 (1), pp. 728-731 (1997).

[3]	T.N. Carlstrom, R.T. Snider, C.B. Baxi, C.L. Rettig, W.A. Peebles, “ITER-EDA Final Report on a Multi-Channel Interferometer and Polarimeter System for Density Measurements,” Phase IV Report for ITER Diagnostic Design: Multi-channel Interferometer/Polarimeter System, (1998). Also published in Diagnostics for Experimental Thermonuclear Fusion Reactors 2, edited by Stott et al., Plenum Press, New York, pp. 193-202 (1998).

[4]	H. Soltwisch, “Current Distribution Measurement in a Tokamak by FIR Polarimetry,” Review of Scientific Instruments 57 (8), pp. 1939-1944 (1986).

[5]	B.W. Rice, “Fifteen Chord FIR Polarimeter System on MTX,” Review of Scientific Instruments 63 (10), pp. 5002-5004 (1992).

[6]	J.H. Rommers, J. Howard, “A New Scheme for Heterodyne Polarimetry with High Temporal Resolution,” Plasma Physics and Controlled Fusion 38 (10), pp. 1805–1816 (1996).

[7]	J.H. Rommers, A.J.H. Donné, F.A. Karelse, J. Howard, “The Multichannel Triple-Laser Interferometer/Polarimeter System at RTP,” Review of Scientific Instruments 68 (1), pp. 1217-1226 (1997).

[8]	H. Park, C.W. Domier, W.R. Geck, N.C. Luhmann, Jr., “Far Infrared Tangential Interferometry/ Polarimetry on the National Spherical Tokamak Experiment,” Review of Scientific Instruments 70 (1), pp. 710-713 (1999).

[9]	T. Kondoh, A.E. Costley, T. Sugie, Y. Kawano, A. Malaquias, C.I. Walker, “Toroidal Interfero-meter/Polarimeter Density Measurement System on ITER,” Review of Scientific Instruments 75 (10), pp. 3420-3422 (2004).

[10]	Y. Kawano, A. Nagashima, T. Hatae, S. Gunji, “Dual CO2 Laser Interferometer with a Wavelength Combination of 10.6 and 9.27 µm for Electron Density Measurement on Large Tokamaks,” Review of Scientific Instruments 67 (4), pp. 1520-1528 (1996).

[11]	Y. Kawano, S-I Chiba, H. Shirai, A. Inoue, A. Nagashima, “First Measurement of Tangential Faraday Rotation of CO2 Laser Wave in a Tokamak Plasma,” Review of Scientific Instruments 70 (2), pp. 1430-1434 (1999).

[12]	Y. Kawano, S-I Chiba, A. Inoue, “Infrared Laser Polarimetry for Electron Density Measurement in Tokamak Plasmas,” Review of Scientific Instruments 72 (1), pp. 1068-1072 (2001).

[13]	V.S. Voitsenya, A.J.H. Donné, A.F. Bardamid, A.I. Belyaeva, V.L. Berezhny, A.A. Galuza, Ch. Gil, V.G. Konovalov, M. Lipa, A. Malaquais, D.I. Naidenkova, V.I. Ryzhkov, B. Schunke, S.I. Solodovchenki, A.N. Topkov, “Simulation of Environment Effects on Retroreflectors in ITER,” accepted for publication in Review of Scientific Instruments 76 (2005).

[14]	H.E. Bennett, “Specular Reflectance of Aluminized Ground Glass and the Height Distribution of Surface Irregularities,” Journal of the Optical Society of America 53, pp. 1389-1394 (1963).

[15]	V.S. Voitsenya, Ch. Gil, V.G. Konovalov, A. Litnovsky, M. Lipa, M. Rubel, A. Sagara, A. Sakasai, B. Schunke, G. De Temmerman, K.Yu. Vukolov, S.N. Zvonkov, P. Wienhold, “The Environment Effect on Operation of In-Vessel Mirrors for Plasma Diagnostics in Fusion Devices, Proceedings of the 12th International Congress of Plasma Physics, Nice, France, October 2004 <http://fr.arxiv.org/abs/physics/0411093> (2004).

[16]	A.J.H. Donné, M.F. Graswinckel, M. Cavinato, L. Giudicotti, E. Zilli, C. Gil, H.R. Koslowski, P. McCarthy, C. Nyhan, S. Prunty, M. Spillane, C. Walker, “Poloidal Polarimeter for Current Density Measurements in ITER,” Review of Scientific Instruments 75 (11), pp. 4694-4701 (2004).

[17]	K. Nakayama, H. Tazawa, S. Okajima, K. Kawahata, K. Tanaka, T. Tokuzawa, Y. Ito, “High-Power 47.6 and 57.2 µm CH3OD Lasers Pumped by Continuous-Wave 9R(8) CO2 Laser,” Review of Scientific Instruments 75 (2), pp. 329-332 (2004).

[18]	S.E. Segre, “Determination of Both the Electron Density and the Poloidal Magnetic Field in a Tokamak Plasma from Polarimetric Measurements of Phase Only,” Plasma Physics and Controlled Fusion 38 (6), pp. 883–888 (1996).

[19]	S.E. Segre, “A Review of Plasma Polarimetry–Theory and Methods,” Plasma Physics and Controlled Fusion 41 (2), pp. R57–R100 (1999).

[20]	S.E. Segre and V. Zanza, “High-Frequency Polarimetric Measurement of Both Electron Density and Poloidal Magnetic Field in Tokamaks,” Plasma Physics and Controlled Fusion 46 (7), pp. 1087–1094 (2004).

[21]	K. Guenther, “Approximate Method to Extract the Pure Faraday and Cotton-Mouton Effects from Polarimetry Measurements in a Tokamak,” Plasma Physics and Controlled Fusion 46 (9), pp. 1423–1441 (2004).

[22]	L. Giudicotti, S.L. Prunty, C. Nyhan, E. Bedin, E. Zilli, L. De Pasqual, “A Polarization Modulation Technique for Far-Infrared Polarimetry in Large Plasmas,” Plasma Physics and Controlled Fusion 46 (4), pp. 681–693 (2004).

[23]	M.A. Walker, R.K. Richards, D.P. Hutchinson, “Transmission Measurements of Selected Liquids in the Far Infrared,” Oak Ridge National Laboratory Report # ORNL/TM-9772 (1985).

[24]	T. Maekawa, T. Minami, K. Makino, S. Tanaka, S. Kubo, M. Iguchi, “Frequency Shift of 1.45 MHz for 337-µm HCN Laser Beam with a Super Rotating Grating,” Review of Scientific Instruments 62 (2), pp. 304-307 (1991).

[25]	T. Akiyama, K. Kawahata, S. Okajima, Y. Ito, S. Tsuji-lio, S. Okamura, “HCN Laser Interferometer/Scattering Measurement on CHS,” presented at the 2005 US-Japan Microwave Imaging Workshop, held at Princeton, NJ on March 21-22, 2005.

[26]	K.C. Lee, C.W. Domier, B.H. Deng, M. Johnson, B.R. Nathan, N.C. Luhmann, Jr., H. Park, “A Stark-Tuned Laser Application for Interferometry and Polarimetry on the National Spherical Torus Experiment,” Review of Scientific Instruments 74 (3), pp. 1621-1624 (2003).

[27]	K. Nakayama, “Development of a Two-Color Interferometer Using Short-Wavelength FIR Lasers for Plasma Diagnostics,” presented at the Princeton Plasma Physics Laboratory Monday Physics Meeting on June 27, 2005, online at <http://nstx.pppl.gov/DragNDrop/Monday_Physics_Meet/2005/06-27-05/Nakayama.pdf>.

[28]	S. Okajima, K. Nakayama, H. Tazawa, K. Kawahata, K. Tanaka, T. Tokuzawa, Y. Ito, K. Mizuno, “Development of Short-Wavelength Far-Infrared Laser for High Density Plasma Diagnostics,” Review of Scientific Instruments 72 (1), pp. 1094-1097 (2001).

[29]	K. Kawahata, K. Tanaka, T. Akiyama, T. Tokuzawa, S. Okajima, K. Nakayama, “Development of New Two Color Laser Diagnostics and CO2 Laser Imaging Diagnostics,” presented at the 2005 US-Japan Microwave Imaging Workshop, held at Princeton, NJ on March 21-22, 2005.

[30]	Tyndall National Institute, NMRC Ireland Division web page, online at <http://www.tyndall.ie/projects/iii-v/schottky.html>.

[31]	Virginia Diodes web page, online at <http://www.virginiadiodes.com/SC2T6.htm>.

[32]	A. Simon, C.I. Lin, “Planar Schottky Barrier Diodes for THz Applications,” 1997 Annual Report of the Technische Universitat Darmstadt, pp. 39-40, available online at <http://www.hf.e-technik.tu-darmstadt.de/forschung/jahresberichte/JB1997/pdf_files/simon_4.pdf> (1997).

[33]	M.C. Gaidis, H.M. Pickett, C.D. Smith, S.C. Martin, R. P. Smith, P.H. Siegel, “A 2.5-THz Receiver Front End for Spaceborne Applications,” IEEE Transactions on Microwave Theory and Techniques 48 (4), pp. 733-739 (2000).

[34] J. Dai and J.A. Harrington, “High-Peak-Power, Pulsed CO2 Laser Light Delivery by Hollow Glass Waveguides,” Applied Optics 36 (21), pp. 5072-5077 (1997).

[35]	D.J. Gibson and J.A. Harrington, “Gradually Tapered Hollow Glass Waveguides for the Transmission of CO2 Laser Radiation,” Applied Optics 43 (11), pp. 2231-2235 (2004).

[36] V.P. Drachev, Yu.I. Krasnikov, P.A. Bagryansky, “Dispersion Interferometer for Controlled Fusion Devices,” Review of Scientific Instruments 64 (4), pp. 1010-1013 (1993).

[37]	F.C. Jobes, N.L. Bretz, “A Prototype Imaging Second Harmonic Interferometer,” Review of Scientific Instruments 68 (1), pp. 709-712 (1997).

[38]	P.A. Bagryansky, A. Khilchenko, A. Lizunov, A. Solomakhin, R. Voskoboynikov, “Dispersion Interferometer for Burning Plasma Experiments,” presented at the Eighth Meeting of the ITPA Topical Group on Diagnostics (ITPA-8) held at Culham, United Kingdom on March 14-18, 2005, available online at <http://www.rijnh.nl/ITPA/ITPA-8.htm> (2005).

[39]	D.L. Brower, W.X. Ding, B.H. Deng, M.A. Mahdavi, V. Mirnov, S.C. Prager, “Fizeau Interferometer for Measurement of Plasma Electron Current,” Review of Scientific Instruments 75 (10), pp. 3399-3401 (2004).