Module 9: Performing HIV Rapid Tests (Demo and Practice)

| Purpose | To provide the participants with necessary knowledge and skills to accurately perform 3 HIV rapid tests and to determine HIV status. |
| --- | --- |
| Pre-requisite Modules | Module 3: Overview of HIV Testing Technologies Module 4: HIV Testing Strategies & Algorithms Module 6: Safety at the HIV Rapid Testing Site Module 7: Preparation for Testing – Supplies & Kits Module 8: Blood Collection – Fingerprick |
| Module Time | 5 ½ hours |
| Learning Objectives | At the end of this module, participants will be able to: Perform 3 HIV rapid tests according to Standard Operating Procedure (SOP) Perform multiple tests simultaneously Accurately interpret individual test results Accurately determine HIV status |

Module Overview

| Step | Time | Activity/ Method | Content | Resources Needed |
| --- | --- | --- | --- | --- |
| 1 | 5 min | Presentation | Module introduction | Slides 1-3 |
| 2 | 15 min | Presentation; Discussion | Slide overview of testing procedures | Slides 4-36; dry room One-pager job aids |
| 3 | 10 min | Activity | Workspace setup | Slide 37; practical room; supplies & test kits |
| Steps 4 & 5 are concurrent sessions. Participants are divided into 2 groups and take turns participating in step 4 and Step 5. |
| 4 | 1 hr 15 min | Demo & practice | HIV rapid tests | Slide 38; practical room; supplies & test kits; known samples |
| 5 | 1 hr 15 min | Presentation; Discussion | Video review of testing procedures | Slides 39-44; videotape; dry room |
| End of concurrent sessions |
| 6 | 10 min | Presentation; Discussion | Video overview of multi-test algorithm | Slides 45-47; videotape; dry room |
| 7 | 2 hrs 15 min | Practice | Conducting multiple HIV rapid tests | Slides 48-49; supplies & test kits; panels of samples; practical room |
| 8 | 5 min | Q&A | Summary | Slide 50 |

Material/Equipment Checklists

- PowerPoint slides or transparencies

- Overhead projector or computer w/LCD projector

- Prepared Flipchart – content outline

- Video and equipment to play the video

- Materials, supplies, and test kits required for rapid testing.  Refer to Section 2: Preparation for quantities needed.

- Panels of blood samples for testing. Refer to Section 2: Preparation for quantities needed and preparation guidance.

- One-pager job aids for HIV rapid tests

- Handouts:

  - Job Aids for the tests in your country’s testing algorithm

  - Video script for HIV rapid tests

  - Practical exercise recording worksheet

Advanced Preparation

- Review Section 2 of Trainer’s Guide – Materials, Supplies, Kits, and Panels.

- Setup trainer Workspace with necessary supplies and kits.

- To better manage the space and facilitate learning, you will need 2 rooms:

- Room A – a dry or clean room for presentation

- Room B – a practical room for demo and practice.

Teaching Guide

| Slide Number | Teaching Points |
| --- | --- |
| 1 | Module 9: Performing HIV Rapid Tests (Demonstration and Practice) DISPLAY this slide before you begin the module. Make sure participants are aware of the transition into a new module. |
| 2 | Learning Objectives STATE the objectives on the slide. STATE that the three HIV rapid test kits that we’re going to demonstrate today are the ones included in the country algorithm; the screening test (name of test), confirmation test (name of test) and tiebreaker test (name of test). |
| 3 | Content Overview EXPLAIN the process and logistics of this module. Two rooms have been set up: a dry room for presentation and a practical room for demonstration and hands-on practice. All participants will watch a presentation together to learn the procedures for three HIV rapid tests. This will take place in the room they are currently in, the dry room (A). They will then all move to the practical room (B) to setup their work stations. Then the participants will split into two groups. Group I will stay in Room B for demo and hands-on practice. This group will learn the individual tests with known samples. Group II will go back to Room A to watch video and review the testing procedures. Then these two groups will switch. After a break, briefly convene the participants in the dry room for the Multiple HIV Tests video and slide presentation to prepare for the next practice session. Then all participants will go to the practical room (B) to practice multi-test algorithm using blind samples. |
|  | ANNOUNCE that there will be a one-hour optional hands-on session for participants who wish for additional help or more practice. It will be offered at (insert time, could be at a lunch break or at the end of a day) |
| Flipchart | WRITE the content outline on a flipchart prior to training. REFER to it frequently to orient participants to where they are in the module. |
| Slide overview of the test procedures (Room A) |
| Customization Notes 4-36 | Instructions for five tests are provided here for reference purpose only. Refer to only those slides that are part of the country algorithm and will be taught in the workshop. Capillus (slides # 4-10) Determine (slides # 11-16) Hema-Strip (slides # 17-23) OraQuick (slides # 24-30) Uni-Gold (slides # 31-36) |
| Customization Notes | If your country uses tests not included in this material, you may still model after the teaching methodology here to teach those tests. Pay attention to: The sequence of instructional events How test demonstration is conducted How hands-on practice is conducted (including specimens provided for the practice) How video is presented |
| 4 | Capillus INFORM the participants that the next few slides address the test procedure for Capillus. |
| 5 | Capillus: Getting Ready EXPLAIN the steps on the slide as you build the slide one step at a time. Use a permanent marker to label test with client’s ID |
| 6 | Capillus: Reagent Preparation EXPLAIN the steps on the slide as you build the slide one step at a time. Add the following information for Step 4: It is important to gently mix reagent. Vigorously shaking will result in air bubbles, which make it difficult to dispense the correct amount in the next step. |
| 7 | Capillus: Collecting Specimen EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following points: The pre-calibrated pipette is included with the kit. The pipette is set to collect 10µl. |
| 8 | Capillus: Applying Specimen to Test Slide EXPLAIN the steps on the slide as you build the slide one step at a time. |
| 9 | Capillus: Getting Results EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following information for Step 11: If a timer is not available, use your watch or clock on the wall to monitor the time. It is a good practice to write down the starting time or projected end time to help you keep track of time. ADD the following information for Step 12: In addition to test results, the tester should also record on appropriate forms the following information: The test date Client identification Name of the person performing the test The name of the test The expiration date Test lot number Control lot number |
| 10 | Capillus: Test Interpretation REMIND the participants how to read the Capillus results. There are two possible test results, reactive or non-reactive. Capillus does not have an internal control line. |
| 11 | Determine INFORM the participants that the next few slides address the test procedure for Determine. |
| 12 | Determine: Reagent Preparation EXPLAIN the steps on the slide as you build the slide one step at a time. Remember to have on hand the chase buffer because it is not included in the kit. When tearing off one test strip it is easier if you bend the unopened strip back and forth, then tear the strip away from the others. Once the test strip is opened, do not touch the testing area. |
| 13 | Determine: Collecting Specimen EXPLAIN the steps on the slide as you build the slide one step at a time. |
| 14 | Determine: Applying Specimen and Buffer to Test Strip EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following information for Step 7: Check to see that the sample is moving along the test strip – this will take a moment. If not, add one more drop of buffer to the strip. But, do not waste buffer. |
| 15 | Determine: Getting Results EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following information for Step 8: If a timer is not available, use your watch or clock on the wall to monitor the time. It is a good practice to write down the starting time or projected end time to help you keep track of time. ADD the following information for Step 9: In addition to test results, the tester should also record on appropriate forms the following information: The test date Client identification Name of the person performing the test The name of test The expiration date Test lot number Control lot number |
| 16 | Determine: Test Interpretation REMIND the participants how to read the Determine results. There are three possible test results: reactive, non-reactive, and invalid. Any visible line in the patient/client window should be read as reactive, when the control line is also visible. Note: A shadow may be seen, but this is not to be understood as reactive specimen. You must see a visible colored line for the specimen to be interpreted as reactive. |
| 17 | Hema-Strip INFORM the participants that the next few slides address the test procedure for Hema-Strip. |
| 18 | Hema-Strip: Getting Ready EXPLAIN the steps on the slide as you build the slide one step at a time. |
| 19 | Hema-Strip: Collecting Specimen EXPLAIN the steps on the slide as you build the slide one step at a time. ENSURE participants understand that either blood from a finger prick or serum from a vial is used, but not both. |
| 20 | Hema-Strip: Buffer Preparation EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following information to Step 5: The buffer vial should be placed in such a way that the foil side is facing up. |
| 21 | Hema-Strip: Applying Specimen and Buffer to Test Strip EXPLAIN the steps on the slide as you build the slide one step at a time. |
| 22 | Hema-Strip: Getting Results EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following information for Step 8: If a timer is not available, use your watch or clock on the wall to monitor the time. It is a good practice to write down the starting time or projected end time to help you keep track of time. ADD the following information for Step 9: In addition to test results, the tester should also record on appropriate forms the following information: The test date Client identification Name of the person performing the test The name of test The expiration date Test lot number Control lot number |
| 23 | Hema-Strip: Test Interpretation REMIND the participants how to read the Hema-Strip results. There are three possible test results: reactive, non-reactive, and invalid. Any visible line in the patient/client window should be read as reactive, when the control line is also visible. |
| 24 | OraQuick INFORM the participants that the next few slides address the test procedure for OraQuick. |
| 25 | OraQuick: Getting Ready EXPLAIN the steps on the slide as you build the slide one step at a time. REMIND participants to have on hand the blue stand that is included in the kit. To prevent contamination, leave the testing device in its pouch, until needed Do not touch the flat pad Carefully place the vial in the stand. Do not snap the vial in the stand as splashing may occur. |
| 26 | OraQuick: Collecting Specimen (Blood or Oral Fluids) EXPLAIN the steps on the slide as you build the slide one step at a time. REMIND participants to discard the loop in a bio-hazard container. |
| 27 | OraQuick: Transferring Specimen (Blood Only) EXPLAIN the steps on the slide as you build the slide one step at a time. EMPHASIZE this step is for blood specimens only. Oral fluids do not need this step. |
| 28 | OraQuick: Inserting Test Device Into Buffer Vial EXPLAIN the steps on the slide as you build the slide one step at a time. Do not remove the device from the vial until you are done reading the results |
| 29 | OraQuick: Getting Results EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following information for Step 8: If a timer is not available, use your watch or clock on the wall to monitor the time. It is a good practice to write down the starting time or projected end time to help you keep track of time. ADD the following information for Step 9: In addition to test results, the tester should also record on appropriate forms the following information: The test date Client identification Name of the person performing the test The name of test The expiration date Test lot number Control lot number |
| 30 | OraQuick: Test Interpretation REMIND the participants how to read the OraQuick results. There are three possible test results: reactive, non-reactive, and invalid. Any visible line within the triangle of the Test(client) area should be read as reactive, when the control line is also visible within the C triangle. |
| 31 | Uni-Gold INFORM the participants that the next few slides address the test procedure for Uni-Gold. |
| 32 | Uni-Gold: Getting Ready EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following information to Step 1: The plastic pipette is included in the kit. |
| 33 | Uni-Gold: Collecting Specimen EXPLAIN the steps on the slide as you build the slide one step at a time. ENSURE participants understand that either blood from a finger prick or serum from a vial is used, but not both. |
| 34 | Uni-Gold: Adding Specimen and Reagent to Test Device EXPLAIN the steps on the slide as you build the slide one step at a time. Check to ensure the sample immediately moves along the test device. If not, add one more drop of buffer. |
| 35 | Uni-Gold: Getting Results EXPLAIN the steps on the slide as you build the slide one step at a time. ADD the following information for Step 6: If a timer is not available, use your watch or clock on the wall to monitor the time. It is a good practice to write down the starting time or projected end time to help you keep track of time. ADD the following information for Step 7: In addition to test results, the tester should also record on appropriate forms the following information: The test date Client identification Name of the person performing the test The name of test The expiration date Test lot number Control lot number |
| 36 | Uni-Gold: Test Interpretation REMIND the participants how to read the Uni-Gold results. There are three possible test results: reactive, non-reactive, and invalid. Any visible line in the patient/client window should be read as reactive, when the control line is also visible. |
| Logistics Notes | Show the following 2 slides before moving participants out of the dry room. |
| 37 | Activity: Workspace Setup INFORM all participants that they are going to setup their own Workspaces in the practical room where the lab session will be conducted. READ the instructions on the slide. |
| 38 | Hands-On Practice: Individual Tests EXPLAIN the instructions for hands-on practice session. EMPHASIZE that participants should not discard their test devices until the instructor has checked them. |
| Workspace setup (Room B) |
| Activity 10 minutes | ASK participants to move to Room B (the practical room). After Workspaces have been set up: DIVIDE the participants into two groups: Group I and Group II. ASK Group I to stay in the practical room for demo and practice while Group II goes back to the dry room for video presentation. |
| Logistics Notes | Group I (demo and practice) and Group II (video) activities happen concurrently. After 1 hour 15 minutes, switch the groups so that Group I will move to the dry room to watch the video while Group II will move to the practical room for demo and practice. |
| Concurrent Session: Demo and Practice (Room B) |
| TIPS | To facilitate retention, follow the sequence below for the demo and practice session: Instructor demonstration (Test 1) Participant practice (Test1) Instructor demonstration (Test 2) Participant practice (Test 2) Instructor demonstration (Test 3) Participant practice (Test 3) To save time, after instructor’s demonstration of one test, participants can start their hands-on practice during the incubation time of the instructor’s test. Therefore, each set of demo and practice should take 20-25 minutes. The total time for the demo-and-practice session will be no more than 75 minutes. |
| Demonstration 15 minutes/test | DEMONSTRATE the procedure for _______ (test 1) using specimens labelled as positive and negative. ASK participants to gather around you while you demonstrate. |
| TIPS | When demonstrating, Commit to memory each step in the testing procedure and relevant notes Have one-pager handy for easy reference Show each step slowly and methodically Talk out loud as you perform, but keep your explanation brief and clear If you repeat the procedure, make sure each time you do exactly the same thing so you don’t confuse the participants Point out commonly made mistakes during incubation time Allow time for Q &A |
| Practice 15 minutes/test | ASK participants to pull out of their manual the one-page job aids for easy reference during the practice session. (Note: There will not be enough room on tabletop for supplies and binder.) REMIND participants of the instructions for the hands-on practice. |
| TIPS | INFORM participants that: If timers are not available, the clock on the wall is for their use during testing practice. They should feel free to raise their hands if they need anything or have a question. They should record results using worksheet provided. MONITOR participant performance during the practice. Provide as much personal attention and assistance as possible. Make sure participants are following the right procedures. |
| Logistics Notes | After one round of demonstration and practice (for one test), repeat the same process for test 2 and test 3. |
| Concurrent Session: Video Presentation and Discussion (Room A) |
| Customization Notes 39-44 | This video includes segments for 6 tests. Show the video segments and slides for relevant tests only. Use the video counter below to help you cue up the right segment. Videotape Section Counter Reading Capillus 14:35 Determine 20:25 Hema-Strip 25:33 OraQuick 30:35 Uni-Gold 36:55 Virocheck 41:07 |
| TIPS | Show the video one segment at a time. Never show the video segments all at once without discussing in between. Use the relevant slide both before and after each segment. The slide serves the following purposes: Before the video, it directs participants’ attention to the key messages in the video. After the video, you can use the questions on the slide to review key learning points. |
| 40 Video | Video: Capillus FOLLOW the procedure for showing the video segment. INFORM participants that they are going to watch a video about (Test Name). POINT OUT the questions on the slide, which you expect them to answer at the end of the video segment. ENCOURAGE participants to take notes while watching the video. PLAY the video. ANSWER any questions participants have. DEBRIEF by asking participants to answer the questions on the slide. |
| 41 Video | Video: Determine FOLLOW the standard procedure (see above) for showing the video segment. |
| 42 Video | Video: Hema-Strip FOLLOW the standard procedure (see above) for showing the video segment. |
| 43 Video | Video: OraQuick FOLLOW the standard procedure (see above) for showing the video segment. |
| 44 Video | Video: Uni-Gold FOLLOW the standard procedure (see above) for showing the video segment. |
| End of Concurrent Session |
| Multiple HIV tests video and slide presentation (Room A) |
| Customization Notes 45 | Depending on the in-country testing algorithm and tests used, the video segment “Multiple HIV Tests” may not apply. However, it provides a good model for you to modify the content with in-country examples. If you choose not to show this video segment, make sure you cover your discussion with all the key points from the video. |
| Video 45 | Video: Multiple HIV Tests FOLLOW the standard procedure for showing video. DISPLAY the slide before and after the video. Videotape Section Counter Reading Multiple HIV Tests 45:30 REFRESH participants’ memory of the in-country testing algorithm. FACILITATE a discussion of the procedures for performing the specific in-country multi-test algorithm. Discussion questions may include: What preparation work is required? How should you set up the testing area? Which test should you do first? Why? How should you set the timer for individual tests? For multiple tests? |
| Customization Notes 46-48 | Customize slides #46-48 by: Displaying your country’s approved algorithm in slide #46. Deleting slide #48 if your country uses a parallel algorithm. Deleting slide #47 if your country uses a serial algorithm. |
| 46 | Country Algorithm REVIEW the country’s algorithm. |
| 47 | Possible Outcomes in a Parallel Algorithm REFER participants to the Possible Outcomes handout in their manual. ASK them to take it out of the manual for easy reference during the next practice. REVIEW the slide with the participants. |
| 48 | Hands-On Practice: Multi-Test Algorithm INFORM participants of the next activity. EXPLAIN the hands-on practice: Participants will conduct tests simultaneously following the testing algorithm. Participants will do this 5 times, each time using the designated blind specimen. Participants will record individual test results and determine the HIV status of each specimen. |
| 49 | Hands-On Practice: Multi-Test Algorithm (Cont’d) READ the instructions on the slide. INFORM participants that they will now move to the practical room for the demo and practice of multi-test algorithm. |
| Logistics Notes | After showing the Multiple HIV Tests video and slides 45-48, move the participants to the practical room for the demo and hands-on practice session. |
| Multiple HIV test hands-on practice session (Room B) |
| Practice 2 hours | INFORM participants that: If timers are not available, the clock on the wall is for their use during testing practice. They should feel free to raise their hands if they need anything or have a question. MONITOR participant performance during the practice. Provide as much personal attention and assistance as possible. Make sure participants are following the right procedures. |
| 15 minutes | DEBRIEF the practice by: Discussing the results Pointing out the commonly-made mistakes you have observed during the practice session Asking participants to share their experience and key learning from the practice Answering any questions participants might have |
| 50 | Summary ASK participants to answer the questions on the slide. ANSWER any questions participants may have. IDENTIFY participants who will be in the optional hands-on session. |