### Post-Sequencing BET Cleanup Protocol

### (384-well plates for MB4000 and ABI3730)

Automated Process (Biomek FX)

Version Number:		3.3

Production Start Date:	12/15/03

Version 3.3 Date:		7/21/05

Authors:	Christine Lansang, Nancy Hammon, Steven Wilson, Catherine Gordon, Arshi Khan, Amber Nivens

Reviewed/Revised by:	Cailyn Spurrell

# Summary

The purpose of this procedure is to automate the BET cleanup protocol using Biomek FX robots. This process is the post-sequencing reaction cleanup procedure for the BigDye and ET-Terminator cycle sequencing reactions of RCA and Fosmid production plates using magnetic beads, 70% ethanol, and TEG reagents. The program uses P30 pipette tips, custom magnets, four plate stackers, a series of deck positions and a wash station. The final product is a plate with clean resuspended, dye labeled DNA fragments to run on the MegaBACE 4500 or the ABI 3730xl sequencer

s.

# Materials & Reagents

| Materials/Reagents/Equipment Disposables 384-well PCR Microplate, Clear 384-well PCR Microplate, Clear 50-ml Centrifuge Tube 50-ml Serological Pipet 25-ml Serological Pipet Biomek AP 384 Pipette tips 250-μl Matrix Pipette tips PlateLoc Pierceable Seal MasterFlex Pharmed Tubing MasterFlex Silicone Tubing Ismatec 3-stop tubing (2.29mm ID) Ismatec 3-stop tubing (2.79mm ID) Clear Adhesive Plate Sealers Tri-Flow Industrial lubricant Reagents Tetra-Ethylene Glycol Seradyn Magnetic Carboxylate-Modified Beads 100% Ethanol (200 Proof) Milli-Q Water Materials/Reagents/Equipment Equipment Biomek FX Vari-Mixer MagneSil Magnetic Separation Unit 384-well High Field Plate Magnet Centrifuge 5810 Spectramax 384 Plus Spectrophotometer PlateLoc Thermal Plate Sealer Easypet 8-channel Pipette (125μl) | Vendor Axygen/ISC BioExpress Applied Biosystems Corning/VWR Falcon/VWR Falcon/VWR VWR Matrix Velocity 11 Cole Parmer Instrument Co. Cole Parmer Instrument Co. Cole Parmer Instrument Co. Cole Parmer Instrument Co. Edge BioSystems - Sigma-Aldrich Agencourt AAPER Alcohol (LBNL) Millipore Milli-Q System Vendor Beckman Coulter/VWR Clay Adams Promega JGI Eppendorf Spectramax Velocity 11 Eppendorf Matrix | Stock Number T-3059-1CS 4343814 430290 357550 357525 BK719222 7151 06644-001 06485-36 96400-24 EW-95603-44 EW-95603-48 48461 - 110175-KG 44152105050350 6810-13289 - Stock Number 719222 - - - - - - - - |
| --- | --- | --- |

# EH&S

JGI employee performing this procedure must wear a lab coat and gloves.  In situation were there might be a chance of an accidental splash to the eyes, safety glasses must be worn.

# Calibration Check

Biomek FX

Tips should be changed every 2-3 days or when they appear dirty. A tip QC should be performed when new tips are installed.  Additionally, the dye QC test requires a plate weight reading. Refer to the QC Protocol for the Biomek FX Liquid Handler.

# Procedure

NOTE: All reagents/stock solutions should be prepared prior to the start of the procedure.

1. Preparing samples for Biomek FX cleanup procedure

  - Remove MegaBACE 4500 and/or ABI 3730 plates from the TEG Backlog deli (the 4oC fridge # 407 in Rm141).

NOTE 1: Twenty plates can be loaded per Biomek FX per run.

NOTE 2: ABI plates need to be loaded with other ABI plates and MegaBACE plates need to be loaded with other MegaBACE plates. Fosmid plates (although they are ABI plates) need to be loaded with other Fosmid plates. However, if running 20 ABI plates, you can have 10 consecutive Fosmid plates and 10 consecutive ABI plates. Do not mix Fosmids and ABI plates within a group of 10 plates.

  - Quick spin the plates in the centrifuge for 1 minute at 1,000rpm.

  - Remove the seals from the plates and throw the seals away in the autoclave waste.  Create stacks of no more than 20 plates per Biomek FX.

NOTE 1: Do not shake plates and dislodge sample.

NOTE 2: If source plates are not going to be loaded onto the Biomeks within 10 minutes, place a clear seal (labeled with correct Biomek FX #) on the top plate in the stack and store plates in the TEG Backlog deli.

2. Creating barcodes for destination plates

  - For MegaBACE 4500 plates, go to the Venonat database (http://venonat.jgi-psf.org/psf/PSF.home) ( click on RCA DB ( Step 5 Scan (MegaBACE) Plates for TEG:

    - Select the correct information from the pull down menus: Operator, TEG Instrument, Bead Lot, TEG Lot, Ethanol Lot, Clamps Used?, Posting Type.

    - Use the barcode scanner to scan in the stack of source plates.

    - Click on SUBMIT & Print the conformation page.

    - Click on Print Barcodes.

    - Select RCA (R140) and print.

    - Affix the barcodes to new Axygen plates (on the numbered and indented side of the plate), creating destination plate stacks that match the order of the source plate stacks.

NOTE 1: Wear gloves when affixing barcodes to destination plates.

NOTE 2: This step may be performed after the Biomek FX program has started running.

NOTE 3: Mark the barcodes for priority plates with a red line. Mark plates that come with a plate request with “PR”.  Mark Microbe QC plates with “MQC”.  Mark any unpaired plate with a blue “S” & when the run is complete place in the bin marked unpaired in the megabase backlog freezer in 140.

NOTE 4: After labeling the destination plates, assure that the order of source plates matches the order of destination plates found in each stack. This can be done by double checking the plates with the printed conformation page, placing the source stack and destination stack side by side and visually scanning over the plates, double checking that the first and last plate of each source and destination stack are matches, or (if you’ve already loaded the source plates onto the Biomek and have started the run), double check that at least one of the first destination plates to be loaded onto the Biomek matches at least one of the first source plates already being processed on the Biomek.

NOTE 5: If destination plates are not going to be loaded onto the Biomeks within 10 minutes, place a clear seal (labeled with correct FX #) on the top plate of each stack.

  - For ABI 3730 plates, go to the Venonat database (http://venonat.jgi-psf.org/psf/PSF.home) ( click on RCA DB ( Step 5 Scan (ABI) Plates for TEG (Associations):

    - Repeat steps 2.1.a – 2.1.e.

    - If using the CyBio Print & Apply to label plates refer to the addendum at the end of this protocol and continue with step 3 Biomek FX Setup of this protocol, otherwise continue with step 2.2.c.

NOTE:  Label one ABI batch per shift using the CyBio Print & Apply.

    - Affix the barcodes to the long side of new ABI plates (the side without an ABI barcode), creating destination plate stacks that match the order of the source plate stacks.

NOTE 1:  Wear gloves when affixing barcodes to destination plates.

NOTE 2: This step may be performed after the Biomek FX program has started running.

NOTE 3:  Mark Fosmid barcodes with a regular green line & priority barcodes with a red line. Mark plates that come with a plate request with “PR” & mark Microbe QC plates with “MQC”.  If a pair is split for some reason mark the ABI barcode with a blue “S” and place in the ABI unpaired bin in the ABI backlog deli.

NOTE 4: After labeling the destination plates, assure that the order of source plates matches the order of destination plates found in each stack. This can be done by double checking the plates with the printed conformation page, placing the source stack and destination stack side by side and visually scanning over the plates, double checking that the first and last plate of each source and destination stack are matches, or (if you’ve already loaded the source plates onto the Biomek and have started the run), double checking that at least one of the first destination plates to be loaded onto the Biomek matches at least one of the first source plates already being processed on the Biomek.

NOTE 5: If plates are not going to be loaded onto the Biomeks within 10 minutes, then place a clear seal (labeled with correct FX #) on the top plate of each stack.

3. Biomek FX Setup

  - Check to make sure that the main water valve of the Biomek FX is in the ON position.

  - Set up the reagent reservoirs:

    - Attach the 500-ml reagent bottles (BET, ethanol, Milli-Q water) to the orange caps connected to the tubing.

NOTE 1: Total of 12 bottles needed for all 4 Biomek FX instruments.

NOTE 2:  Use freshly made BET reagent for the first shift of the day, and recycled BET reagent for the 2nd & 3rd shift of the day.  Refer to Reagent/Stock Preparation to see how to prepare BET solutions.

Checkpoint: Check to make sure that the waste bottle is in place.

    - Snap the Ismatec Return and Supply applicators down on the Ismatec pump.

Checkpoint: Check to see if the Ismatec tubing needs to be moved or changed before snapping down the applicators. If it is a Monday, remember to move and date the MasterFlex tubing or enter a work request if it needs to be changed.

NOTE: Refer to the Biomek Instrument Maintenance section for moving or changing the Ismatec tubing or moving the MasterFlex tubing.

    - Turn the Ismatec pump on to the forward direction.

Checkpoint: Make sure the Ismatec pump speed is set at 25, the reagent reservoirs are in the correct positions, and that the reagents are flowing into the reservoirs.

  - If this is the first run of the day, initialize the Biomek carousel by using the remote keypad (Stacker Carousel Pendant) connected to the carousel.

    - Press the Remote Local button. The display on the keypad should read “Keypad Control Only.”

    - Press the Initialize button. A message will appear on the keypad reading “Press initialize key to initialize or press another key to cancel.” Press the Initialize button again.

    - The carousel will slowly rotate 360(. When finished rotating press the Remote Local button to return the stacker control to the Biomek FX. The display should read “Press the Remote Local Key.”

  - Load the desired number of source plates into Stacker A, starting at the bottom of the hotel with labels facing inward (A1 position out).

NOTE 1:  Each stacker may hold a maximum of 20 plates.

NOTE 2: Manual control of the stacker for loading can be accomplished by using the Stacker Carousel Pendant connected to the carousel. The display on the keypad should read “Keypad Control Only” for manual control.  If the display reads “Press the Remote Local Key,” press the Remote Local button on the keypad to bring the display to read “Keypad Control Only.”  Once in manual control, use the Rotate Right and Rotate Left buttons on the keypad to rotate the carousel.

  - Load the set of destination plates into Stacker C.

NOTE: This step does not need to be done right away. It can be done within the first 45 minutes of starting the run.

  - Open the Biomek FX program on the desktop.

  - Open the appropriate program file:

- Click on File ( Open ( Methods ( TEG_FLAT+MAGS+BCR_ SW_0909.04.bmt

- Click on Open.

  - If this is the first run of the day, make sure that the Biomek power is turned on and then home all axes by going to Tools ( selecting Manual Control ( select Home All Axes.

NOTE 1: Make sure the grippers on POD 2 are up before you home the axes.

NOTE 2: A pop-up box with a series of conditions will appear, make sure all conditions are met before hitting OK.

NOTE 3: Homing all axes takes approximately 30 seconds. When the pods are no longer moving you can exit the Home All Axes program.

  - After checking the levels of the reagent reservoirs, click on the RUN (green arrow) button.

NOTE: The display on the Stacker Carousel Pendant should read “Press the Remote Local Key” for Biomek FX control.  If the display reads “Keypad Control Only,” press the Remote Local button on the keypad to return the display to “Press the Remote Local Key.”

  - Upon prompt, enter the number of source plates to be run and select OK.

  - Another prompt will ask for the number of MegaBACE plates.

- If running MegaBACE plates:

      - Enter the number of MegaBACE source plates when prompted and select OK.

NOTE: ABI and MegaBACE plates cannot be mixed. If the number of MegaBACE plates is greater than 0, all plates will be run as MegaBACE plates.

- If running ABI plates:

- Leave the default 0 for the number of MegaBACE plates and select OK.

- Look at the Biomek FX deck layout and make sure it matches the setup on the computer diagram.

Checkpoint: Make sure that the reservoirs are full, the clamp switch is on, the Ismatec pump is on and running, the FX controller device box switch is on automatic, and the water source switch is on.

- When ready, click OK, and observe Biomek FX to make sure tips are aspirating and transferring liquid.

NOTE 1: The entire process takes approximately 60 minutes for every 10 plates loaded. When running 20 plates, the total run time is approximately 2 hrs for ABI plates and 2 hrs 15 min for MegaBACE plates.

NOTE 2: The Biomek FX has a light curtain, which protects the deck area during operation. If this curtain is crossed during operation, the protocol will stop and prompt a light curtain error. To reset the light curtain, click on the “OK” button on the prompt window.

NOTE 3: Plate counts are recorded in an excel spreadsheet entitled “TEG Plate Count.”  At some point during the run enter in the correct number of ABI or MB plates processed for that batch. The “TEG Plate Count” spreadsheet can be found in Octopus(Prod Seq(RCA(TEG Plate Count.

Checkpoint: It is important to look for consistent levels of reagents at the respective stages. One of the most important steps to observe is the ethanol wash procedure. After adding and removing the ethanol from the plates, there should be no ethanol left over in the wells, only dry magnetic beads on the bottom of wells.

  - When the BET cleanup is complete, carefully unload plates from Stackers B and D as to not mix up the order.

NOTE: The source plates will be in stacker B and the destination plates will be in stacker D.

  - After each set of 10 plates, a Barcode Verification and Submission program window will open:

- If the window is gray, then all barcodes were successfully scanned and no mix-ups occurred:

  - Fill in the requested information: Operator’s initials, and any comments.

  - Click on Proceed.

  - Click on Upload.

- If the window is yellow, the barcode scanner did not scan one or more barcodes:

  - Carefully identify the plates that need to be scanned into the form (matching source and destination plates) and scan them in manually using the handheld scanner. Select the appropriate reason from the drop down menu of why the barcode was not scanned in on the biomek.

  - Repeat steps 3.13.a.i – 3.13.a.ii.

- If the window is red, then a mix-up occurred. Notify a supervisor or instrumentation specialist immediately to determine how to proceed. Normally, if the mix-up is simple in nature and can be fixed with 100% confidence, the operator can correct the mix-up by switching the barcodes on the destination plates. If the mix-up is complex or cannot be fixed with 100% confidence, then all of the plates involved must be failed. Check the Do not submit box next to the affected plates, and fill in the requested information upon submitting. These plates must also be scanned into the Enter Fails or Comments form in the RCA DB as failures. Print out a copy of these failed plates and give to Sequencing Chemistry. Mark on the sheet if these are ABI or MegaBACE plates. An Incident Report must also be written.

  - When the BET cleanup is complete and the barcodes verified:

- Discard the source plates from Stacker B into the autoclave waste.

- Prepare and store the plates for sequencing:

NOTE: For passing on unpaired plates, label the JGI barcode with a blue “S” and place it in the appropriate refrigerator or freezer specified below within the “unpaired” bin.

- For the MegaBACE plates:

  - Seal the plates with clear plate seals.

  - Store the plates (match FW and RV pairs) in the MegaBACE Backlog freezer (the -20oC freezer 201 in Rm140) in bins labeled with MB & the BET cleanup date.

NOTE: In some situations, if you don’t have time to seal each plate, you can temporarily store plates in the TEG backlog deli. You should seal the top plate of each stack and keep plates there for no longer than 1-2 hours.

- For the ABI plates:

  - Seal the plates with the PlateLoc foil sealer.

  - Visually check each plate to make sure all wells have sample.

  - Store the plates (match FW and RV pairs) in the ABI Backlog deli (the 4oC fridge #406 in Rm141) in bins labeled with ABI & the BET cleanup date.

  - The Biomek FX can now be loaded for another run.

4. Biomek FX Cleanup

If there are no more runs for the day, then the following steps should be carried out:

  - Switch the Ismatec pump to the “reverse” direction and prop the reservoirs at an angle so that their reagents are fully emptied. Stop the pump once lines are empty.

  - Remove the orange caps from the reagent bottles.

    - On Monday-Thursday consolidate the remaining BET solution from all 4 biomeks into one bottle (labeled BET & current date) and store it in the 4° C TEG deli located in Rm 141.

NOTE: After consolidating the remaining BET solution, make sure to track the amount of solution in the BET Tracking spreadsheet.  The “BET tracking” spreadsheet can be found in Octopus(Prod Seq(RCA(TEG Plate Count.\\Octopus\ProdSeq\RCA\TEG Plate Count.xls

    - On Friday dump the remaining BET solution into the red Hazardous Waste Container.

NOTE 1: If the Hazadous waste container is empty, you will need to mark it with the    date that liquid started being collected in the container.

NOTE 2: Make sure to note the amount of BET solution that was thrown out under the Comments section in the BET tracking spreadsheet.

4.3	       Replace the BET bottle with the Milli-Q water bottle.

4.4       Release the water and ethanol pump white applicators and continue to pump the water in   the forward direction to clean the BET reservoir with water.

NOTE: Pump in enough water to cover the bottom of the BET reservoir.

4.5        Reverse the water flow from the BET reservoir and prop the reservoir at an angle to ensure the removal of any residual water.  Stop the pump once lines are empty.

4.6        Remove the BET cap from the Milli-Q water bottle, dump the Milli-Q water bottle contents into the sink, and attach the 70% ethanol bottle to the BET cap.

4.7         Pump the ethanol in the forward direction to clean the BET reservoir.

NOTE: Pump in enough ethanol to cover the bottom of the BET reservoir.

4.8        Wipe the BET, water, and ethanol reservoirs with damp paper towels.

NOTE: Make sure to wipe BET trough separately than the other troughs as to not contaminate the other troughs with residual beads.

4.9	      Reverse the ethanol flow from the BET reservoir and prop the reservoir at an angle to ensure the removal of any residual ethanol.  Stop the pumps once the lines are empty.

4.10      Remove the BET cap from the ethanol bottle and release the BET Return and Supply white applicators.

4.11      Wipe all plate flatteners with 70% ethanol.

4.12       Pour the 70% Ethanol bottle contents into the red Hazardous Waste Container.

NOTE: If the waste container is empty, you will need to mark it with the date that liquid started being collected in the container.

4.13     Remove the orange cap from the waste bottle and dispose of the waste into the designated red Hazardous Waste Container.

  - Place the waste bottle back on the Biomek.

NOTE: Replace the waste bottle with a new autoclaved 2L bottle every Friday.

4.15      Place all bottles in the autoclave bin located next to the sink.

4.16     Close out the software window by selecting File ( Exit. If prompted to save changes, click on NO.

# Reagent/Stock Preparation

NOTE: Wear a lab coat and gloves when preparing reagents. You MUST wear goggles when preparing 70% Ethanol.

Washed Bead Reagent

NOTE: Three bottles of beads, needed to make 16 tubes of washed bead reagent, should be taken out of the QC deli 3 days ahead of time and placed on the Vari-Mixer.

- Add 15ml of thoroughly mixed Seradyn beads to a 50-ml Falcon tube.

- Add 35ml of Milli-Q water.

- Place the tube on the Promega magnet for 15 minutes or until the solution is almost clear.

- With the tube on the magnet, pour off the water.

- Repeat steps 2 – 4 two more times.

NOTE: Rotate the tube 180o after each wash to ensure that all of the beads get thoroughly washed.

- Add 15ml of Milli-Q water to the Falcon tube.

- Label tube: Washed beads, Bead lot number, date, operator initials.

- Place on a Vari-Mixer until ready for use.

Fresh BET Solution Preparation (prepared daily during first shift)

- The following reagents are required to process 240 Axygen 384-well plates (3 batches of 20 plates on each of the 4 Biomeks). In a 2-L bottle combine the following reagents:

NOTE: You will have approximately 349mls of BET solution left over at the end of the 3rd batch.

| Reagent | Volume (16X) |
| --- | --- |
| 100 % Ethanol (200 proof) | 1024 ml |
| Milli-Q water | 112 ml |
| Tetra Ethylene Glycol | 102.4 ml |
| Washed Reagent Bead | 32 ml |

- Consult the BET and 70% EtOH Solution Preparation sheet located at the bottom of this section to determine how much solution should be prepared based on the number of plates that need to be processed.

- Dispense the appropriate volume of ethanol and water into the 2-L bottle.

- Add the appropriate amount of TEG. After dispensing TEG with a serological pipet, aspirate the TEG/water/ethanol mixture into the pipette to remove any residual TEG on the walls of the pipet.

- Add the washed bead reagent. Cap the bottle and mix by inverting and shaking the bottle.

NOTE:  If reagents are not going to be used within the next 20 minutes store the BET in the ABI Backlog deli in Rm141.

Recycled BET Solution Preparation (prepared daily for 2nd & 3rd shift)

- Obtain the previous day’s consolidated BET solution from the TEG deli located in Rm 141.

- Consult the BET and 70% EtOH Solution Preparation sheet located at the bottom of this section to determine how much solution should be prepared based on the number of plates that need to be processed.

- Subtract the amount of leftover BET solution from the total amount of BET needed to determine the amount of fresh BET to prepare.

- Dispense the appropriate volume of ethanol and water into the 2-L bottle.

- Add the appropriate amount of TEG. After dispensing TEG with a serological pipet, aspirate the TEG/water/ethanol mixture into the pipette to remove any residual TEG on the walls of the pipet.

- Add the washed bead reagent.

- Add the consolidated BET solution from the previous day. Cap the bottle and mix by inverting and shaking the bottle.

NOTE:  If reagents are not going to be used within the next 20 minutes store the BET in the ABI Backlog deli in Rm141.

70% Ethanol Wash Solution (prepared fresh daily)

Consult the BET and 70% EtOH Solution Preparation sheet located in the BET Cleanup bay to determine how much solution should be prepared based on the number of plates that need to be processed. Add the following reagents to a 2-L bottle (to process 240 Axygen 384-well plates) and mix by inverting and shaking:

| Reagent | Volume (22X) |
| --- | --- |
| 100 % Ethanol (200 proof) | 1540 ml |
| Milli-Q water | 660 ml |

NOTE 1:  EtOH bottles can be found and signed out of the Flammable Cabinet in Rm 141.

NOTE 2:  If reagents are not going to be used within the next 20 minutes store the 70% EtOH in the ABI Backlog deli in Rm141.

NOTE 3: Graduated cylinders must be covered with foil when not in use. Funnels should be washed once per week.  Graduated cylinders should be washed and autoclaved once per week.

|  |  | BET and EtOH Solution Preparation |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| BET Solution | Volume | 2X | 4X | 5X | 8X | 10X | 16X | 20X | 22X | 24X | 26X | 28X |  |
| Ethanol 200 proof | 64 ml | 128 | 256 | 320 | 512 | 640 | 1024 | 1280 | 1408 | 1536 | 1164 | 1792 |  |
| Milli-Q Water | 7 ml | 14 | 28 | 35 | 56 | 70 | 112 | 140 | 154 | 168 | 182 | 196 |  |
| Tetra Ethylene Glycol | 6.4 ml | 12.8 | 25.6 | 32 | 51.2 | 64 | 102.4 | 128 | 140.8 | 153.6 | 166.4 | 179.2 |  |
| Washed Seradyn beads | 2 ml | 4 | 8 | 10 | 16 | 20 | 32 | 40 | 44 | 48 | 52 | 56 |  |
| Total Volume (ml) | 79.4 | 158.8 | 317.6 | 397 | 635.2 | 794 | 1270.4 | 1588 | 1746.8 | 1905.6 | 2064.4 | 2223.2 |  |
| 70% EtOH | Volume |  |  |  |  |  |  |  |  |  |  |  |  |
| Ethanol 200 proof | 70 ml | 140 | 280 | 350 | 560 | 700 | 1120 | 1400 | 1540 | 1680 | 1820 | 1960 |  |
| Milli-Q water | 30 ml | 60 | 120 | 150 | 240 | 300 | 480 | 600 | 660 | 720 | 780 | 840 |  |
| Total Volume (ml) | 100 | 200 | 400 | 500 | 800 | 1000 | 1600 | 2000 | 2200 | 2400 | 2600 | 2800 |  |

# Instrument Setup & Maintenance

Biomek FX

Moving the Ismatec Tubing:

NOTE: When moving or changing the Ismatec tubing, make sure the Biomek is not running.

    - Switch the Ismatec pump off.

    - One at a time, release an Ismatec pump white applicator, dissociate it from its tubing, and move the white applicator over one notch.

    - After all of the white Return and Supply applicators have been moved, sparingly spray the bottom of the tubing that is now under the white applicators with Tri Flow industrial lubricant.

NOTE: Using liberal amounts of lubrication spray can reduce tubing lifetime.  Tubing particles can gather beneath clamp and tear tubing.

    - Switch the Ismatec pumps off.

    - Re-attach the Ismatec pump white applicators and snap down onto the Ismatec pump rotating device.

NOTE: Make sure to place the applicators back in the same order.

    - Switch the Ismatec pump on to the forward direction.

    - Mark that the tubing has been moved.

Changing the Ismatec Tubing:

  - Switch the reagent Ismatec pump to the reverse direction and prop the reservoirs at an angle so that their reagents are fully emptied.

  - Once all reservoirs and tubing are emptied, switch the Ismatec pump off.

  - One at a time, release an Ismatec pump white applicator, dissociate it from its tubing, remove the old tubing and place it in the autoclave waste.

  - Wipe tubing residue from each white clamp with a dry paper towel.  This will help extend tubing lifetime.

  - Replace with new tubing of the same color and re-attach the white applicator to the tubing.

NOTE: All supply lines should have purple and gray tubing; all return lines should have purple and white tubing.

  - After all of the tubing has been changed, sparingly spray the bottom of the tubing that is now under the white applicators with Tri Flow industrial lubricant.

  - Switch the Ismatec pump off.

  - Re-attach the Ismatec pump white applicators and snap down the applicators onto the Ismatec pump rotating device.

NOTE: Make sure to place the applicators back in the same order.

  - Switch the Ismatec pump on to the forward direction.

  - Place a piece of tape onto the Ismatec pump motor labeled with what day the tubing needs to be moved (one week later) and what day the tubing needs to be changed (2 weeks later).

Moving and Changing the MasterFlex Tubing:

NOTE: When moving or changing the MasterFlex tubing, make sure the Biomek is not running.

      - At the beginning of every week, release the MasterFlex tubing clamps, move the MasterFlex tubing over one notch, and reclamp the MasterFlex tubing clamps.  Mark the date it was moved on the tubing.

      - There are four spaces that the MasterFlex tubing can be moved. When all spaces have been filled (approx. 1 month), enter a work request to have the MasterFlex tubing changed by the Instrumentation group.

# Troubleshooting

PROBLEM 1:  Small number of Tips bent during run without error

If the tips are bent without error, the operator will have to snap the method, raise the pod (if necessary), (see b. Moving Pod Via Manual Control), replace bad tips with good ones and restart the snapped method.

1. Wait until there is no liquid in the tips.

2. Snap the method by pressing the green snap button next to the stop button.

NOTE: The method will then be saved from this point (i.e. The Biomek will resume from the point that it was interrupted).

        - Determine if any of the plates being processed have been contaminated due to the bent tips (reagent on the top of the plate or plates being in wrong locations). If so, fail these plates in the database system. Scan them into the Enter Fails or Comments form in the RCA DB, select a failure mode, add an appropriate comment, and then click on submit.

          - Mark the plates that need to be failed with a sharpie so they will be removed from the production line later. Do not remove them from the deck or stackers at this time.

NOTE: Remember to select the Do Not Submit checkbox for any failed plates when the barcode validation window opens.

          - Print out a copy of these failed plates and give a copy to Sequencing Chemistry so that the plates can be resequenced. Write on the print out whether the plates are MB or ABI.

        - Moving the Pods via manual control during a run:

          - After the method is snapped, go to TOOLS ( select manual control.

NOTE: You will not be able to manually move the pods in the middle of a run.  You must SNAP the method first.

          - Select Advanced controls.

          - Select Pod 1 (if tip problem)/ Pod 2 for gripper errors.

        - To move the pod up, select Move Z-Max (you may also click or select the up arrow button to move the pod upwards).

NOTE: Z-Axis moves the pods up and down, Y-Axis moved the pods back to front, and X-Axis moves the pods left to right.

        - Check to see if tips are bent.

        - If necessary, replace any bent tips and/or reposition any plates.

          - Locking pliers can be used to remove bent tips.

          - Replace the bent tips with new ones from the SPARE TIPS box.

        - Close the Advanced Control (Close) and Manual Control (Exit) windows.

        - Select the RUN button (green play button) and the method will resume.

PROBLEM 2:  LARGE NUMBER OF TIPS BENT DURING RUN (WITH ERROR – LIQUID IN TIPS)

If the Biomek crashes while liquid is loaded in the tips, the operator must abort the method, unload the bent tips and start the run over from the beginning of the current step.

    - Abort the run by selecting the Abort button on the error window.

    - Unload the bent tips: Go to File ( New Method ( Run (or File ( Open ( Unload Tips). This will unload the bent tips into the tip box.

    - Remove the old tip box from the deck and dispose of in autoclave bag. Replace with a new tip box (Label with your initials and date). Perform a tip QC now or some time that same day by following the QC Protocol for the Biomek FX Liquid Handler.

    - Instrument crash during BET addition: The operator should put dummy plates into Stacker A to replace those that already have BET added to them. Run the normal method, but put the original plates on the deck after BET is added to the dummy plates.

    - Instrument crash during EtOH Wash: The operator should put dummy plates into Stacker A to replace those that have already been washed. Run the program ETOH_Wash_and _Transfer_SW_091604.bmt. After the dummy plates have been washed, put the original plates on the deck.

    - Instrument crash during Final Transfer: The operator should put dummy plates into Stacker A and C to replace those that have already been transferred.  Run the program Resuspension and Transfer_SW_091604.bmt

  - If neither the source plate nor the destination plates have been scanned, the operator must put the original plate on the deck after the dummy plate has been transferred, but before it is scanned.

  - If the source plate has been scanned but not the destination, the operator must put both the source and destination plates on the deck after the dummy plate transfers.

  - If the source and destination plates have been scanned, the operator can remove these plates from the set to be run with the Final Resuspension and Transfer program.

PROBLEM 3: LARGE NUMBER OF TIPS BENT DURING THE RUN (WITH ERROR – NO LIQUID IN TIPS)

If the Biomek crashes with no liquid in the tips, the operator can change the tips and continue from where the instrument left off.

      - Snap the current method by pressing the green snap button next to the stop button.

      - Unload the bent tips: Go to File ( Open ( unload_tips_midrun_SW_051004.bmt.

      - Run this method. This will unload the tips, but not clear the world file of all plate locations.

      - Remove the old tip box from the deck and dispose of in autoclave bag. Replace with a new tip box (Label with your initials and date). Perform a tip QC now or some time that same day by following the QC Protocol for the Biomek FX Liquid Handler.

      - Reload the new tips: Go to File ( Open ( reload_tips_midrun_SW_051004.bmt.

      - Run this method. This will load and wash tips without clearing the world file of plate locations.

      - Go to File ( Load continuation. This is your method continuation that was created by the snap method.

      - Press RUN.

PROBLEM 4: PLATE CARRIED TO WASH STATION

            - If the pod carries the plate to the wash station, snap the method.

            - Use the manual controls (Refer to Problem 1) to move the pod up allowing you to move the pod up from the wash trough.

            - Verify that the plate has not been contaminated (look for uneven well volumes and possibly splashing on the top of the plate.)  If you think this plate was damaged or contaminated, mark it with a sharpie (so that you can fail it at the end of the run and place it in the appropriate position upon the deck.)  If needed, replace any bent tips that could be catching on the plate. Then resume the method.

NOTE: Remember to select the Do Not Submit checkbox for any failed plates when the barcode validation window opens.

PROBLEM 5: GRIPPER ERROR

      - Sometimes a gripper error occurs saying that the “gripper failed to retract” or that the “gripper failed to extend.”  When this happens, hit retry on error prompt window. You may need to do this several times.

      - If this does not work, you must abort the run.

      - Go to Manual Controls ( select Advanced Manual Control ( Select POD 2 (click on Retract Gripper ( reload the plates and restart the run.

      - Put a work request into the system to inform Instrumentation of the problem.

PROBLEM 6: LIGHT CURTAIN ERROR

This type of error occurs when the light curtain is broken. When this happens, the run will stop and there will be a RED error window prompt. Click OK. This will resume the program from where it was interrupted.

PROBLEM 7: UNLOCKING CLAMPS ON MAGNETS AND PLATE FLATTENERS

During the run, if a plate needs to be adjusted while on the magnet or plate flattener, you can manually unlock clamps and fix the plate problem.

  - Press the pause button on the timer and either press the pause button on the Biomek software or break the light curtain to pause the run.

  - To open the clamps, there is a switch located on the deck that says, on/off.  Turning the switch to the off position will open the clamps.

  - You can then remove or adjust plates.

  - Turn switch to the on position.

  - Resume the run by pressing the run icon or clicking OK on the light curtain error.

  - Unpause the timer so that future errors will be sent to the pager.

PROBLEM 8: STACKER ERRORS

  - Press the pause button in the Biomek software (or break the light curtain) and pause the timer.

  - Press the Remote Local button on the Stacker Carousel Pendant in order for the display to read “Keypad Control Only.”

  - Move the appropriate hotel to the shuttle by rotating the carousel either left or right, using the Stacker Carousel Pendant.

  - Restack plates by moving the shuttle in (press Shuttle In on the keypad) and then moving the stack up (press Stack Up on the keypad).

  - Ensure that all source plates are in order with their corresponding destination plates in Stacker C.

- Plates that end up in Stacker B will be in the reverse order as those in Stacker C.

- Plates that end up in Stacker A will be in the same order as those in Stacker C.

  - Resume the run and unpause the timer.

  - If plates are being stacked crooked into the stacker notify instrumentation.  The stacker may need to be leveled or the shuttle may be misaligned.

NOTE:  Carefully examine plates that were stacked crooked to verify if the sample was splashed causing cross-contamination.  If this occurs, fail the plate in venonat and request the Chemistry resequence the affected plates.

PROBLEM 9: SUBMITTING PLATES TO THE DATABASE

Sometimes a Biomek will have problems connecting to the database.

      - First, try resubmitting the barcodes by clicking on Connect to DB.  Submission may take awhile.  There will be a prompt telling you that it is OK to submit the barcodes. Continue with submission as normal.

      - If this does not work, you can manually submit the barcodes by following the troubleshooting steps below:

        - Press the ALT + F4 button to close the open barcode window.

        - Go to the desktop open up the folder Shortcut to Logs.

        - Look for the file with the name barcodes today’s date and the time of the run in question. It should look something like this “barcodes04-08-2004 15.45.21b.log!”

        - Copy and paste the file or files in question onto a disk.  You can do this by inserting a 3 ½ in floppy into the A drive open my computer and then drag from the shortcut folder into the A: drive.

        - Take the disk to another Biomek computer (select one that the internet connection is working.)  Put disk into the A: drive open by going to my computer.

        - Drag the barcode file(s) in question into the Shortcut to Logs folder.

        - Right click on the barcode files in question select Rename.

        - Click after the “!”, then remove the “!” from the file name and press ENTER. You will receive a prompt asking if you are sure you want to change the file name select YES.

NOTE: At this point the file name should look like: barcodes04-08-2004 15.45.21b.log.

        - In the shortcut to logs folder click on the “up one level button” (this button is a folder with an upward arrow). After selecting this you should see a folder called GI Custom Programs.

        - Run the barcode-parsing program: Double click on the file named “barcode_upload_biomek.exe.”  (A shortcut to this program may be on the desktop of the computers operating the Biomeks). A processed barcode file window will then pop up showing the association between the source plate and the ABI barcode verify that everything is correct and make necessary comments then select proceed and save as normal.  This file will then be uploaded to the database and is now considered processed.

        - Return to the original Biomek, open the Shortcut to Logs folder.

        - Move the files that you just processed into the Processed_logs folder.

      - If the Biomek has scanned a portion of the barcodes in a run and the method is then snapped or restarted the Biomek software will generate multiple log files. This will cause the barcode-parsing software to present multiple windows to the user and often times the barcodes will be out of order. Follow the steps below to correct this problem:

        - Do not submit the barcodes to the database.

        - Press “Alt+F4” to close the barcode verification screen. If another barcode-parsing window opens, close it in the same manner.

        - On the desktop, open the Shortcut to Logs folder.

        - Find the file named “barcodes”; today’s date and the time of the run in question it should look something like this “barcodes04-08-2004 15.45.21b.log!”

NOTE: This file should contain either the first or second portion of barcodes in the Biomek run.

        - Open the file; verify that the barcodes in this file are the ones that appeared on the barcode verification screen.

NOTE 1: In these logs the barcode parser program scans the first 4 source plates then the first 4 destination plates. So the first source plate scanned will then be paired with the first destination plate scanned.

NOTE 2: You will most likely see that there are not enough entries to match the source and the destination to all of the plates.  You will then need to copy these barcodes by highlighting them and pressing “Ctrl+C” (close this file by going to File and selecting Close).

        - You will then need to look for a second barcode log file in the same folder Short cuts to logs.  It should have the same date as the previous file and approximately the same time.  It will also end in an exclamation point.  Open the file and verify that these are the remaining plates in the run. This second file might look like: “barcodes04-08-2004 16.23.33.log!”

        - Paste the first set of barcodes into this file by pressing “Ctrl+V”.  Keep the one space line at the top of the document.  You should now have a file consisting of 4 scanned source plates, then 4 scanned destination (ABI) plates, then another 4 source plates, and another 4 destination plates and finally 2 source and 2 destination. Verify that the scanning order is correct by checking the time that each scan took place.

        - Save this by clicking File, then Save. Then exit out of the file screen. By going to file and Close.

        - The file now needs to be renamed.  In the “shortcut to logs” folder find the file in that has been modified.  Right click on the title with the Mouse select Rename.  The file name will then become highlighted and the cursor will be blinking.  Remove the exclamation point from the file name.

        - Move the un-modified file (barcodes04-08-2004 15.45.21b.log! in this case) to the “Processed_Logs” folder.

        - Go to the desktop and open the shortcut to barcode upload program.  The barcodes from the batch in question should appear in the correct order and appropriately matched.  Verify then Proceed as normal by selecting Proceed, entering the batch info and Save.

PROBLEM 10: Home all axes

Each axis (X, Y, Z & D) must be homed when the Biomek power is turned off and on.

            - Go to the Tools menu ( select Manual Control ( select Home all Axes.

            - A window will open asking if there are any tips on POD1. If so, they will be shucked. Unload tips back into tip box by running the unload_tips_midrun method before continuing.

            - Click OK.  Stop button will become available.

            - Click Exit when machine is stationary.

# SOP Approval

| DEPARTMENT | APPROVED BY | DATE |
| --- | --- | --- |
| Lab Supervisor |  |  |
| Research & Development |  |  |
| Instrumentation |  |  |
| QC |  |  |
| Purchasing |  |  |
| EH & S |  |  |
| Informatics |  |  |
| Seq Assessment & Analysis |  |  |
| Dept Head of Prod Seq |  |  |

# Appendix

AUDIT TRACKING

5/27/04 – Damon Tighe and Helene Feil performed an audit between 8:30AM – 12:30PM. The operator was Catherine Gordon.

PROCEDURAL CHANGES

7/26/04 – Initializing Biomek Carousal and Homing All Axes.

8/17/04 – No longer need to spin down ABI destination plates after Standard has been added. We now visually check every plate after standard has been added and seal the plates. We only spin down if there are bubbles or well volumes don’t look consistent.

12/12/04 – Start taking tip QC’s daily at 10pm (the start of the 3rd shift), or immediately after the tips have been changed(PCN).

1/14/05-.  We quit dispensing standard in the A1 well (12/20/04) of the ABI plates.  Mark unpaired plates with a blue “S” (12/22/04) and place these plates in a bin marked “unpaired” for the appropriate platform.

Updated program names that we are currently using on the biomek.

02/16/05- Begin taking weekly tip QC’s instead of daily.  QC will occur at the end of 2nd shift or the beginning of 3rd shift at 10pm every Wednesday. (PCN)

5/10/05 – Changes to Ismatec tubing shift and change.  Too much lubrication was being used and causing premature tears in tubing.  SOP changed so that less lubricant is used & pump rollers are no longer sprayed.

ADDENDUM

Cybio Print & Apply Addendum to Post sequencing BET cleanup for ABI plates:

- Follow steps 1.0 thru 2.2a on the Bet Post sequencing BET Protocol.  After scanning the source plates into the database, load them into stacker A of the biomek.  Load the same number of pre-barcoded ABI plates into stacker C of the Biomek and start the current program.  If unclear on how to start the biomek refer to Section 3.0 Biomek FX Setup of the Protocol.

- After the biomek has finished the run be sure to save and upload the barcodes in the Barcode Submission program displayed at the end of each run on the biomek.  Discard the source plates from stacker B, and remove the ABI destination plates.  Seal the plates with the Velocity 11 plate sealer (with the pierceable seal.)

NOTE:  Do not remove the source or destination plates from the Biomek FX before they have been uploaded to the database by the Barcode Submission program.

- After sealing these plates they are now ready to be labeled using the Cybio print and apply.  Stack up to 40 ABI plates in the Cybio stacker labeled Load 1.  If there are more than 40 plates, load up to 40 additional plates in the stacker labeled Load 2.  The blank side of the plate should be facing out of the stacker; this is where the barcode will be applied (The ABI barcode will be loaded facing the back of the stacker.)

Note:  You cannot keep loading or load more than 40 plates into each stacker. If more than 40 plates are placed into either load 1 or load 2 stackers the Unload stackers will overflow.  Possibly allowing plates to fall.

  - Open the Cybio program by doubling clicking on the Icon on the desktop labeled Cybio Control.

  - Then go to FILE(Local Disk C:(Cybio Software(2005.01.24 BG Two Stack short Labels.  (Below is the opening screen of the program)

  - To start the Cybio You can either go to EXECUTE(Program Execute (F5) or click on the icon with an arrow going down with lines across.  (It is the 15th icon from the right side)

  - After selecting this icon the following message will be displayed as the cybio initializes.  The shuttle will move first to the right and then to the left of the track.

Wait for a moment.  After the Cybio has finished initializing the following prompt will be displayed when the shuttle is in the far left position.

Verify that the shuttle is empty then select Ok.

  - The Cybio will now begin removing the plates from the stacker labeled Load 1, scanning the plate, affixing the barcode, and restacking the plates into the Unload stackers.  After the plates are stacked into the unload stacker with the label affixed to the side of the ABI plate, the operator can then unload them and mark the plates with the appropriate marking according to the Color code chart posted through out the lab and place these plates in the ABI backlog refrigerator in Room 141.

Note:  If the Cybio fails to label any of the plates try replacing the plate in the load stacker for it to label the plate again.  If this fails verify that all of the barcodes have been submitted to the Database on the biomeks.  Sometimes there is a lag time after submitting the barcodes to the database and the Cybio accessing this information.  If this occurs hold the plates in the TEG backlog and wait for 15 mins to an hour and try again.  If it still fails enter a work request.

Note:  If the Cybio crashes while running and the operator has to restart the program, the first stacker Load 1 must be refilled with 40 plates if you want the program to go to the second stacker (Load 2).

  - When the program is finished a message will appear stating that program has been executed without errors.

Note:  If only 40 plates are loaded into Load 1 you will receive the following message. This only means that it was unable to get a plate from Load 2.