Pierre Auger Project Electronics

Front End (FEv2.x) Board

Kwame Bowie

September 1, 2000

##### Revised: 6/13/2001

- Introduction

The Front End (FE) board is the digitization and triggering board for the Surface Detector Electronics system. The Surface Detector Electronics system consists of a CPU board (CPU), a Slow Controls board (SC), a Time Tagging board (TG), an Ethernet board (ETH), a GPS receiver (GPS), and the Front End (FE) board. The FE board is directly linked via the Local Bus to the all of the subsections. The block diagram Fig. 1 shows the basic layout.

Figure 1:  Surface Detector Electronics Interconnections

- Operation

The FE version 2 board consists of the FEv2.x board developed at Fermilab and the Surface Detector Trigger granddaughter-board, developed by Michigan Technological University (MTU). Together these two boards function as the digitization board, the data storage board, and the trigger generation board for the Surface Detector. The main functions of the board are listed below:

- FE Board must take 6 channels of analog data from photo multiplier tubes and convert it to digital data.

- FE Board must store this digital data in memory deep enough to hold 25.6 (s of data (6.4 (s pre-trigger and 19.2 (s post-trigger).

- FE Board must transmit this data to the CPU based upon trigger conditions.

The following are the specifics of the board operation based upon the Surface Detector and Front End board specification.

- The analog to digital data conversion (ADC) is to be done at a sampling rate of 40MHz and with a precision of 10 bits per channel.

- There will be six (6) channels of analog inputs from the photo multiplier tubes (PMTs). The digitized data from these channels shall be used generate the trigger(s).

- The digitized data shall be stored onboard and read out at known intervals after trigger generation.

- The stored data shall be comprised of ¼ pre-trigger data and ¾ post-trigger data.

- There will be two (2) types of triggers: fast and slow based upon amplitude, slew rate, and time over threshold measurements.

- The memory will be double buffered so that an event may be accepted while data is being read out.

- Data readout will comply with the PPC403 CPU specification for DMA transfer in Fly-By-Burst Mode.

- The FE board must be air-cooled  with no fan.

- The board must operate at temperatures between –20(C and +60(C.

To ensure that prototypes were available for initial testing, two (2) versions of the FE board have been produced: Version1 and Version 2. All Front End Boards receive the following nomenclature: FevM.N  where M is the version number and N is the design iteration. This nomenclature is used when referring to the different versions of the FE board from this point on.

### 2.1	Version 1.x

The first version, Version 1.x, of the FE board consists of the FE motherboard and an analog daughter card. FEv1.x supports the use of a digital daughter card by providing a connector and signal paths for a digital daughter card. The digital daughter card was to be developed in parallel at a different location, but that option was not pursued. The digital daughter card was to produce a digital trigger signal, which could be logically OR’ed with two other triggers (Analog and External) to control data flow into the buffers.

To provide maximum flexibility in available configurations, the FEv1.x  board design was split into three (3) separate boards: the FE motherboard, the Analog Daughter Card, and the Digital Daughter Card. By separating the FE Board design into distinct parts, several engineers could work on various portions of the board in parallel. This has two desirable effects: first by making distinct parts on separate boards, a delay or design change on one board will have minimal effect on the others as the design processes become independent. Second, by separating the boards, performance analysis may be done with different FE configurations. It becomes possible to have several interchangeable analog front ends or digital daughter boards and evaluate the performance of each in the actual system to get a true performance metric.

### 2.2	Version 2.x

Version 2.x of the Front End board will have the same physical dimensions and stacking- connector positioning, but will eliminate the need for first-in-first-out memory chips (FIFOs) or a PLD. In place of the FIFOs and PLD will be a PGA socket for an ASIC designed to integrate all of the digital logic into a single chip. As a result, the digital daughter card connector along with all digital logic will be gone from Version 2.x. Also, the ASIC will handle a more advanced triggering scheme, which will be based solely upon the digitized data so no analog trigger will be present.

- Input signals

Input signals to the FE board originate from three (3) PMTs. Each PMT outputs two (2) analog signals. The three PMTs are set up to have very different output gains. To simplify nomenclature we simply call the two PMTs high gain and low gain. The PMT outputs appear at the FE board as six (6) single-ended, negative-edged signals on 50-Ohm cables from the PMT bases. The cables are terminated at high frequency SMA sockets. The high gain channels have an input voltage range of 0 to –0.3V and the low gain signals have a voltage span from 0 to –1.5V.

- 6 single-ended signals

- 50 Ohm cable impedance from bases

- High gain range (0 to –0.3V)

- Low gain range (0 to –1.5V)

The FE board has the ability to handle the 6 input signals in either single ended form or in a pseudo-differential form via jumper selection. The pseudo-differential method provides the ideal solution in low-noise environments as it eliminates possible ground offsets between the FE board and the PMT bases. To connect with the coax cable, we have selected standard, right-angle, PCB-mount, SMA connectors AMP part #: 221790-1.

4.      Shaping/Amplification/Filtering

4.1 Overvoltage protection

Before any filtering, shaping, or amplification is done, the amplifiers must be protected from over-voltage conditions. This actually serves 2 purposes: first, it prevents large voltage spikes from exceeding the maximum op-amp voltage ratings. Reliability will certainly be a major concern in this project as the stations are predicted to be serviceable only at very long intervals. Replacing op-amps after any large event or noise spike is not an option. The second reason for including surge protection on the analog input signals is to lower the power consumption within the analog portion of the FE board. To achieve the over-voltage protection goal, we use series-connected dual pin diodes. For the high gain channels we only use 2 per channel since we can clamp after one diode drop since the signal is limited to 300mV. However, for the low gain channels, the signal range is much larger (1.5 V) so we need at least 3 diode drops in the signal direction before we clamp the signal. So for the low gain channels, each channel receives 3 dual pin diodes. The pin diodes that were initially selected for clamping were the Agilent HSMP3822 series connected dual pin diodes. However, due to availability and pricing, a replacement had to be selected. The replacement part that was chosen was part number MA4P274ST from Richardson Electronics. The only important parameters when selecting the replacement parts were price, effective series resistance, and effective capacitance. All of these parameters were deemed acceptable for the Richardson replacement part.  The input circuits are shown below in Figures 4.1 and 4.2.

Figure 4.1: High gain channel input circuit

Figure 4.2: Low gain channel input circuit

4.2 Amplification

The very next step after surge protecting the inputs is to amplify the inputs. Amplification is necessary to enable the analog inputs to match the ADC input range. The ADC device that we have selected has a 0 to 1V input span, so the inputs must be amplified accordingly. The term amplification is used as a generalization of the operations performed on the inputs. The operations are actually polarity inversions, amplifications, and/or attenuations. All three of these terms may be generalized and considered amplifications by a gain factor. The sign and absolute value of this gain factor determines whether the amplification is a polarity inversion, an amplification, or an attenuation. In our case, all high gain channels have a gain factor of -10/3 and all low gain channels have a gain factor of -2/3. Thus the high gain channels are actually inverted then amplified, whereas the low gain channel is inverted then attenuated. The Analog Devices AD8011 Current Feedback Op-Amp was selected because it provides excellent gain vs. bandwidth characteristics and is a very low power device, dissipating nominally 1mA of current. The 1st op-amp in each channel is responsible for amplification. Thus, by examining the input circuits in figures 4.1 and 4.2, we can see that the gain is roughly -2.7 for the high gain channels and  -0.46 for the low gain channels.

4.3 Filtering

After the amplification stage, the inputs are filtered. This filter is a simple, 2-pole, anti-aliasing filter. The signals that we intend to recover span a wide range of frequencies and have spectral components that also span many frequency ranges. The desired signal range is from DC to the maximum frequency resolvable by the ADCs. However, the station resides in a relatively high-noise environment and must reject the broadband RF noise. An anti-aliasing filter provides good rejection of unwanted noise signals above the Nyquist frequency. The sampling frequency of the ADC is 40 MHz so a simple 2-pole active filter with cutoff frequency at 20 MHz works very well. The circuit in figure 5.3 shows the 2-pole active filter used for both the high gain and the low gain channels. The filter itself has a gain of 1.44, and this brings the input-to-output gain of the amplifier and filter sections to 3.86 for the high gain channels and 0.67 for the low gain channels.

Figure 4.3: Filtering circuit

- Analog to Digital Conversion

The signal lines are single-ended analog waveforms which represent the filtered, shaped, and amplified PMT outputs. These six signals are the inputs for the analog-to-digital converters (ADCs). Since the signals are single ended, they are all referenced to the board’s analog ground. By all signals referencing the same analog ground, it becomes possible for small return currents to appear at the inputs. This was an area of intense scrutiny as there is no consistent documentation on the best way to handle the ground connections for Analog Devices ADCs. The final circuit that was agreed upon has the ADC’s ground connections isolated from each other, and only returning to their own ground planes in order to prevent digital ground currents from returning to the analog input. Figure 5 below shows the ADC circuit that is used in the final FEv2.2 board. Also visible in figure 5 is the pedestal selection circuit. The pedestal selection circuit is a simple voltage divider between VREF an AINN. The circuit simply shifts the midpoint of the ADC span downwards. The desired pedestal was 5% of maximum range, which corresponds to roughly 50 counts. Theoretically a 9/10 voltage divider should provide this correct range. However, in reality there is already a small offset of 20 to 40 counts. This is probably a result of the low pass filters averaging high frequency noise. If this noise is not symmetric about analog ground (which it probably isn’t) then the averaging leaves a DC component of the signal. To counteract this initial offset, the voltage divider had to be set to shift the midpoint less than theoretical predictions would suggest. As a result, the voltage divider for the FEv2.2 was selected to be a ratio of roughly 96/100. The actual values were R23  = 562 Ohms and R51  = 13 KOhms.

Figure 5:   Analog-to-Digital Converter circuit

The sampling device that was selected was the Analog Devices AD9203 flash ADC chip. This chip was selected due to its very low power consumption with a high conversion rate and decent precision. The AD9203 utilizes both edges of the SYSCLK signal. The ADC inputs are clocked into the flash converter’s sample and hold amplifiers on every trailing edge of the SYSCLK signal. The digital outputs are latched onto the data bus at each rising SYSCLK edge. To enable such a high sampling rate at decent precision and low power, the AD9203 incorporates a pipelined flash architecture. The intrinsic delay of each input sample is 7 rising SYSCLK edges.

- Digital Interface

The FEv2.x boards must interface with the PowerPC (PPC) 403GCX local station CPU in order to transfer the accumulated data. The transfers occur once a data buffer is filled and ready to transmit the ADC data. The PPC CPU is signaled that the FE is ready to send data by the FE asserting the EVTCLK(F/S) signal on the bus. The PPC processor then determines when it would like to conduct a data transfer. To ease the load on the processor, the DMA fly-by-burst memory transfer mode was selected. By deciding to use the DMA transfer mode, the ALU and registers of the CPU can still execute instructions uninterrupted, but cannot access the data bus.  The fly-by-burst mode is the fastest and simplest of all of the DMA transfer modes, which is very important as the CPU data bus is locked out at this time. The transfer uses 1 wait state for the memory to enable adequate time for the FE boards to place the necessary data on the bus. With one wait state, the data is put onto the bus at a 20MHz rate. Each buffer consists of 2 FIFOs therefore reading both of these 1024 word buffers will take roughly 102.4 us. However, after many CPU write cycles, the CPU intervenes to allow the DRAM to refresh and row switch. So a full data transfer ends up taking a little longer than 125us.  Specifics of the DMA fly-by-burst protocol can be obtained from the PPC403 User’s Manual. One very important note about the digital interface to the CPU is that the PPC devices use IBM’s bit ordering. This caused a slight problem as in most cases D31 is assumed to be the highest order bit, whereas with the PPC CPU, bit D31 is the lowest order bit. The bit swapping is done on the circuit board layout as software bit swapping is computationally expensive and the PGA socket’s pins were already defined. Also, current versions of the trigger boards differ from the FE board spec. The FE spec. allocates 1024 words of data for each event, but the trigger board implements 768 words due to the memory limitations of the current PLD selected for the trigger board. However, the pre-trigger/post-trigger relationship remains the same. This is not a problem but should be noted to ensure that the CPU knows when valid data ends.

- History

This section is dedicated to explaining the differences between the FE version 2 board, both cosmetic and otherwise. It is important to be able to visually determine which revision of the FEv2.x board one is examining. Without knowing which version of the FEv2.x board is being analyzed, it is impossible to find the correct schematics to troubleshoot it.  Although the FEv2.2 board achieves all of the design specs., the other iterations are more than capable for development purposes. For that reason, it is perceived that all versions of the FE board may remain within limited use at different locations.

7.1	FEv2.0

The FEv2.0 board although functional and useful for development was riddled with problems. The first, and possibly the most serious was the fact that the entire data bus was bit-swapped. This problem made the FEv2.0 unacceptable for production even if no other problems were found simply because bit-swapping code had to be developed. Although this code does not seem to be a considerable problem, as the code is quite easy to write, it puts a tremendous load on the CPU. The bit-swapping was a true bit swapping and not a simple endian-ness switch. For that reason, every bit of every word had to be moved. For 768 data words, with 32 bits each, this gives a CPU processing overhead of at least 24568 operations for each DMA transfer. The CPU operates at only twice the DMA transfer speed so this processing overhead was much too high. As a result, it was immediately apparent that there would be a revision, FEv2.1. Another problem that presented itself was the fact that the trigger board could not access memory fast enough to send data to the data bus without dropping the first word. As a result, some signals were dropped in favor of DRAM control signals, so that the trigger board would have enough time to prepare data for the DMA transfer. These changes were implemented in kludges for the FEv2.0 board but were added to the FEv2.1 revision design. The final problems with the FEv2.0 board were related to the analog amplification and shaping section. The first of the analog-related problems was noise. The FEv2.0 was significantly noisier than the specification tolerated. True system-level noise estimates could not be done at Fermilab due to a lack of hardware, but with a stable power supply and data read from the actual socket, it was determined that rms errors due to noise were less than 2 LSBs. The design spec demanded rms noise within the actual system to be limited to less than 1 LSB. It was found that the major problems with the FEv2.0 with respect to noise were twofold. First, the analog signal lines were too far from the ADCs that were measuring the analog signals. Secondly, the grounds were handled poorly. The analog signal traces traveled over the digital ground plane, which exhibits quite a bit of high frequency ground bounce. Even worse, a high Q inductor was placed within the analog ground return path effectively peaking any small ground perturbations that do occur. The true nature of these grounding problems was not fully understood until the FEv2.1 board was tested. The last major problem with the FEv2.0 board was the unacceptable amount of cross talk. One of the major problems with the original FEv1.0 board was that it provided only one analog ground reference point and return path for all channels. The ideal arrangement of the analog daughter board should’ve been: analog signal – analog ground – analog signal –analog ground, etc. However, the connector was actually arranged: analog signal – digital ground – analog signal – digital ground. This did not allow each analog channel to have its own reference point and ground return path. As a result, a large signal on a given channel could cause small analog ground perturbation of a few millivolts. While this perturbation would not be seen by the driven channel as it is in phase and just provides a rejected common-mode component, the un-driven channels all reference the same ground point and as a result, see the ground ripple as a signal of a few millivolts which is significant at our scale. There was no easy way to kludge this fix, so solutions to this problem were untested until the next design iteration.

The FEv2.0 boards are very easy to distinguish from the other version 2 FE boards. These are the only FEv2.x boards with an analog daughter card connector and surface mount pads for comparators and a PLD.

The silk-screened text says that this board is Rev. C, designed October 6, 2000.

The schematic is listed as FEv2_0.pdf.

7.2	FEv2.1

The FEv2.1 was the first redesign of the FE version 2 board and remedied a few of the more definite errors. The databus bit swapping and DRAM control signals were correctly fixed in the FEv2.1 board. Also, the cross talk problem was remedied by using a single analog ground plane for all signals referencing analog ground. These were problems that had the sources already identified and solutions already developed. As a result, these problems were fixed without any problems in the FEv2.1 board. The only problem that did not have its sources clearly identified and its solutions completely developed was the noise problem. The FEv2.0 board while in the actual system was said to exhibit 5 to 6 LSBs of rms noise. This was unacceptable and needed to be fixed for the engineering array (EA). The sources were not completely identified, so the solutions were only based on hypotheses and the FEv2.1 was designed to be a test bed for possible fixes to noise problems. This led to the use of many redundant parts and jumpers to allow for a maximum number of different board configurations. Prior to finishing assembly of FEv2.1, Fermilab received a local station CPU board. This was essential because even with the FEv2.0, the noise from the FE board run without the station and PLD was almost acceptable for the EA. When the FEv2.1 was tested from a stable supply without the trigger board, noise was less than 1 LSB so testing of noise performance would absolutely require the real system setup to get an adequate noise estimate. Also, it was determined that to ensure a channel was working, it would be beneficial to include a pedestal. Changes to the pedestal value should be implemented by changing component values. It was debated whether it would be better to set all channel’s pedestals at a single point or whether it would be better to be able to individually trim each channel’s pedestal. A compromise was reached by implementing both solutions and selecting the best one via jumpers. A comprehensive list of changes between FEv2.0 and FEv2.1 is as follows:

- fixed data bus bit ordering

- removed DMADXFER, TC0, and, TC1 were replaced with CAS3, LA23, and DRAMWE

- referenced all analog related parts to AGND on a single analog ground plane (including both ADC grounds)

- added voltage divider for pedestal setting

- added reference and amplifier for alternative pedestal setting

- added ground jumper to select chassis or DGND for AGND return path

- changed amplifier supply rails from +/- 5V to +/- 3.3V to keep outputs within ADC input range even if pin diodes fail. This change eliminated DC-DC converters

- added pull-up network for EXT_TRIGGER as it is active-low (negative logic)

- removed remnants of analog daughter board (24-pin connector, digital GND plane on lower half of board).

These are all of the changes that were made in the FEv2.1 board design. After several days of testing different configurations with the real system (local station CPU and trigger board), the best configuration for this board was determined that performed almost well enough for the EA (but not quite). As a result, another design had to be done. The next design (FEv2.2) would incorporate all of the design changes that were needed from FEv2.0 to FEv2.1 and take advantage of all that was learned while configuring the FEv2.1 board to perform at its best performance point. While configuring, the FEv2.1 board, it was determined that the major noise sources were radiated and conducted digital noise from the station and the trigger board. Without both the trigger board and station, the FEv2.1 had rms errors of 0.5 to 0.9 LSBs. With the station powering the FEv2.1 board, the rms error jumped to between 1.0 and 1.2 LSB rms. Finally, when the trigger board was added, the errors jumped to 1.3 to 1.9 LSBs rms error. Taking this knowledge in, it was decided that it might be a good idea to re-separate the ADC ground connections and more importantly to add buffers on the digital outputs. These changes should handle most of the conducted noise, and would appear in the designs for the FEv2.2. To conquer the radiated noise problems, the main technique was spacing and shielding within the layout. By moving the analog outputs closer to the ADCs, less radiation is picked up between these sections. Also, shielding the data bus and moving the amplifier and ADC sections further from the databus, help reduce radiated RF noise and pick-up. Finally, for FEv2.2 it was decided that the amplifier and reference were more trouble than they were worth as that circuit performed worse than the voltage divider at the ADCs (probably due to trace length).

The distinguishing characteristic of FEv2.1 is the large number of throughhole-jumpers at the ADCs.

Also, the digits 02-01 appears in the upper left hand corner of the board.

The board is labeled on the silkscreen info section as 5 February 2001 Rev. 2_1

The schematic file for this version is FEv2_1.pdf.

7.3 FEv2.2

FEv2.2 is the final version of the FE board, and shall be produced in large quantities (50 boards) for the EA. The FEv2.2 finally solved the noise problems that could not be solved with FEv2.1. Most of the problems were diagnosed after several weeks of analyzing FEv2.1. However, the solutions could only be proposed and not tested because of FEv2.1 layout constraints. FEv2.2a is the prototype for FEv2.2 and is a test board that incorporates several jumper options. The jumperable options on the FEv2.2a board are as follows:

- DGND to DGND or AGND at each ADC

- DGND and AGND tied at each ADC

- Pseudo-differential input or single-ended input at SMA connector

- AGND return through DGND or CHASSIS

- Buffer or transceiver

These are the only jumperable options for the FEv2.2a, and the FEv2.2 board eliminates options 1 and 2, as it is the actual board to enter the small production run for the EA. These jumpers were eliminated to enable the FEv2.2 board to be assembled more easily as the SMT jumpers do not have parts associated with them.

The design changes from the FEv2.1 to the FEv2.2 board were very slight; mainly noise suppression and correcting the surge protection on the low-gain channels. The only design changes visible from the schematics were the removal of the reference-amplifier circuit, the separation of the grounds at the ADCs, the insertion of buffers at ADC outputs, and the clock driver using all fanouts. The design changes were deemed redundant and more power hungry, but were more in line with good design practice than the previous FEv2.x versions. The main changes to the FEv2.2 design were layout related changes. The initial purpose of FEv2.2 was to shift the layout so that shielding would be simpler, as it was known that a good portion of the rms noise was radiated digital noise. The remainder of the changes were made to ensure that good design practices were followed. The layout changes focused around three key goals: keep the ADCs close to the amplifiers, move the amplifiers and ADCs away from the databus, and shield extensively. To accomplish these goals, extensive shuffling of the layout had to be done. First, the layout was migrated to 8 layers to enable room for traces while shifting and shuffling parts. Second, ½” thick rectangular perimeters surrounded the ADC and amplifier sections to enable separate shield cans to be placed over each of these sections. Then both the ADCs and amplifiers were shifted as far right and close together as possible to ensure that the databus is as far away from these sections as possible. The ADCs were then shifted down substantially so that they were very close to the amplifiers. Then layer stack was setup such that the clock and low order data bits were not on the same layer as the analog signals to prevent high frequency radiated switching noise.  Finally, the data bus lines were shielded at the connector by a ground plane on every other layer starting with the top and bottom. All of these changes brought the rms noise down to 0.3 to 0.5 LSBs of noise. This meets the design spec. for the production run next year. As a result, minimal design changes will be needed for the FEv2.2 board as it becomes part of the unified board.

The distinguishing characteristics of the FEv2.2 boards are the large ground shields on both the top and the bottom of the board. The FEv2.2a board differs from the FEv2.2 board by eliminating the 2-pin and 3-pin SMT jumpers at each ADC and shifts the analog ground shield planes slightly to the left to give more space for the channel 0 ADC.

Schematic file FEv2_2.pdf is the schematic for FEv2.2a

Schematic file FEv2_2_2.pdf is the schematic for FEv2.2

# Text boxes

SURFACE DETECTOR ELECTRONICS BOARD

Pwr.

Supl.

GPS

C3

C3

## CPU

## TG

## ETH

## SC

C1

C2

C1

## FE

C2