# E1:  Investigation Of Errors

| Pattern of errors | Possible causes | Suggested Investigation Steps |
| --- | --- | --- |
| HFP and HFN | 1. Unusable microscope | 1. Examine a 3+ using that microscope |
|  | 2. Staining problems | 2. Check stains and staining procedure |
|  | 3. Technician cannot recognize AFB | 3. Test with clear-cut pos. / neg. and good microscope |
|  | 4. Gross neglect | 4. Exclude other causes |
| A single HFP | 1. Administrative error | 1. Compare lab-register with QC-listing: correct slide number & result? |
|  | 2. As for more frequent HFP | 2. Exclude causes of more frequent HFP |
| Regularly a HFP with or | 1. Poor registration routine | 1. Check accuracy of lab-register and other record keeping |
| without LFP | 2. Staining problems/Fading | 2.Check stains and staining procedure, consider restaining for rechecking |
|  | 3. Technician unclear on AFB appearance | 3. Look for inconsistent results of suspects (regularly single pos. / low positive) in lab register |
| Rare LFP | To be expected | No investigation unless numbers increase |
| Many LFP, with | 1. Problem with controllers | 1. Evaluate controllers |
| or without occasional HFP | 2. Technician unclear on AFB appearance | 2. Recheck special sample of LFP from laboratory register |
|  | 3. Contaminated stain reagents | 3. Test stain with known negative smears |
| Single HFN | 1. Administrative error | 1. Compare lab-register with QC-listing: correct slide number & result? |
|  | 2. Very thick smears and/or poor light | 2. Evaluate quality of smear preparation, check microscope |
|  | 3. Gross neglect | 3. Exclude other causes |
| Frequent HFN and/or | 1. Staining problems/Fading | 1.Check stains and staining procedure, consider restaining for rechecking |
| Many LFN | 2. Poor smearing-technique | 2. As above, single HFN |
|  | 3. Problems with microscope | 3.Check microscope with positive slide |
|  | 4. Careless microscopy | 4. Exclude other causes |
|  | 5. Contaminated stain reagents/water | 5. Test stain with known negative smears |
| Very high proportion LFN | Contaminated meth. blue or rinse water | As above |
| Many QE | 1. Poor staining | As above |
| (too low gradings) | 2. Problems with microscope | As above |