SMOS

SIX-MONTH CALIBRATION PROCEDURES

Facility:		EF

Frequency:	Twice a year

Mentor: 		Michael Ritsche

NOTE: Prior to all testing and verification, Flag 7 must be set high to prevent false data from being recorded.  This Flag should be reset when complete.

1.	05103 RM Young Wind Monitor

A.  Wind speed

To check that the wind speed sensor is measuring properly, the following equipment is needed.

1.	9/16" wrench

2.	#18810 Anemometer drive

3.	A computer or CR10KD

4.	Two people

Tilt the tower down to access the wind monitor. Refer to Appendix C for Tower lowering procedure. Remove the propeller from the propeller shaft and install the torque nut provided.  The anemometer drive connects to the wind monitor body and the motor connects to the torque nut.  Next, select one of the following RPM levels on the anemometer drive and start the unit.  Wait for the RPM level to stabilize.

RPM	RPM Reading

500            ___________

750            ___________

1000            ___________

Compare the reading from the anemometer drive and the sensor output as read from the computer or the CR10KD.  The RPM’s measured by the CR10 (Location 10) should be within +/- 1% of that indicated in the above table.

B.  Wind Direction

Using the angle fixture and a computer or a CR10KD, check that the sensor output (Location 2) matches the direction settings determined by the angle fixture #18212.  The angle fixture is inserted between the orientation ring and the wind monitor.  Once the angle fixture is installed, make sure that the notches of the orientation ring, angle fixture and wind monitor all match up.  Use the positioning arm to hold the tail of the wind monitor in the correct position.  Rotate the wind monitor through the positions listed below and compare them to directions measured by the data logger (Location 2).

Quadrant		Measured value

0			   ___________

90			   ___________

180			   ___________

270			   ___________

355			   ___________

The accuracy of this measurement is +/- 5.0 degrees, with a resolution of 1.0 degree.

NOTE: The absolute accuracy of this measurement depends on the exactness of the orientation of the sensor with respect to true north.

2.	HMP-35C Temperature & Relative Humidity

To verify the accuracy and installation of the HMP-35C sensor the following tools will be needed.

1.	HMI31/HMP35 Temperature and Relative Humidity

Meter, recently calibrated.

- Computer or CR10KD

- Portable Aspiration Unit.

Remove the sensor from the radiation shield and inspect, clean or replace the humidity cap.  Verify that the sensor is working properly using the following procedures.

The verification of the temperature sensor is done with the Vaisala HMI31/HMP35 T/RH meter.  Mount the portable aspiration unit on the SMOS tower approximately chest high.  Place the HMP-35C T/RH sensor and the HMI31/HMP35 T/RH Meter in the holders of the portable aspiration unit.  Turn on the HMI31/HMP35 meter.  Allow the sensors to stabilize such that the temperature and RH readings are changing by less that +/- 0.2 Deg C or +/- 1% RH, then wait another 5 minutes before taking readings.  The air temperature measurement is stored in the data logger in location 3 and the RH readings are stored in location 4.  You can monitor this location with either a computer or the CR10KD.  Compare the data logger measurement to the read-out of the HMI31/HMP35.  The temperature readings should match within +/- 0.5 C.  The resolution of the data logger measurement is 0.01 C.  The RH readings should match within +/- 2 % RH from 0 - 90 % RH or +/- 3 % RH from 90 - 100 % RH.  The resolution of the data logger measurement is 1.0 % RH.

If the Bi-weekly check is being done at the Central Facility the portable aspirator is not needed.  The SMOS site at the Central Facility has an aspirated shield.  The HMI31/HMP35 Meter may be held near the intake of the aspirated shield.  Take care to position the HMI31/HMP35 T/RH meter out of direct sunlight.

NOTE:  If the probe or the T/RH Meter has been exposed to a temperature significantly different from the ambient air temperature and RH, it may take an appreciable time for the sensor to stabilize.  They may require rapid waving before they will stabilize.

Test Results:

Measurement	Cal Standard	SMOS Sensor	Pass/Fail

Temperature	____________	___________	_________

Humidity		____________	___________	_________

3.  UDG01 Snow Depth Gauge (CURRENTLY NOT DONE)

NOTE:  The grass or vegetation under the gauge should be cut short and kept manicured to obtain accurate snow depth readings.

To start this test set Flag 8 high.  This is done either using the F8 key while monitoring the input locations in PC208W or using the CR10KD.   To use the CR10KD press *6AD8 to set the flag.  The flag will automatically reset in two minutes.  Once Flag 8 is high the UDG01 snow depth gauge will start making measurements once per second for two minutes.  The snow depth measurement is found in location 17 in the data logger, this measurement should be within +/- 25 mm of zero when nothing is under the snow depth gauge.  The snow depth gauge has a resolution of 0.5 mm.  A cardboard box of a known height at least 2 feet by 2 feet square is then placed under the gauge.  The measurement with the box under the gauge should show the height of the box within +/- 25 mm.

4.	2500E Rain/Snow Gauge

Two things need to be verified on this sensor.  The first is that the sensor records each tip of the bucket, and the other is to check that the heater portion of the rain gauge functions.

A.  Rain fall

The rainfall verification is accomplished with the 500 milliliter volumetric flask and plastic cup provided with the system.

Note this test should be started just after the even hour to insure that the system does not reset the number of counts during the test (This location resets every two hours).

Fill the volumetric flask to the indicator line and place the plastic cup in the top of the rain gauge.  Pour the 500 ml of water into the cup and allow the contents to slowly drain through the rain/snow gauge.  Using the data logger to measure the counts of the rain gauge the number of tips for 500 ml of water should be 25 to 27 counts; this value is stored in Location 16.  This test requires that the amount of water used is precisely 500 ml.  Once most of the water has drained from the cup, slowly pour the last of the water through the rain gauge.

B.  Heater

To verify the operation of the heater in the rain/snow gauge, the following test should be performed.  Place an ice cube in a plastic bag and hold the ice on the thermostat for 1 minute, and continue to hold in place while checking the heater (See Rain Gauge Manual in Installation Manual).  Feel the heater element at the base of the rain gauge and ensure that it has started to get hot.  The heater should be warm to the touch within 30 to 45 seconds after the thermostat has been tripped; the thermostat is set to trip at 37 degrees F.

5.  PTB-201 Barometric Pressure

The PTB-201 barometer is located inside of the enclosure.  A computer or the CR10KD can be used to monitor the location where the barometric pressure is being displayed (LOC: 18).  Establish communications with the data logger and monitor location 18 for barometric pressure in millibars.  Use the PTB-220 Digital Barometric pressure transducer to read the current barometric pressure in millibars and compare the two values.  The two values should be with in +/- 0.5 millibars of each other; the resolution of the sensor is 0.1 millibars.  The barometric pressure that is measured and recorded in the station (Location 8) is in Kilopascals and

is 1/10 of the millibars readings used in the verification of the sensor.

SURFACE METEOROLOGICAL OBSERVATION SYSTEM

SENSOR VERIFICATION FORM

SITE ID:_____________  LOCATION:____________________________

Tested by:__________________________  Date:_____/_____/_____

Started: __________________GMT   Ended: _________________GMT

1.	05103 Wind Monitor

A.	Wind Speed

RPM      Measured RPM (LOC: 10)	PASS/FAIL

500		 _________			_________

750		 _________			_________

1000		 _________			_________

The measured RPM should be within +/- 1% of the motor drive RPM.

Acceptable: YES  /   NO

B.	Wind Direction

Quadrant	  Measured value (LOC: 2) PASS/FAIL

0		   ___________			 _________

90		   ___________			 _________

180		   ___________			 _________

270		   ___________			 _________

355		   ___________			 _________

The accuracy of this measurement is +/- 5.0 degrees, with a resolution of 1.0 degree.

Acceptable: YES  /   NO

2.	HMP-35C Temperature & Relative Humidity

Measurement	Cal Standard	SMOS Sensor	Pass/Fail

Temperature	____________	___________	_________

Humidity		____________	___________	_________

(Temperature readings are in LOC: 3, RH in LOC:4)

The temp readings once stabilized should match within +/- 0.5 C.  The resolution of the data logger measurement is 0.01 C.

The RH readings once stabilized should match within +/- 2 % RH from 0 - 90 % RH, +/- 3 % RH from 90 - 100 % RH.  	The resolution of the measurement is 1.0 % RH.

Acceptable: YES  /   NO

3.	UDG01 Snow Depth Gauge  (not done at this time)

(LOC: 17)

Object Height Measured:__________   UDG01:______________

This measurement should be within +/- 25 mm.

Acceptable: YES  /   NO

4.	2500E Rain/Snow Gauge

A.	Rain, 500 ml should produce 25 to 27 tips

Number of tips recorded:________________ (LOC: 16)

Acceptable: YES  /   NO

B.	Heater

Acceptable: YES  /   NO

5.	Barometric Pressure

PTB-220 Reading:____________

PTB-201 Reading:____________ (LOC: 8)

The readings should agree within +/- 0.05 kPa or 0.5 mb.

Acceptable: YES  /   NO

6.	Comments: