Principal Investigator: 	Max Loewenstein

Co-Investigators: 	Jimena Lopez, Jeffrey Grose

Organization:	Mail Stop 245-5

Atmospheric Chemistry and Dynamics Branch

Ames Research Center

National Aeronautics and Space Administration

Moffett Field, CA 94035-1000

Instrument Description

Argus is a two channel, tunable diode laser instrument set up for the simultaneous, in situ measurement of CO (carbon monoxide), N2O (nitrous oxide) and CH4 (methane) in the troposphere and lower stratosphere. The instrument measures 40 x 30 x 30 cm and weighs 21 kg.

The instrument is an absorption spectrometer operating in rapid scan, second-harmonic mode using frequency-modulated tunable lead-salt diode lasers emitting in the mid-infrared. Spectra are co-added for two seconds and are stored on a solid state disk for later analysis. The diode laser infrared beam is shaped by two anti-refection coated lenses into an f/40 beam focused at the entrance aperture of a multi-pass Herriott cell. The Herriott cell is common to both optical channels and is a modified astigmatic cell (New Focus Inc., Santa Clara, California).

The aspherical mirrors are coated with protected silver for optimal infrared reflectivity. The cell is set up for a 182-pass state for a total path of 36m. The pass number can be confirmed by visual spot pattern verification on the mirrors observed through the glass cell body when the cell is illuminated with a visible laser beam. However, instrument calibration is always carried out using calibrated gas standards with the Argus instrument operating at its infrared design wavelengths, 3.3 and 4.7 micrometers respectively for CH4 and CO/N2O detection. The electronic processing of the second harmonic spectra is done by standard phase sensitive amplifier techniques with demodulation occrring at twice the laser modulation frequency of 40 kHz. To optimize the second-harmonic signal amplitude in a changing ambient pressure environment the laser modulation amplitude is updated every 2 seconds to its optimal theoretical value based upon the measured pressure in the Herriott cell.

Calibration

Argus is regularly calibrated in the laboratory against a CMDL (Climate Monitoring and Diagnostics Lab) whole air standard for CH4, N2O and CO. Flow-mixing of this standard with a zero gas provides us with absolute calibration and linearity data as inputs to the in situ flight recorded analysis. During field operations Argus is calibrated in the lab both before and after every flight operation with the instrument maintained in its flight setup configuration throughout the calibrations.

Performance Quality at 18 km

| Tracer | Typical Mixing ratio (ppb) | Data Rate (Hz) | Precision (ppb) | Accuracy (%) 1-sigma |
| --- | --- | --- | --- | --- |
| CO | 35 | 0.5 | 1.3 | 3 |
| CH4 |  | 0.5 | 9.8 | 1.5 |
| N2O (estimated) |  | 0.5 | 3.9 |  |

Data Analysis

The post-flight data reduction uses a non-linear, least squares Marquardt-Levenberg fitting procedure where theoretical fitting function is the second Fourier component of the modulated Voigt absorption line-shape. One parameter returned by the fitting procedure is the molecule number density (CO, N2O and CH4) in the Herriott cell. This is converted to mixing ratio using the measured gas density in the cell. Argus then reports molecule mixing ratio at 0.5 Hz data rate to the data archive.