Building 510B

P.O. Box 5000

Upton, NY 11973-5000

Phone 631 344-3761

Fax 631 344-2739

lastname@bnl.gov

managed by Brookhaven Science Associates
for the U.S. Department of Energy

date:		May 23, 2003

to:		Helen Kerch, DOE

from:		Steve Shapiro,	Igor Zaliznyak

subject:	Answer to the referee reports on HYSPEC Proposal

We are pleased that two referees judged the HYSPEC proposal to be an excellent one, and have expressed their strong support for the construction of the instrument. While the negative response of the third referee is certainly disappointing, we believe that the referee’s concerns resulted mainly from some deficiencies in the text and the layout of the proposal, but not in its essence, nor in the instrument’s conceptual design. Below, we present arguments which address in detail the issues raised by the third referee, and which support our conclusion.

- The introductory remark, that the proposal “presents a less detailed description of the instrument…” has raised our concern of a possibility that somehow, by an error or omission, the proposal was not available to the referee in its entirety. While in the main text of the proposal we indeed included only a brief, 10-page overview of the instrument’s conceptual design and main operational characteristics, a very detailed, 15-page formal description of the instrument was attached in Appendix C, “HYSPEC Top Level Specifications” (TLS).  It is referred to in the Introduction, and in the Section 3 of the main document.

- On the technical side, the main concern of the referee is that HYSPEC “has not been optimized to take into account the other instruments in the SNS…”, and, specifically, that he sees “the reason for (HYSPEC) long (secondary) flight path is to obtain the resolution that matches ARCS, CNCS and SEQUOIA”. Referee further suggests that “optimizing the instrument for its strength – intensity – would also result in  … savings”, and proposes to reduce the HYSPEC secondary flight-path to 2 – 2.5 m, simultaneously  increasing the detector coverage, ie the surveyed momentum space. To the referee’s opinion, this could result in a very significant cost savings without serious impact on the instrument performance.

These suggestions raise a number of vital issues in the conceptual design of HYSPEC and, to that matter, of any thermal-neutron spectrometer for single crystal studies, which, perhaps, were lacking clear exposition in the text of the proposal. In fact, there are several very important reasons for which the secondary neutron flight-path in such spectrometer has to be long, and in the case of HYSPEC can not be shorter than 4 m. In short, these are,

  - a spectrometer performs best if the energy resolution of the analyzer and the monochromator are matched; the shorter is the analyzer flight-path, the more out of balance and less optimized is the spectrometer for the energy transfers of interest

  - reducing the length of the secondary flight-path at any (coarse or fine) fixed analyzer resolution requires proportional decrease in the sample illumination time and, therefore, results in a proportionally less efficient use of the long neutron pulse supplied by the coupled H2 moderator

  - it follows from the translational invariance of the crystalline material that an increase in the reciprocal space surveyed by the instrument beyond one Brillouin zone (BZ) does not result in the proportional increase in instrument throughput

  - efficient separation in the detector of up and down neutron polarization in the supermirror-bender type transmission polarization analysis setup previewed for HYSPEC requires that the secondary flight-path is at least 4 m long

It is essentially the throughput (intensity) optimization, which is achieved by matching the instrument with the moderator that requires long secondary flight-path in the HYSPEC design. A 4.5 m long analyzer is roughly matched with the FWHM of the coupled H2 moderator pulse for the coarse, but acceptable, ΔE/E resolution of about 10% for Ei= [10,60] meV. Relaxing the resolution beyond 10% is unacceptable for most experiments, while reducing the secondary flight-path within the resolution constraint requires shorter sample illumination time, so that less neutrons provided by the moderator are used. In fact, in a recent development the SNS Instrument Team in close interaction with HYSPEC IDT has proposed a new SNS floor layout, where HYSPEC is located on beamline 14B and fits within the SNS building even with 4.5 m long secondary flight-path.

A spectrometer that is designed to operate in the cold neutron range could have shorter secondary flight path, because the time of flight of the scattered cold (slow) neutrons in the analyzer is longer and allows for the longer sample illumination times. Such is the case of FOCUS at PSI cited by the referee, which is optimized at Ei=5 meV and has 2.5 m long analyzer. To match the FOCUS elastic resolution at Ei=50 meV with similar sample illumination time, a thermal neutron spectrometer would require 25 m long secondary flight-path. While sufficiently long secondary flight-path in HYSPEC can, in principle, be used to achieve a higher energy resolution at a cost of lower intensity, this is not the regime that the instrument was designed for. This is illustrated by the performance comparison for HYSPEC, CNCS, ARCS and SEQUOIA by Granroth and Abernaty, that the referee had referred to. Contrary to the referee’s impression, though, HYSPEC in its current design is not capable of operating at high resolutions those other instruments are optimized for. To achieve such resolution, HYSPEC setup was deliberately changed by Granroth and Abernathy, by reducing the opening in the disc chopper. This has a detrimental impact on the intensity, and is not previewed in the HYSPEC design, although, even with this handicap, HYSPEC noticeably outperforms the other instruments in the 10 - 30 meV thermal neutron range, where it is best optimized.

Finally, we comment on the referee remarks that neither the immediate polarized beam capability nor the “path that detectors cut through Q-E space”, which “are the same for all instruments” make HYSPEC unique among the SNS instruments.

  - While we are ready to share the referees enthusiasm on the perspectives of 3He filter-polarizers, and do allow for an upgrade to one of those once they become available in the HYSPEC design, we can not rely entirely on their future availability. In fact, 3He polarizers with large angular acceptance which are suitable for use in the instruments such as HYSPEC probably will not be operational within a decade or more, which means that for at least first 5-10 years of SNS operation HYSPEC is bound to be the only polarized beam inelastic spectrometer there. In addition, the maintenance and operational costs of 3He filter-polarizers are likely to be high, in which case their eventual availability will also depend on an additional substantial funding. HYSPEC polarized beam setup is essentially maintenance free.

  - The phase space probed by HYSPEC for each fixed incident energy Ei is indeed the same as measured by the corresponding detectors in any other inelastic instrument at SNS, a Ei = const scan. It is well-known, though, that such measurement has a significant handicap that the momentum space probed shrinks rapidly with the increasing energy transfer E; to cover the same amount of reciprocal space at higher E one has to increase Ei. HYSPEC’s intended mode of operation is such that a number of measurements with different Ei are performed which cover roughly one Brillouin zone (but always the same one), for extended range of energy transfers. This is achieved by rotating the sample and properly positioning the detector bank, so that the measurement is focused on the same region of the reciprocal space for all Ei. A similar mode of operation is possible for other TOF spectrometers (although, it is severely restricted for the Fermi-chopper instruments, where only few discrete Ei are possible). Provided the sample is properly aligned as incident energy changes, at each Ei the Brillouin zone of interest would be measured by different parts of a large detector bank. In HYSPEC, a smaller detector bank is used with the high efficiency, as it follows the movement of the sample momentum space with changing Ei, and always probes the same part of it, which is of interest to the experimenter.

- While we are pleased that referee acknowledges the strong scientific case presented in the HYSPEC proposal, we find his reservations that “the focus of the IDT is quite narrow” and “there is very little participation of the general university research community” surprising. Indeed, the HYSPEC IDT was organized by scientists at Brookhaven National Laboratory, and therefore has a very significant BNL representation. And while the situation when “the great bulk of the members are at the existing neutron sources at the national labs” does not seem un-appropriate for an IDT that builds a neutron spectrometer at a DOE laboratory, we note that HYSPEC IDT also includes participants from almost all American universities that have neutron scattering program in condensed matter. In fact, they account for half of the non-BNL membership of the HYSPEC IDT. In addition to the U. Missouri and Iowa University mentioned by the referee, MIT, Stanford, Rutgers, and McMaster University of Canada are represented. The real issue here is the shrinking support for the neutron scattering effort at the US universities. Princeton group previously lead by Bernhard Keimer is inexistent, formerly mighty MIT neutron scattering program is now reduced to one tenure-track professor and a postdoc, Martin Greven’s group at Stanford shifts focus to X-ray scattering, Rutgers group of Valery Kiryukhin, who studied in neutron scattering with Keimer and Birgeneau, has now only a rudimentary neutron program. The scientific effort previewed in the HYSPEC budget is essentially aimed at addressing this problem, and it is gratifying that the referee acknowledges its importance. The funding will be used to revive the neutron scattering efforts at the participating universities, in particular, by hiring both postdocs and graduate students to work on the HYSPEC-related scientific projects. We also envision that a neutron scattering program in condensed matter could be conceived in the Physics Department of our own SUNY at Stony Brook.

Addressing the referee suggestion to open the IDT to “a much broader range of participation” we note that HYSPEC IDT membership policy is quite open, it does not require any qualifications, only approval by the Executive Committee. In fact, since the IDT was established it has grown by a number of members, and nobody was refused the participation.

- We conclude by addressing the referee comments on the HYSPEC estimated budget. For us, it is a painful issue because it was prepared with a very limited funding and without the engineering support appropriate for the project of this scale. The budget is essentially based on the good faith estimates that we were able to make at the time. For the few components which are common with other instruments at the SNS we obtained the cost estimates from the SNS Instrument Systems. These costs are reasonably well defined, and we are working closely with the SNS to update them as information of the new procurements become available. However, more than half of the construction costs are HYSPEC-specific and can not be simply copied from other spectrometers. Realistic budget estimate requires drafting a detailed specifications for each of the components based on the HYSPEC TLS, and obtaining the cost estimates from the manufacturers. This implies significant engineering effort, which we believe could be completed within 6 to 12 month from the time it is funded.

To answer the most disappointing remark of the referee, that “the budget for the instrument seems excessive”, we note that main savings ($3-4 M) that the referee proposed result from reducing the spectrometer secondary flight-path from 4.5 to 2 m. As we discussed in detail above, this would have a detrimental and untolerable impact on the instrument performance. More than two-fold loss of flux-on-sample at any given resolution is apparent, and many other disadvantages follow. Also, a wide 4 cm guide used by HYSPEC does not allow employing a single disc chopper in place of the counter-rotating pair, thus eliminating another source of savings identified by the referee.

In fact, for proper comparison with the funded instruments, such as ARCS and CNCS, one has to look at the HYSPEC FY 2003 estimated construction cost, which is k$14,750 (there was an unfortunate error of $1.2M in inflating the costs to FY2008, that was correctly noticed by the referee). Taking into account the higher level of mechanical complexity of the Hybrid Spectrometer, and also the fact that this price tag includes the polarized beam option which will add up to $1M or more to the cost of the other instruments, we believe that the budget does not look excessive.

With best regards,

Igor Zaliznyak

Steve Shapiro

Appendix. More on HYSPEC design considerations.

There is currently no much experience in designing and building neutron spectrometers optimized for single-crystal studies on the pulsed spallation sources, MAPS at ISIS being essentially the first attempt of this kind. HYSPEC is a novel design, which appears to offer significant new advantages for these particular studies. Naturally, it also incorporates the well-known guiding principles of the design of a time-of-flight (TOF) spectrometer on a pulsed neutron source, which are tested by time. Below, we discuss in more detail some of the considerations that shaped the HYSPEC design, in connection with the concerns raised by the referee.

- HYSPEC is optimized for studies of the dynamic properties of crystalline materials in the thermal neutron range [10,60] meV. Translational invariance is a fundamental physical property that sets crystals apart from other materials, such as glasses and liquids. It has important consequences for the design of a single crystal inelastic spectrometer. The most immediate result of the translational invariance (lattice periodicity) is that the momentum in the crystal is replaced by the quasi-momentum, which is defined only within the Brillouin zone (BZ). The physical properties are periodic in the momentum space, and, while covering large part of the reciprocal space is undoubtedly useful ( eg for structure refinement or background characterization), the response in a single Brillouin zone essentially provides the bulk of the information on the dynamics of the crystal. In other words, while a spectrometer that is capable of measuring the phonon dispersion throughout 8 Brilluoin zones does yield more information compared to the one that covers one BZ, the gain is far from the naively expected 8-fold increase in throughput. As a result, the historical common wisdom, that efficiency of the TOF spectrometer grows proportionally to the analyzer angular acceptance, and that a loss in the flux on sample can be easily compensated by the increase in a surveyed momentum space, does not hold. A reliable measurement requires that a data with a satisfactory signal to background ratio are collected in at least one Brillouin zone. In fact, because neutron scattering cross-section varies strongly with Q, from one Brillouin zone to another, combining the data from different BZ is not straightforward and, in many cases the increase in the surveyed momentum space does not increase the data collection rate at all. This is in contrast with the disordered systems, such as liquids or glasses, where the collected information is not periodic in the momentum space, and, therefore, grows roughly in proportion with the increase in reciprocal space surveyed by the spectrometer. Because non-crystalline systems were for decades the primary object of studies with the TOF spectrometers, it became accepted in their design that the solid angle covered by the detectors is maximized at the expense of the secondary flight-path, which is made as short as possible, while consistent with the resolution that the instrument is intended for.

Here is how the important consideration formulated above applies to HYSPEC. Even for a typical metal, with the lattice constant as short as about 3.1415, and the Brillouin zone as large as 2 Å-1, the analyzer acceptance of 60 degrees covers, in elastic scattering, about 1.3 BZ for Ei = 15 meV neutrons. Because the data collection rate cannot be efficiently increased by increasing the detector angular coverage once more than one Brillouin zone is surveyed, the only way to optimize the instrument is to use as large fraction of neutrons produced in the moderator pulse, as possible. This implies that the sample illumination time, τ, is increased, at least to match the full width at half maximum (FWHM) of the neutron pulse prodced by the moderator. For a given analyzer energy resolution, the increase in τ requires that the secondary flight-path is increased proportionally. Because the pulse FWHM of the coupled hydrogen moderator is rather long, up to 100 μs (it decreases with the increasing incident neutron energy, but so does the time of flight in the analyzer), matching it with the pulse on sample requires a secondary flight-path of 4-5 m for energy resolution of 5% to 10%. In fact, the counter-rotating disc choppers in the design of HYSPEC do not allow for τ < 40 μs which would be required to achieve the high energy resolutions comparable to CNCS or SEQUOIA.

- For the optimum performance of an instrument energy resolutions of the primary and the secondary spectrometer (monochromator and analyzer) have to be matched for some useful value of energy transfer E. This condition requires that the time of flight in the monochromator for neutrons with energy Ei and in the analyzer for neutrons with energy Ei–E are roughly equal. If the monochromator is twice longer than the analyzer, the instrument is balanced for the energy transfer of roughly half of the incident energy. This is close to the optimum setup, and is implemented in the case of MAPS at ISIS. Of the spectrometers currently previewed for the SNS, the shortest, 17.5 m primary flight-path is planned for SEQUOIA. Combined with the 6 m long analyzer it makes this instrument balanced within the useful energy transfer range, at E ≈ 2/3Ei. For an instrument at 20 or more meters from the moderator, such as HYSPEC, the analyzer is bound to be appreciably shorter than the monochromator, and the spectrometer is not balanced for most energy transfers. However, the longer is the secondary flight-path (analyzer), the less out of balance, and therefore better optimized, the instrument is. Because analyzer controls the energy resolution for low E under such conditions, the longer secondary flight-path also leads to better elastic energy resolution. Our goal in developing the HYSPEC design, though, was to optimize the instrument by balancing the setup as much as possible, and not to achieve high elastic energy resolution, as referee suggested.

- Additional restriction on the length of the secondary flight-path is imposed by the efficient polarization analysis scheme based on the transmission polarizers, envisaged for HYSPEC. Because spatial separation of neutrons with up and down polarizations in the detector is required in order to measure them simultaneously, thus doubling the instrument efficiency, the secondary flight-path should not be shorter than about 4 m. It is important to mention that issue of the floor space that is required to accommodate the HYSPEC long secondary flight path was successfully resolved with the instrument relocation to the beamline 14B, and therefore no additional expense required by going to the outside building is envisaged anymore.

While we are ready to share the referees enthusiasm on the perspectives of 3He filter-polarizers, and do allow for an upgrade to one of those once they become available in the HYSPEC design, we can not rely entirely on their future availability. In fact, 3He polarizers with large angular acceptance, suitable for use in the instruments such as HYSPEC, probably will not be available within a decade or more, which means that for at least first 5 to 10 years of SNS operation HYSPEC is bound to be the only polarized beam inelastic spectrometer there.

- Finally, referee notes that “the path that the detectors cut through Q-E space are the same for all the instruments”, and that because of smaller detector coverage HYSPEC “appears to be less flexible…”. The issue of the detector coverage was addressed above; HYSPEC allows more than a whole Brillouin zone of a crystalline material to be surveyed at once. More importantly, the remark of the referee is based on the assumption that HYSPEC is operated in exactly the same way as the other inelastic spectrometers, ie essentially collecting data in the constant-Ei scans. This, however, is not the intended mode of operation for the hybrid spectrometer! Actually, the constant-Ei mode of operation is notoriously inefficient, which is why all or almost all inelastic studies with triple-axis spectrometers are made with the fixed final energy. Nevertheless, a direct geometry TOF spectrometer is bound to perform a scan with fixed incident energy, and analyze Ef. The main disadvantage is that the momentum space surveyed by the analyzer shrinks at higher energy transfers, and the instrument throughput decreases very fast, where it is most needed! HYSPEC offers a unique opportunity to perform a measurement, where the incident energy is varied in steps, much like in a 3-axis spectrometer, so that the volume in the phase space surveyed by the instrument is matched with the scattering intensity. For example, to characterize a spin-wave dispersion with bandwidth 40 meV in a typical HYSPEC experiment, one would perform a number of measurements, a “scan”, with Ei = 5,10,15,20,30,40,60 meV, spending different amount of time at each incident energy, as required to obtain similar statistical error on the data for each Ei. By rotating the sample, and properly positioning the detector bank, the measurement will always concentrate on the same region of reciprocal space. The combined dataset will result in a high-quality measurement, whose resolution and statistical error will depend on the energy transfer much less than those of the equivalent traditional, constant-Ei data.

# Text boxes

Memo