| Name(s) of Risk Team Members: Diane Cabelli, Michael Schueller | Point Value → Parameter ↓ | 1 | 2 | 3 | 4 | 5 |
| --- | --- | --- | --- | --- | --- | --- |
| Job Title: 11-C Target Foil Change Job Number or Job Identifier: DJ-JRA26-Cyclotron | Frequency (B) | <once/year | <once/month | <once/week | <once/shift | >once/shift |
| Job Description: Potential dispersible radiation, RWP | Severity (C) | First Aid Only | Medical Treatment | Lost Time | Partial Disability | Death or Permanent Disability |
| Training and Procedures List: As defined in ESR CO-11-13, RWP for routine maintenance on EBCO Cyclotron, SAD for EBCO Cyclotron | Likelihood (D) | Extremely Unlikely | Unlikely | Possible | Probable | Multiple |
| Approved by: R. Colichio Date: 7/24/06 Rev. #: 0 |  |  |  |  |  |  |
| Stressors (if applicable, please list all): None. | Reason for Revision (if applicable): | Comments: Space is large and open so there are no significant tripping hazards, biggest hazard is a cut from an edge of the foil, gloves are worn but they must be maximized for dexterity. No mechanical parts so any possible cuts involve physical pressure only. |
|  | Before Additional Controls |  | After Additional Controls |  |
| Job Step / Task | Hazard | . Control(s) | Stressors Y/N | # of People A | Frequency B A | Severity C | Likelihood D | Risk* AxBxCxD | Control(s) Added to Reduce Risk | Stressors Y/N | # of People A | Frequency B | Severity C | Likelihood D | Risk* AxBxCxD | % Risk Reduction |
| Open shields to target chamber, radcon survey to verify dose rate and contamination levels are appropriate, verify training, PPE for all personnel in room and ensure they are signed on the RWP. | Shields are ca 10 tons, any encounter with moving shields is hazardous | Shields opened hydraulically where button is positioned such that operator is far from shielding and in direct vision line of shielding so no other person can be inadvertently in the way of the shield | N | 2 | 2 | 5 | 1 | 20 |  |  |  |  |  |  |  |  |
|  | Exposure to radioactive material | No entrance until FST verifies that contamination and radiation levels are below accepted limits. RWP, PPE | N | 1 | 2 | 1 | 2 | 4 |  |  |  |  |  |  |  |  |
| Verify target is not under pressure, valve off deionized cooling water to side 1. Take airline and blow out residual water into reservoir (continual flow, not single pass water) | Exposure to high pressure. | Pressure gauge to read pressure externally. If windows are blown, there is obviously no pressure on target. Pressure volume is first accessed through a 1/16" Swagelok connection, so even if the target were pressurized, the threat is minimal., PPE | N | 1 | 2 | 1 | 1 | 2 |  |  |  |  |  |  |  |  |
| Cut and clean 2 target foils (0.001 inch aluminum and 0.001 inch Havar-Fe/Co/Ni/Cr alloy) | cuts | Can be cut with scissors. | N | 1 | 2 | 1 | 2 | 4 |  |  |  |  |  |  |  |  |
| Take off water jacket to expose target, unscrew to separate flange, He cooling section and target body. Remove old targets and place in plastic bag for survey by FST, change gloves | Radiation exposure | All activities carried out in designated radioactive materials area, aluminum is activated from beam but has a 5 second half life, Havar may be active. Havar foil is quickly removed and placed behind lead shielding. PPE | N | 1 | 2 | 1 | 2 | 4 |  |  |  |  |  |  |  |  |
| Reassemble target using new foils, replace at target site in cyclotron, turn on cooling water, check for water leaks, turn on vacuum. | No hazards | No controls | N | 1 | 2 | 1 | 1 | 2 |  |  |  |  |  |  |  |  |
| Close shields to target chamber | See first section | See first section | N | 1 | 2 | 5 | 1 | 10 |  |  |  |  |  |  |  |  |
| Further Description of Controls Added to Reduce Risk: |
| *Risk: | 0 to 20 | 21 to 40 | 41-60 | 61 to 80 | 81 or greater |
|  | Negligible | Acceptable | Moderate | Substantial | Intolerable |