summary of safety and effectiveness

I.  General Information

Device Generic Name:	Dual Chamber Implantable 
Cardioverter Defibrillator (ICD) with Cardiac Resynchronization Therapy

Device Trade Name:	InSync® ICD Model 7272 Dual Chamber Implantable Cardioverter Defibrillator with Cardiac Resynchronization Therapy and the Model 9969 Application Software

Applicant’s Name and Address:	Medtronic, Inc.

710 Medtronic Parkway

Minneapolis, MN  55432-5604

PMA Number:			P010031

Date of Panel Recommendation:		March 5, 2002

Date of Notice of Approval to Applicant:	June 26, 2002

II.  Indications for Use

InSync® ICD Model 7272 Device

The InSync® ICD Model 7272 is indicated for ventricular antitachycardia pacing and ventricular defibrillation for automated treatment of life threatening ventricular arrhythmias.  The system is also indicated for the reduction of the symptoms of moderate to severe heart failure (NYHA Functional Class III or IV) in those patients who remain symptomatic despite stable, optimal medical therapy (as defined in the clinical trials section), and have a left ventricular ejection fraction less than or equal to 35% and a QRS duration greater than or equal to 130 ms.

III.  System Description

Description of InSync® ICD Model 7272

The InSync® ICD Model 7272 Dual Chamber Implantable Cardioverter Defibrillator (ICD) System is a multi-programmable, implantable cardioverter defibrillator with biventricular pacing for cardiac resynchronization therapy (CRT) that monitors and regulates a patient’s heart rate by providing ventricular tachyarrhythmia therapy and single or dual chamber bradycardia pacing.  Ventricular tachyarrhythmias therapies are for the treatment of ventricular tachycardia (VT) and ventricular fibrillation (VF), arrhythmias associated with sudden cardiac death (SCD).  Cardiac resynchronization therapy uses simultaneous biventricular electrical stimulation to synchronize ventricular contractions.

The Model 7272 InSync® ICD System consists of:

- InSync® ICD Model 7272

- Model 9969 Application Software for use with the previously approved:

          - Model 9790C Programmer

          - Model 2090 Programmer

- Commercially Available Components

          - Attain LV Model 2187 Cardiac Vein Pacing Lead

          - Attain CS Model 2188 Coronary Sinus Pacing Lead

          - Attain Model 6215 Venogram Balloon

          - Attain Models 6216 and 6218 Left Heart Delivery Systems

          - Right Atrial, Right Ventricular and Cardioversion/Defibrillation leads, implant support instruments and accessories

IV.  Contraindications

InSync® ICD Model 7272 Device

Do not use the InSync® ICD system in:

- Patients whose ventricular tachyarrhythmias may have transient or reversible causes, such as:

          - Acute myocardial infarction

          - Digitalis intoxication

          - Drowning

          - Electrocution

          - Electrolyte imbalance

          - Hypoxia

          - Sepsis

- Patients with incessant VT or VF

- Patients who have a unipolar pacemaker

V.  Warnings and Precautions

See device labeling.

VI.  Adverse Events

Per the investigational plan, an adverse event is defined as any undesirable experience (sign, symptom, illness, or other medical event) occurring to a patient whether it is considered to be device related or not. Adverse events were classified as complications or observations based on the following definitions:

- A complication is defined as an adverse event that results in invasive intervention, or directly results in the death or serious injury to the patient, the explant or repositioning of the ICD or lead, or the termination of significant device function regardless of other treatments.

- An observation is defined as an adverse event that does not result in invasive intervention, the death or serious injury to the patient, the explant or repositioning of the ICD or lead, or the termination of significant device function.

On the following pages, Table 1 summarizes all reported adverse events in NYHA Class III and IV patients; and Table 2 summarizes NYHA Class III and IV patient deaths.

The InSync® ICD study was a prospective, multi-center, randomized double-blind, parallel arm, controlled clinical trial to assess the safety and effectiveness of biventricular pacing for heart failure therapy in ICD patients.

A.  Adverse Events

Table 1 provides a summary of all adverse events that occurred in NYHA Class III and IV patients from implant through the 6-month randomization period who had an implant attempt or were implanted with an InSync® ICD system.  1499 events were reported in 351 of 424 patients implanted or attempted with an InSync® ICD system (1861 total device months). Of the events, 459 were complications and 1040 were observations.

Table 1. Adverse Events During the Randomization Period

|  | # of Events (# of patients) | % Complications (# of patients) | Complications per 100 device-months (# of events) | % Observations (# of patients) | Observations per 100 device-months (# of events) |
| --- | --- | --- | --- | --- | --- |
| Total Adverse Events | 1499 (351) | 49.4 (208) | 24.7 (459) | 75.8 (319) | 55.9 (1040) |
| ICD related events |
| Abnormal impedance measurement | 2 (2) | 0.2 (1) | 0.1 (1) | 0.2 (1) | 0.1 (1) |
| Atrial or ventricular arrhythmias | 4 (4) | 0.0 (0) | 0.0 (0) | 1.0 (4) | 0.2 (4) |
| Electrical reset of ICD | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| ICD discharge | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Inappropriate VT/VF detection | 5 (5) | 0.0 (0) | 0.0 (0) | 1.2 (5) | 0.3 (5) |
| Pacemaker mediated tachycardia | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Phantom shocks | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Pocket infection/ seroma/hematoma | 2 (2) | 0.2 (1) | 0.1 (1) | 0.2 (1) | 0.1 (1) |
| Subtotal: ICD related events1 | 19 (17) | 1.0 (4) | 0.2 (4) | 3.3 (14) | 0.8 (15) |
| LV lead related events |
| Cardiac perforation | 2 (2) | 0.5 (2) | 0.1 (2) | 0.0 (0) | 0.0 (0) |
| Coronary Sinus dissection | 5 (5) | 0.7 (3) | 0.2 (3) | 0.5 (2) | 0.1 (2) |
| Elevated pacing thresholds/ Failure to capture, loss of capture | 12 (11) | 2.1 (9) | 0.5 (9) | 0.7 (3) | 0.2 (3) |
| Lead dislodgment | 33 (31) | 6.7 (28) | 1.6 (30) | 0.7 (3) | 0.2 (3) |
| Muscle stimulation (diaphragm, chest wall, pectoral) | 32 (29) | 2.9 (12) | 0.6 (12) | 4.5 (19) | 1.1 (20) |
| Pain (chest; neck) | 2 (2) | 0.0 (0) | 0.0 (0) | 0.5 (2) | 0.1 (2) |
| Palpitations | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Pericardial effusion | 2 (2) | 0.2 (1) | 0.1 (1) | 0.2 (1) | 0.1 (1) |
| Ventricular arrhythmias | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Subtotal: LV lead related events | 90 (77) | 12.6 (53) | 3.1 (58) | 7.1 (30) | 1.7 (32) |
| RA lead related events |
| Atrial arrhythmias | 2 (2) | 0.2 (1) | 0.1 (1) | 0.2 (1) | 0.1 (1) |
| Atrial standstill | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Elevated pacing thresholds/ Failure to capture, loss of capture | 7 (6) | 1.0 (4) | 0.2 (4) | 0.7 (3) | 0.2 (3) |
| Far-field R-wave sensing | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Lead dislodgment | 6 (6) | 1.4 (6) | 0.3 (6) | 0.0 (0) | 0.0 (0) |
| Oversensing, undersensing | 2 (2) | 0.0 (0) | 0.0 (0) | 0.5 (2) | 0.1 (2) |
| Subtotal: RA lead related events | 19 (18) | 2.9 (12) | 0.6 (12) | 1.7 (7) | 0.4 (7) |
| RV lead related events |
| Cardiac perforation | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Elevated pacing thresholds/ Failure to capture, loss of capture | 2 (2) | 0.5 (2) | 0.1 (2) | 0.0 (0) | 0.0 (0) |
| Heart block | 3 (3) | 0.5 (2) | 0.1 (2) | 0.2 (1) | 0.1 (1) |
| Lead dislodgment | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Muscle stimulation (diaphragm, pectoral) | 2 (2) | 0.5 (2) | 0.1 (1) | 0.2 (1) | 0.1 (1) |
| Oversensing | 4 (4) | 0.0 (0) | 0.0 (0) | 1.0 (4) | 0.2 (4) |
| PVCs | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Poor DFTs | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Subtotal: RV lead related events | 15 (15) | 1.9 (8) | 0.4 (8) | 1.7 (7) | 0.4 (7) |
| System related events |
| Elevated pacing thresholds/ Failure to capture, loss of capture | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Inappropriate VT/VF detection | 2 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (2) |
| Muscle stimulation - diaphragm | 2 (2) | 0.0 (0) | 0.0 (0) | 0.5 (2) | 0.1 (2) |
| Phantom shocks | 2 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (2) |
| Programmer malfunction | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Twiddler’s syndrome | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Subtotal: System related events1 | 9 (7) | 0.5 (2) | 0.1 (2) | 1.2 (5) | 0.4 (7) |
| Implant tools related events |
| Cardiac perforation/pericardial effusion | 4 (4) | 0.7 (3) | 0.2 (3) | 0.2 (1) | 0.1 (1) |
| Chest pressure/tightness | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Coronary Sinus dissection | 8 (8) | 1.4 (6) | 0.3 (6) | 0.5 (2) | 0.1 (2) |
| Guide catheter damage | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Guidewire fracture | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Heart block | 7 (7) | 0.2 (1) | 0.1 (1) | 1.4 (6) | 0.3 (6) |
| Hemo/Pneumothorax | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Lead stylet stuck | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Ventricular arrhythmias | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Subtotal: Implant tools related events | 25 (22) | 3.1 (13) | 0.8 (14) | 2.4 (10) | 0.6 (11) |
| Possibly device related |
| Atrial arrhythmias | 12 (11) | 0.5 (2) | 0.1 (2) | 2.1 (9) | 0.5 (10) |
| Bradycardia/heart block/ junctional rhythm | 3 (3) | 0.0 (0) | 0.0 (0) | 0.7 (3) | 0.2 (3) |
| Chest pain/angina pectoris | 7 (7) | 0.5 (2) | 0.1 (2) | 1.2 (5) | 0.3 (5) |
| Dizziness | 4 (4) | 0.0 (0) | 0.0 (0) | 1.0 (4) | 0.2 (4) |
| Fever | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Hypertension | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Hypotension | 3 (3) | 0.7 (3) | 0.2 (3) | 0.0 (0) | 0.0 (0) |
| Ventricular arrhythmias | 9 (8) | 0.2 (1) | 0.1 (1) | 1.7 (7) | 0.4 (8) |
| Subtotal: Possibly device related | 49 (36) | 2.1 (9) | 0.6 (11) | 6.9 (29) | 2.0 (38) |
| Procedure related/Possibly procedure related |
| Arm/hand numbness/ swelling | 6 (6) | 0.0 (0) | 0.0 (0) | 1.4 (6) | 0.3 (6) |
| Atrial arrhythmias | 6 (6) | 0.5 (2) | 0.1 (2) | 1.0 (4) | 0.2 (4) |
| Coronary Sinus dissection | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Dizziness | 2 (2) | 0.2 (1) | 0.1 (1) | 0.2 (1) | 0.1 (1) |
| Drainage, pocket site | 3 (3) | 0.2 (1) | 0.1 (1) | 0.5 (2) | 0.1 (2) |
| Eccymosis, groin | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Eccymosis, pocket site | 2 (2) | 0.0 (0) | 0.0 (0) | 0.5 (2) | 0.1 (2) |
| Fatigue, tiredness | 3 (3) | 0.2 (1) | 0.1 (1) | 0.5 (2) | 0.1 (2) |
| Fever | 6 (5) | 0.5 (2) | 0.2 (3) | 0.7 (3) | 0.2 (3) |
| ICD migration | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Heart failure decompensation | 5 (5) | 1.2 (5) | 0.3 (5) | 0.0 (0) | 0.0 (0) |
| Hematoma, groin | 2 (2) | 0.0 (0) | 0.0 (0) | 0.5 (2) | 0.1 (2) |
| Hemo/Pneumothorax | 3 (3) | 0.5 (2) | 0.1 (2) | 0.2 (1) | 0.1 (1) |
| Hypotension | 8 (7) | 1.7 (7) | 0.4 (8) | 0.0 (0) | 0.0 (0) |
| Inadequate cardiac output | 1 (1) | 0.2 (1) | 0.1 (1) | 0.0 (0) | 0.0 (0) |
| Junctional rhythm | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Nausea/vomiting | 11 (11) | 1.4 (6) | 0.3 (6) | 1.2 (5) | 0.3 (5) |
| Pain pocket site | 101 (96) | 1.7 (7) | 0.5 (9) | 21.4 (90) | 4.9 (92) |
| Pain (arm, back, chest, groin, head, shoulder) | 62 (61) | 1.0 (4) | 0.1 (2) | 13.5 (57) | 3.1 (58) |
| Pericardial effusion/Pericarditis | 4 (4) | 0.2 (1) | 0.2 (1) | 0.7 (3) | 0.2 (3) |
| Phlebitis | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Pleural effusion | 6 (6) | 0.2 (1) | 0.1 (1) | 1.2 (5) | 0.3 (5) |
| Pocket infection/seroma/ hematoma/swelling | 38 (38) | 2.4 (10) | 0.5 (10) | 6.7 (28) | 1.5 (28) |
| Rash | 3 (3) | 0.2 (1) | 0.1 (1) | 0.5 (2) | 0.1 (2) |
| Renal insufficiency/failure | 4 (4) | 0.7 (3) | 0.2 (3) | 0.2 (1) | 0.1 (1) |
| Respiratory arrest/ failure | 2 (2) | 0.5 (2) | 0.1 (2) | 0.0 (0) | 0.0 (0) |
| Thrombosis | 2 (2) | 0.5 (2) | 0.1 (2) | 0.0 (0) | 0.0 (0) |
| Atrial or ventricular arrhythmias | 6 (6) | 1.0 (4) | 0.2 (4) | 0.5 (2) | 0.1 (2) |
| Subtotal: Procedure related/Possibly procedure related1 | 317 (208) | 13.3 (56) | 4.0 (74) | 42.5 (179) | 13.1 (243) |
| Heart failure decompensation |
| Atrial arrhythmias | 3 (3) | 0.7 (3) | 0.2 (3) | 0.0 (0) | 0.0 (0) |
| Cardiac related2 | 12 (12) | 1.9 (8) | 0.4 (8) | 1.0 (4) | 0.2 (4) |
| Edema | 6 (5) | 0.2 (1) | 0.1 (1) | 1.2 (5) | 0.3 (5) |
| Electrolyte imbalance3 | 3 (3) | 0.2 (1) | 0.1 (1) | 0.5 (2) | 0.1 (2) |
| Elevated liver enzymes | 1 (1) | 0.0 (0) | 0.0 (0) | 0.2 (1) | 0.1 (1) |
| Heart failure decompensation | 127 (88) | 14.3 (60) | 4.4 (82) | 9.7 (41) | 2.4 (45) |
| Pleural effusion | 2 (2) | 0.5 (2) | 0.1 (2) | 0.0 (0) | 0.0 (0) |
| Renal insufficiency/failure; elevated creatinine | 13 (12) | 1.4 (6) | 0.4 (7) | 1.4 (6) | 0.3 (6) |
| Respiratory related4 | 26 (25) | 1.9 (8) | 0.5 (9) | 4.0 (17) | 0.9 (17) |
| Subtotal: Heart failure decompensation1 | 202 (120) | 17.1 (72) | 6.3 (118) | 16.2 (68) | 4.5 (84) |
| Not device or procedure related | 754 (259) | 21.9 (92) | 8.5 (158) | 56.3 (237) | 32.0 (596) |

1The following adverse events were reported in three or fewer patients: ICD related: Pain pocket site, shoulder pain/discomfort. Possibly device related:  Back pain/discomfort, diaphoresis, dyspnea/shortness of breath, palpitations, peripheral edema, renal insufficiency/failure, shoulder pain/discomfort. Procedure related/possibly procedure related: Adrenal insufficiency, anemia, anxiety, bloody sputum, COPD, confusion, decreased oxygen saturation, density on chest x-ray, drainage of old device pocket site, dysphagia, electromechanical dissociation (transient), elevated white blood cell count, eye abrasion, headache, nose bleed, occluded brachiocephalic vein, palpitations, sinus tachycardia, sleep problems, subclavian vein collapse, urinary problems. Heart failure decompensation: Anxiety, decreased appetite, fatigue, heart block, heart transplant, syncope, ventricular arrhythmias, weight gain.

2 Reported cardiac related events: cardiac arrest, cardiogenic shock, cardiomyopathy, chest pain/angina pectoris, inadequate
 	 cardiac output.
3 Reported electrolyte imbalances: hyperkalemia, hyponatremia.
4 Reported respiratory events: Bloody sputum, dyspnea/shortness of breath, COPD, respiratory failure.

A total of 82 deaths were reported in the 434 enrolled NYHA Class III and IV patients. Of these deaths, 59 occurred in 364 patients who were implanted with an InSync® ICD system and randomized to either the control (CRT OFF) or treatment group (CRT ON) Table 2 summarizes the investigator’s death classifications.

Table 2.  Summary of Death Classifications for Randomized NYHA Class III    and IV Patients

| Investigator Classification | Control | Treatment |
| --- | --- | --- |
| Sudden Cardiac | 5 | 4 |
| Non-sudden Cardiac | 17 | 20 |
| Non-Cardiac | 7 | 3 |
| Unknown | 1 | 2 |
| Total | 30 | 29 |

In addition, 23 deaths occurred in patients who were not randomized: 3 deaths in patients in whom an implant was not attempted; 14 deaths in patients in whom an implant attempt was unsuccessful; and 6 deaths in patients who were implanted with an InSync® ICD system but were not randomized. Of these 23 deaths, 4 were classified by the investigator as sudden cardiac, 12 as non-sudden cardiac, 3 as non-cardiac and 4 as unknown.

B.  Possible Adverse Events

Possible adverse events associated with ICD systems with Cardiac Resynchronization Therapy include (in alphabetical order):

- acceleration of arrhythmias (caused by ICD)

- air embolism

- bleeding

- chronic nerve damage

- erosion

- exacerbation of heart failure

- excessive fibrotic tissue growth

- extrusion

- fluid accumulation

- formation of hematomas or cysts

- inappropriate shocks

- infection

- keloid formation

- lead abrasion and discontinuity

- lead migration / dislodgment

- myocardial damage

- pneumothorax

- potential mortality due to inability to defibrillate or pace

- shunting current or insulating myocardium during defibrillation; thromboemboli

- venous occlusion

- venous or cardiac perforation.

The adverse events related to the use of transvenous leads include, but are not limited to, the follow patient-related conditions when the lead is being inserted and/or repositioned

- cardiac dissection

- cardiac perforation

- cardiac tamponade

- coronary sinus dissection

- endocarditis

- fibrillation or other arrhythmias

- heart block

- heart wall or vein wall rupture

- infection, muscle or nerve stimulation

- myocardial irritability

- pericardial effusion

- pericardial rub

- pneumothorax

- thrombolytic and air embolism

- thrombosis

- valve damage (particularly in fragile hearts).

VII.  Alternative Practices and Procedures

The present established therapies for the treatment of heart failure and sudden cardiac death and the associated signs and symptoms include pharmacological therapy, heart transplantation, other approved biventricular ICDs or other surgical procedures.

VIII.  Marketing History

The InSync® ICD device and Attain Models 2187/2188/4189* leads are currently distributed commercially outside the United States.  Specifically, these devices are approved for sale in the European Community, Australia, Canada, and Latin America (Argentina, Uruguay).

Neither the device nor leads have been withdrawn from the market in any country for any reason related to the safety and effectiveness of the device.

- Summary of Pre-clinical Studies InSync® ICD System

A.  Safety and Risk Analysis

A safety and risk analysis of the InSync® ICD system was conducted to identify potential system hazards and to identify appropriate mitigating actions.  This analysis included:

Hazard Analysis

Hazard analysis identifies possible hazards from use of the InSync® ICD system, determines the probability of occurrence and documents mitigating actions to minimize risk.  Hazard analysis was performed on the InSync® ICD system, including review of all hazards and mitigating actions, and the remaining risk was determined to be acceptable.

Failure Modes Effects Analysis (FMEA)

FMEA was performed on the InSync® ICD system to identify and mitigate any potential design, test, or process issues that could adversely influence the safety and/or performance of the device.

Reliability Prediction Analysis

An analysis of the expected field performance on the InSync® ICD was performed based on the field performance of similar Medtronic devices.  The predicted field performance of the InSync® ICD is a failure rate of 0.035% per month at a 90% confidence level.

- Non-clinical Laboratory Studies – InSync® ICD Model 7272

Major Components

The InSync® ICD Model 7272 is a derivation of the GEM DR Model 7271 and GEM II DR Model 7273 systems with biventricular pacing capability added for cardiac resynchronization.  The InSync® ICD Model 7272 components are identical or similar to those used in the GEM DR Model 7271.

Table 3.  Major Components of the InSync® ICD Model 7272

| Component | Comparison to Model 7271 GEM DR | Qualification |
| --- | --- | --- |
| Connector Module | Polyurethane module that contains lead ports. The module is bonded to device can Similar to Model 7271 GEM DR. LV lead port added to the connector module | InSync® ICD Model 7272 Connector Qualification Testing |
| Battery (2) | Same lithium-silver vanadium oxide cells (Li/SVO) developed by the Promeon Division of Medtronic | Qualified with Model 7271 GEM DR |
| High Voltage Capacitors (2) | Two HV output capacitors provide energy for cardioversion and defibrillation therapies. Same as Model 7271 | Qualified with Model 7271 GEM DR |
| Antenna | The antenna sends/receives device communications through bi-directional telemetry Same component used in Model 7271 GEM DR | Qualified with Model 7271 GEM DR |
| Transformer | The transformer converts low voltage from the battery to high voltage for the HV capacitors Same component used in Model 7271 GEM DR | Qualified with Model 7271 GEM DR |
| Reedswitch | The reedswitch is a magnetically controlled mechanical switch that, once closed, signals the device that telemetry communications with the external programmer can occur Same component used in Model 7271 GEM DR | Qualified with Model 7271 GEM DR |
| Biocompatibility | The materials used in the InSync® ICD device are identical to those used in the Medtronic GEM DR Model 7271. These materials include: titanium, silicone adhesive, silicone rubber, and polyurethane. No new materials or processes were introduced with the InSync® ICD Model 7272 that would introduce new issues of biocompatibility. | Qualified with Model 7271 GEM DR |

Connector Module

Qualification activities performed on the InSync® ICD Model 7272 tested the electrical and physical integrity of the connector module, adhesive interfaces, and block / MBC-to-feedthrough welds (qualified by similarity to GEM Model 7227Cx).  The connector design met all connector-specific requirements specified in the InSync® ICD product specification and the IS-1 and DF-1 international standards for connectors (ISO 5841-3 and ISO 11318, respectively).  Qualification demonstrated that the connector module met these requirements with 90% reliability at a 90% confidence level.

Table 4.  Connector Test

| Connector Test | Sample Size | Acceptance Criteria | Result |
| --- | --- | --- | --- |
| Insertion / Withdrawal (Go Gauge & Lead Connector) | 22 | < 2.0 lbf, Go Gauge), < 13.15 lbf, Lead Connector) | Passed Passed |
| IS-1 Electrical Leakage Impedance | 22 | > 50 K ohms | Passed |
| DF-1 Electrical Isolation | 22 | < 50 mA | Passed |
| DF-1 Current Carrying | NA | < 65 V | Qualified by Similarity |
| Suture Hole Pull Force | 22 | > 6 lbs | Passed |

Device Qualification

Device qualification testing was performed to ensure that the InSync® ICD Model 7272 performs acceptably in typical shipping, handling and operating environments.  The device qualification testing is summarized below.  The results demonstrated that the InSync® ICD Model 7272 will perform acceptably in typical shipping, handling, and operating environments and is qualified for its intended use.

Table 5.  Device Qualification Testing

| Test | Sample Size | Acceptance Criteria | Results |
| --- | --- | --- | --- |
| Environmental | 10 | Temperature Storage: Meets Section 26.2 of European Standard prEN45502-2-2 Mechanical Vibration: Meets Section 23.2 of European Standard prEN45502-2-2. Mechanical Shock: Show no visible signs of damage which affects function of device after a shock having a change in velocity (dV) of 118 inches per second and a duration of 1ms in each of six axes. Simulates a drop of 45cm (18”) to a hard surface. | Meets Acceptance Criteria |
| Electromagnetic Compatibility | 22 | Electromagnetic Interference: Meets requirements of the 1975 AAMI Pacemaker Standard. Also meets requirements at additional frequencies, including radiated continuous wave and pulsed electromagnetic fields and conducted continuous wave sinusoidal currents. | Meets Acceptance Criteria |
|  | 3 | Cellular Phone: Not susceptible to interference from analog or digital cellular telephones at distances >=8 cm, including the following systems: AMPS, TDMA-50 (NADC), GSM, DCS, and CDMA. |  |
|  | 22 | X-ray: Testing waived based on similarity to GEM DR Model 7271. |  |
|  | 22 | Electrosurgical Cautery: Must withstand spark cutting, spark coagulating, and sine cutting modes and energies. |  |
|  | 22 | Transthoracic Defibrillation: must withstand 1000V and 1500V. |  |
| Design Verification Testing | 3 | Because of hardware similarity between InSync® ICD and GEM DR, verification testing was conducted only on hardware components (tests on the CRT therapy, unmodified circuits and device immunity to noise and ESD). Testing confirmed conformance with product specifications. | Meets Acceptance Criteria |

C.  Non-clinical Laboratory Testing – Firmware/Software

1.  Firmware Verification Testing

Firmware verification testing functionally verifies that each firmware requirement for the Model 7272 InSync® ICD is met.  The verification test plan is developed in parallel with the firmware design and firmware implementation phases.

Table 6.  Summary of Firmware Verification Testing

| Summary of Firmware Verification Testing | Sample Tested | Result |
| --- | --- | --- |
| Firmware verification testing was performed as defined by the Verification Test Plan Rev. 003 | Firmware | Passed |

2.  Software Verification Testing

Software verification testing was performed on the Model 9969 Application Software on both the Model 9790 and 2090 Programmers.

Table 7.  Summary of Software Verification Testing

| Summary of Software Verification Testing | Sample Tested | Result |
| --- | --- | --- |
| Software verification testing was performed as defined by the Verification Test Plan. Testing consisted of functional testing, Software Install testing, stress testing, and regression testing. | Software | Passed |

3.  System Testing

System testing of the InSync® ICD system (InSync® ICD Model 7272, Model 9969 Software, Model 9790 and 2090 programmers, and accessories) was performed to ensure that all system components work together appropriately under simulated clinical situations.

Table 8.  Summary of System Testing

| Summary of System Testing | Sample Tested | Result |
| --- | --- | --- |
| System testing evaluates the compatibility, interaction and functional operation of the InSync® ICD system components when used together as a system. | InSync® ICD System Components | Passed |

D.  Biocompatibility Testing

All blood-contacting materials used in the InSync® ICD pulse generator are identical to those used in PMA-approved Medtronic pulse generators. These currently-marketed pulse generators have previously been subject to standard biocompatibility evaluations which conform to FDA Tripartite Biocompatibility Guidance (4/87), FDA Blue Book Memorandum #95-1 (5/95), and the International Standard ISO-10993.

All biocompatibility testing was conducted in compliance with Good Laboratory Practices (GLP).

E.  Sterilization Information

The 100% ethylene oxide (EtO) sterilization process used to sterilize the implantable pulse generator has been previously approved in other PMA applications.  All processes used to sterilize a product are validated and qualified according to the major guidelines and standards. Package qualification testing was performed on both the device to ensure suitability for their intended purpose.

F.  Animal Studies

GLP studies were conducted in 3 canines with the InSync® ICD Model 7272.  The purpose of the Model 7272 InSync® ICD study was to determine that the device provided acceptable performance for sensing and detection of cardiac rhythms, appropriate rejection of sinus arrhythmias, and correct delivery of appropriate programmed therapy.  In each of the studies, appropriate pacing, sensing, and thresholds were documented.

G.  Conclusion Concerning Non-clinical Laboratory Testing

Medtronic conducted a hazard analysis on all new features and critical components and then conducted testing to evaluate these and other device features.  All test results were found to be acceptable.

H.  Shelf Life

Based on battery longevity testing, the shelf life for the InSync® ICD pulse generator was established and approved at 18 months from battery attachment date.

X.  Summary of Clinical Study

A.  Study Design

The InSync® ICD study was a prospective, multi-center, randomized, double-blind, parallel arm, controlled clinical trial to assess the safety and effectiveness of transvenous atrial-based synchronous biventricular pacing for heart failure therapy (cardiac resynchronization therapy – CRT) in patients who are indicated for an ICD. Patients in NYHA Classes II, III and IV were enrolled in the clinical study (for the purposes of FDA approval, only data from NYHA Class III and IV patients are presented unless otherwise noted).  The products being evaluated include the Model 7272 InSync® ICD and the Model 9969 application software for use with the Model 9790C and 2090 programmers and the commercially available Models 2187 and 2188 Attain LV leads and the Model 4189* Attain LV Lead.

Note:  The 4189 Attain Lead is not currently market approved in the U.S.  Lead performance data in the clinical trial did not support reasonable assurance of the safety and effectiveness of the 4189 Lead.  However, the data obtained with the 4189 Lead are presented for completeness here.  Based on similarities in the design and function of the 4189 and 2187/2188, these two sets of leads are equally compatible with the Model 7272 InSync® ICD to deliver biventricular pacing.

Figure 1 depicts the InSync® ICD study design. Patients were randomized within 7 days of a successful implant and were in double-blind follow-up for 6 months. After the six-month follow up, all patients had CRT turned on.

Figure 1.  InSync® ICD Study Design

Patient Selection Criteria

Inclusion Criteria

ICD Indication

- Patients who have had at least one episode of cardiac arrest (manifested by loss of consciousness) due to ventricular tachyarrhythmia;

or

- Patients having recurrent, poorly tolerated, sustained VT that occurs spontaneously or can be induced;

or

- Canada Only:  Patients who have had a prior myocardial infarction, left ventricular ejection fraction less than or equal to 35%, a documented episode of non-sustained VT and inducible ventricular tachyarrhythmia.

And Heart Failure as defined by:

- Patients having symptomatic congestive heart failure (NYHA Class II, III, IV) with evidence of ventricular dysynchrony demonstrated by:

- QRS duration ( 130 ms,

- LV Ejection Fraction ( 35%,

- LV End Diastolic Diameter ( 55 mm.

- Stable medical regimen for at least 1 month prior to enrollment (to include at least ACE inhibitors, or ACE inhibitor substitute) patients currently on stable beta blockade regimes for at least 3 months prior to enrollment were allowed in the study.  Patients may not be started on beta-blockers during the 6-month randomization period.

- 18 years of age or older.

Exclusion Criteria

- Baseline Six Minute Hall Walk distance greater than 450 meters (Class III and IV only).

- Unstable angina, or acute MI, CABG, or PTCA within the past 3 months.

- Recent CVA or TIA (within the previous 3 months).

- Intermittent positive inotropic drug therapy (intermittent is defined as more than two infusions per week).

- Patients having indications or contraindications for standard cardiac pacing.

- Systolic blood pressure of less than 80 mm or greater than 170 mm.

- Resting heart rate greater than 140 bpm.

- Serum Creatinine greater than 3 mg/dl.

- Hepatic function (serum) greater than 3 times the upper limit of normal.

- Primary valvular disease.

- Severe primary pulmonary disease.

- Chronic atrial arrhythmias (or cardioversion for afib within the past month) or paroxysmal atrial fib event within the previous month.

- Post heart transplant (patients who are waiting for a heart transplant are allowed entry into the study).

- Enrolled in any concurrent study that may confound the results of this study.

- Patients who are not expected to survive for 6 months of study participation.

- Women who are pregnant or with child bearing potential and who are not on a reliable form of birth control.

- Patients with mechanical right heart valves.

- Patients with VT associated with reversible causes.

Study Blind

The InSync® ICD trial was double-blinded to reduce the effect of bias. Implanting physicians and staff responsible for programming the InSync® ICD were required to be aware of the mode to which the device was programmed.  The EP staff (electrophysiologist, EP nurses, etc.) performed all tests that required viewing of the programmed pacing mode or the patient’s ECG.  The heart failure staff (heart failure physician, nurses, etc.) and the patients remained blinded and did not view these materials.  Neither the patient nor the heart failure physician had knowledge of the pacing mode assigned.  Because of their blinded status, the evaluations of QOL, NYHA classification, and 6-minute hall walk were conducted by the heart failure staff.  All non-electrophysiological patient management decisions were made by the blinded heart failure physicians and nurses.

B.  Clinical Study Results

1.  Patient Population

Study Enrollment

The first successful implant of the Model 7272 InSync® ICD cardiac resynchronization system was performed on October 4, 1999. Fifty-three centers provided data for this clinical report (48 in USA, 5 in Canada) and 638 patients (Class II, III and IV) had implant attempts. There were 570 patients with successful implants, of which 555 were randomized. The patients received the commercially available LV Lead Models 2187 or 2188 or the Model 4189* LV Lead which is not a subject of this PMA approval.  In total, there were 283 patients randomized to the control arm (OFF), and 272 patients randomized to the treatment arm (ON).  Of the 283 control patients, 106 were NYHA Class II and 177 were NYHA Class III and IV. Of the 272 treatment patients, 85 were NYHA Class II and 187 were NYHA Class III and IV treatment patients. For the purposes of FDA approval, results from NYHA Class III and IV patients only will be presented unless otherwise noted.

Figure 2 provides summary information on the study patient population.

1 Refer to page 20 for additional information.
2 Refer to page 33 for detailed information.
3 Refer to page 20 for detailed information.

Figure 2. InSync® ICD Study Patient Population

Patients Enrolled But Implant Not Attempted

Of the 434 NYHA Class III and IV patients enrolled in the InSync® ICD study, 10 NYHA Class III and IV patients were consented, entailing enrollment, but an implant was not attempted.  Table 9 summarizes the reasons that an implant was not attempted in these patients.

Table 9. Patients Enrolled But Implant Not Attempted – 
NYHA Class III and IV Patients

| Reason implant was not attempted | Total |
| --- | --- |
| Did not satisfy VT inclusion criterion | 2 |
| Did not satisfy LVEF inclusion criterion | 2 |
| Unable to obtain venous access | 2 |
| Death prior to implant attempt | 2 |
| Developed indications for standard cardiac pacing | 1 |
| Cancer diagnosed; chemotherapy initiated | 1 |

Patients Not Randomized

Reasons for ten NYHA Class III and IV patients who were implanted with an InSync® ICD system but not randomized are summarized below:

- Patient unstable/pacing required (4)

- LV lead dislodgement (3)

- Respiratory arrest (1)

- Heart transplant prior to randomization (1)

- Patient death prior to randomization (1)

Crossovers

Biventricular Pacing OFF to ON

Fifteen NYHA Class III and IV patients were randomized to biventricular pacing OFF and had an early crossover to biventricular pacing ON prior to the 6-month follow-up visit. The reasons for the programming changes are described below:

- Worsening heart failure requiring CRT (11)

- Inadvertent programming (2)

- Bradycardia (1)

- AV nodal ablation (1)

Biventricular Pacing ON to OFF

Ten NYHA Class III and IV patients were randomized to biventricular pacing ON and had an early crossover to biventricular pacing OFF prior to the 6-month follow-up visit. The reasons for the programming changes are described below:

- Inadvertent programming (6)

- LV lead dislodgment (2)

- Diaphragmatic stimulation (2)

2.  Baseline Demographic Data

Baseline demographic data for randomized NYHA Class III and IV patients are summarized in Table 10.  P-values < 0.05 are noted with an asterisk.

Table 10.  Comparison of Control and Treatment Patients at Baseline –
NYHA Class III and IV Patients

| InSync® ICD Comparison of Demographics by Treatment Assignment All Randomized NYHA Class III/IV Patients |
| --- |
|  | Control (OFF) (n=177) | Treatment (ON) (n=187) | P-value |
| Gender (n,%) Male Female | 136 (76.8%) 41 (23.2%) | 142 (75.9%) 45 (24.1%) | 0.90 |
| Age (years) Mean ( Standard deviation n(%) | 67.4 ( 9.4 177 (100.0%) | 66.6 ( 11.3 187 (100.0%) | 0.88 |
| New York Heart Classification (n,%) III IV | 158 (89.3%) 19 (10.7%) | 165 (88.2%) 22 (11.8%) | 0.87 |
| QRS Width (ms) Mean ( Standard deviation n (%) | 162 ( 22 177 (100.0%) | 165 ( 22 187 (100.0%) | 0.15 |
| Ejection Fraction (%) Mean ( Standard deviation n (%) | 20.1 ( 6.1 176 (99.4%) | 20.5 ( 7.0 187 (100.0%) | 0.74 |
| LVEDD (mm) Mean ( Standard deviation n (%) | 70.9 ( 9.1 174 (98.3%) | 70.4 ( 9.0 186 (99.5%) | 0.51 |
| MN Living with HF Score Mean ( Standard deviation n (%) | 55.2 ( 22.9 174 (98.3%) | 56.5 ( 22.9 185 (98.9%) | 0.56 |
| 6 –Minute Hall Walk (meters) Mean ( Standard deviation n (%) | 245 ( 118 175 (98.9%) | 243 ( 130 183 (97.9%) | 0.87 |
| Peak VO2/kg (ml/kg/min.) Mean ( Standard deviation n (%) | 13.5 ( 3.9 135 (76.3%) | 13.4 ( 3.6 144 (77.0%) | 0.96 |
| Exercise Time (seconds) Mean ( Standard deviation n (%) | 511 ( 229 136 (76.8%) | 473 ( 213 143 (76.5%) | 0.27 |
| Supine Heart Rate (bpm) Mean ( Standard deviation n (%) | 71.5 ( 12.9 171 (96.6%) | 70.9 ( 12.4 186 (99.5%) | 0.82 |
| Baseline CV Medical History (n, %) (non-exclusive) n recorded Coronary Artery Disease Congenital Heart Disease HF etiology: ischemic HF etiology: non-ischemic Myocardial Infarction Anterior Lateral Posterior Inferior Non Q-wave Cardiomyopathy Primary electrical disease Valvular Disease Hypertension Chronotropic incompetence Syncope/presyncope Other | 177 135 (76.3%) 0 (0%) 133 (75.1%) 49 (27.7%) 112 (63.3%) 46 (26.0%) 13 (7.3%) 3 (1.7%) 32 (18.1%) 15 (8.5%) 166 (93.8%) 5 (2.8%) 0 (0%) 90 (50.8%) 2 (1.1%) 46 (26.0%) 28 (15.8%) | 186 122 (65.6%) 2 (1.1%) 119 (64.0%) 67 (36.0%) 100 (53.8%) 48 (25.8%) 16 (8.6%) 5 (2.7%) 31 (16.7%) 6 (3.2%) 176 (94.6%) 6 (3.2%) 0 (0%) 84 (45.2%) 0 (0%) 49 (26.3%) 31 (16.7%) | 0.03* 0.50 0.02* 0.09 0.07 1.00 0.70 0.72 0.78 0.04* 0.82 1.00 1.00 0.29 0.24 1.00 0.89 |
| Prior Cardiac Intervention (n, %) (non-exclusive) n recorded CABG Valve replacement Ablation Coronary Artery Intervention Repair/Correction of Cong. Abnormality ICD currently implanted None | 173 87 (50.3%) 9 (5.2%) 7 (4.0%) 48 (27.7%) 0 (0%) 51 (29.5%) 48 (27.7%) | 182 82 (45.1%) 13 (7.1%) 6 (3.3%) 44 (24.2%) 2 (1.1%) 54 (29.7%) 55 (30.2%) | 0.34 0.51 0.78 0.47 0.50 1.00 0.64 |
| Spontaneous Arrhythmia History (non-exclusive) n recorded Atrial Fibrillation Atrial Flutter AV Nodal Reentrant Tachycardia Sustained Monomorphic VT Sustained Polymorphic VT Nonsustained VT Ventricular Flutter/Fibrillation Torsades de Pointes Long QT Syndrome Bradycardia1 Right Bundle Branch Block Left Bundle Branch Block Others | 177 31 (17.5%) 7 (4.0%) 0 (0%) 67 (37.9%) 10 (5.6%) 52 (29.4%) 28 (15.8%) 1 (0.6%) 0 (0%) 101 (57.1%) 24 (13.6%) 126 (71.2%) 32 (18.1%) | 187 49 (26.2%) 10 (5.3%) 0 (0%) 64 (34.2%) 7 (3.7%) 73 (39.0%) 26 (13.9%) 1 (0.5%) 0 (0%) 105 (56.2%) 24 (12.8%) 141 (75.4%) 39 (20.9%) | 0.06 0.62 0.62 0.51 0.46 0.06 0.66 1.00 1.00 0.92 0.88 0.41 0.51 |
| Other Medical History (n, %) (non-exclusive) n recorded Endocrine (includes diabetes) Pulmonary (includes COPD) Renal | 177 32 (18.1%) 37 (20.9%) 33 (18.6%) | 187 24 (12.8%) 49 (26.2%) 28 (15.0%) | 0.19 0.27 0.40 |
| Medications (n, %) (non-exclusive) n recorded ACE Inhibitor Anti-Arrhythmic Anti-Depressant Anti-Coagulant Beta Blocker Ca++ Blocker Diuretic Positive Inotrope Nitrate Other Medication | 177 155 (87.6%) 57 (32.2%) 35 (19.8%) 143 (80.8%) 103 (58.2%) 10 (5.6%) 167 (94.4%) 129 (72.9%) 59 (33.3%) 108 (61.0%) | 187 172 (92.0%) 79 (42.2%) 36 (19.3%) 144 (77.0%) 118 (63.1%) 13 (7.0%) 173 (92.5%) 142 (75.9%) 68 (36.4%) 108 (57.8%) | 0.17 0.05 1.00 0.44 0.39 0.67 0.53 0.55 0.58 0.59 |

* p < 0.50

1 Includes sinus bradycardia, sick sinus syndrome, 1º, 2º and 3º atrioventricular block and bradycardia-tachycardia syndrome.

- Primary Effectiveness Endpoint Results

The primary endpoints of the study were improvement in NYHA Class, Quality of Life and 6-minute Hall Walk.  The study results were analyzed using the Hochberg criteria in that there would be a statistical success for effectiveness if one of the three endpoints was statistically significant at alpha =0.0167, if any two were significant at alpha = 0.025, or if all three were significant at alpha = 0.05.

Quality of Life

Patient quality of life (QOL) was assessed with the Minnesota Living with Heart Failure (MLWHF) questionnaire. A lower score indicates an improvement in quality of life. Both the patient and study personnel administering the questionnaire were blinded to the randomization assignment. Table 11 summarizes the paired baseline and 6 month QOL results for NYHA Class III and IV patients.  P-values were based on the paired median difference.

Table 11. MLWHF QOL Paired Baseline and 6-Month Data 
NYHA III/IV Patients

|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P- value |
| --- | --- | --- | --- |
| 6-month change in QOL score | 153 patients with paired data Baseline: median 57.0 Mean 55.1 + 22.5 6-month: median 43.0 mean 42.0 + 22.9 Median Paired Difference: – 11.0 | 162 patients with paired data Baseline: median 55.0 mean 54.2 + 23.1 6-month: median 34.0 mean 35.9 + 23.4 Median Paired Difference: – 17.5 | 0.02 |

Figure 3 presents the median change in QOL for the control and treatment groups at baseline through 6 month follow up.

Figure 3. MN Living With Heart Failure Data Summary

NYHA Classification

NYHA functional classification was determined by a physician who was blinded to the patient’s randomization assignment.  Table 12 summarizes the paired baseline and 6 month NYHA results for NYHA Class III and IV patients. The P-value was based on the median paired difference.

Table 12. NYHA Paired Baseline and 6-Month Data– NYHA III/IV Patients

|  | Control Group (OFF) | Treatment Group (ON) | P- value |
| --- | --- | --- | --- |
| 6-month change in NYHA Classification | 158 patients with paired data Baseline: median 3.0 Mean 3.09 + 0.29 6-month: median 3.0 mean 2.51 + 0.6 Median Paired Difference: – 0.5 | 165 patients with paired data Baseline: median 3.0 mean 3.10 + 0.30 6-month: median 2.0 mean 2.31 + 0.72 Median Paired Difference: – 1 | 0.01 |

Figure 4 presents the median change in NYHA class for the control and treatment groups at baseline through 6 month follow up.

Figure 4. NYHA Classification Data Summary

6-Minute Hall Walk

Patients performed a 6-minute hall walk test in order to assess sub-maximal exercise capacity. Study personnel administering the test were blinded to the patient’s randomization assignment.  Table 13 summarizes the paired baseline and 6 month 6 minute hall walk results for NYHA Class III and IV patients. The P-value was based on the median paired difference.

Table 13. 6-Minute Hall Walk Paired Baseline and 6 Month Data– 
NYHA III/IV Patients

|  | Control Group (OFF) | Treatment Group (ON) | P- value |
| --- | --- | --- | --- |
| 6-month change in 6 minute hall walk (meters) | 150 patients with paired data Baseline: median 275 Mean 254 + 112 6-month: median 334 mean 324 + 114 Median Paired Difference: 53 | 153 patients with paired data Baseline: median 254 mean 252 + 128 6-month: median 340 mean 339 + 106 Median Paired Difference: 55 | 0.43 |

Figure 5 presents the median change in 6-minute hall walk distance for the control and treatment groups at baseline through 6 month follow up.

Figure 5. 6-Minute Hall Walk Data Summary

4.  Primary Safety Endpoint Results

ICD-Related Complications

Table 14 summarizes the 5 NYHA Class III and IV patients who experienced 7 complications at 3 months post-implant related to the Model 7272 InSync® ICD.

Table 14. InSync® ICD-Related Complications – NYHA III/IV Patients

| Event Type | Number of events | Events in the First 3 Months |
| --- | --- | --- |
| Pain at pocket site | 1 | 1 |
| Pocket seroma/hematoma | 2 | 2 |
| Abnormal impedance measurement | 1 | 1 |
| Dizziness | 1 | 1 |
| Electrical reset of ICD | 1 | 1 |
| Pocket infection | 1 | 1 |
| Pacemaker-mediated tachycardia* | 1 | 0 |
| TOTAL | 8 events (in 6 patients) | 7 events (in 5 patients) |

*after 3 months

Table 15 summarizes the freedom from ICD-related complications at three months post-implant.

Table 15. Freedom from InSync® ICD-Related Complications
 During First 3 Months – NYHA III/IV Patients

| # Patients With ICD-Related Complications | # Implanted Patients | Estimated survival | Lower 95% confidence limit | 95% lower bound criterion |
| --- | --- | --- | --- | --- |
| 5 | 371 | 98.6% | 97.6% | 89.0% |

Attain Model 4189 LV Lead-Related Complications

Table 16 summarizes Model 4189* LV lead-related complications in NYHA Class III and IV patients at six months.  The Model 4189* lead is not the subject of this approval decision.

Table 16. Model 4189-Related Complications  – 
NYHA III/IV Patients

| Event Type | Number Of Events In All Patients | Events In All Patients In The First 6 Months |
| --- | --- | --- |
| Lead dislodgement | 31 | 29 |
| Extra cardiac stimulation | 12 | 11 |
| Elevated pacing thresholds | 4 | 4 |
| Failure to capture | 3 | 3 |
| Muscle stimulation – pectoral | 1 | 1 |
| Pericardial effusion | 1 | 1 |
| TOTAL | 52 events (in 46 patients) | 49 events (in 44 patients) |

Table 17. Freedom from Model 4189
Complications During First 6 Months – NYHA III/IV Patients

| Number of Patients With Events | Number of Implanted Patients | Estimated Survival | Lower 95% Confidence Bound | 95% Lower Bound Criterion |
| --- | --- | --- | --- | --- |
| 44 | 315 | 85.1% | 81.7% | 75% |

5.  Attain Models 2187 and 2188 LV Lead-Related Complications

The following table presents the summary of the market approved Model 2187 and 2188 complications in NYHA Class III and IV patients.

Table 18. Model 2187/2188 Related Complications 
NYHA III/IV Patients

| Event Type | Number of Events | Events in the First 6 Months |
| --- | --- | --- |
| Lead dislodgment | 6 | 3 |
| Failure to capture | 1 | 1 |
| Elevated pacing thresholds | 1* | 1* |
| TOTAL | 8 (in 8 patients) | 5 events (in 5 patients) |

*Related to Model 2188

Table 19. Freedom from 2187/2188 Related Complications 
During First 6 Months – NYHA III/IV Patients

| Number of patients with complications | Number of Implanted Patients | Estimated survival | Lower 95% confidence limit | 95% Lower Bound Criterion |
| --- | --- | --- | --- | --- |
| 5 | 56 | 89.9% | 82.9% | 75% |

System-Related Complications

Table 20 summarizes the system-related complications for NYHA Class III and IV patients.

Table 20. System-Related Complications – NYHA III/IV Patients

| Device Relatedness | Event Type | # Events | # Events at 6 mo. |
| --- | --- | --- | --- |
| InSync® ICD related | Pain at pocket site | 1 | 1 |
|  | Abnormal impedance measurement | 1 | 1 |
|  | Pocket seroma/hematoma | 2 | 2 |
|  | Pocket infection | 1 | 1 |
|  | Dizziness | 1 | 1 |
|  | Electrical reset of ICD | 1 | 1 |
|  | Pacemaker mediated tachycardia | 1 | 1 |
|  | Subtotal: | 8 | 8 |
| LV lead related (2187/88) | Lead dislodgment | 6 | 3 |
|  | Failure to capture | 1 | 1 |
|  | Elevated pacing thresholds | 1 | 1 |
|  | Subtotal: | 8 | 5 |
| LV lead related (4189) | Lead dislodgment | 31 | 29 |
|  | Muscle stimulation – diaphragm | 12 | 11 |
|  | Elevated pacing thresholds | 4 | 4 |
|  | Failure to capture | 3 | 3 |
|  | Muscle stimulation – pectoral | 1 | 1 |
|  | Pericardial effusion | 1 | 1 |
|  | Subtotal: | 52 | 49 |
| RA lead related | Lead dislodgment | 6 | 6 |
|  | Elevated pacing thresholds | 2 | 2 |
|  | Failure to capture | 2 | 2 |
|  | Far-field R-wave sensing | 1 | 1 |
|  | Subtotal: | 11 | 11 |
| RV lead related | Muscle stimulation – diaphragm | 1 | 1 |
|  | Elevated pacing thresholds | 1 | 1 |
|  | Failure to capture | 1 | 1 |
|  | Chronic RV lead discovered implanted in cardiac vein | 1 | 1 |
|  | Subtotal: | 4 | 4 |
| System related | Pocket infection | 1 | 1 |
|  | Twiddler’s syndrome | 1 | 1 |
|  | Subtotal: | 2 | 2 |
| Total |  | 85 events (in 69 pts.) | 79 events (in 65 pts.) |

Table 21 summarizes the freedom from system-related complications at six months post-implant.

Table 21. Freedom from System-Related Complications During First 6 Months – 
NYHA III/IV Patients

| # Patients With System-Related Complications | # Implanted Patients | Estimated survival | Lower 95% confidence limit | 95% lower bound criterion |
| --- | --- | --- | --- | --- |
| 65 | 371 | 81.1% | 77.6% | 67.0% |

6.  Primary LV Lead Effectiveness Endpoint Results

LV Lead Implant Success

Four hundred and twenty-four NYHA Class III and IV patients underwent implant attempts. 374 patients were successfully implanted with an InSync® ICD system (85% received a Model 4189*, 14% received a Model 2187 and 1% received a Model 2188). Three hundred and sixty-five patients were successfully implanted on their first attempt. Eleven of the unsuccessful patients underwent a second implant attempt. Nine of those patients were successfully implanted the second time.

Table 22. Implant Success Rate – NYHA III/IV Patients

| Number of Patients With Implant Attempts | Number of Successful Implants | Successful Implant Rate | Two-sided 95% Confidence Interval | 95% Lower Bound Criterion |
| --- | --- | --- | --- | --- |
| 424 | 374 | 88.2% | 84.8%-91.1% | 83% |

Table 23 summarizes the reasons for the unsuccessful implants (patients may have more than one reason).

Table 23. Summary of Unsuccessful Implants 
(Model 2187, 2188 and 4189 LV Leads) – NYHA III/IV Patients

| Reason For Unsuccessful Implant | Number (not mutually exclusive) |
| --- | --- |
| Dislodgement/unstable position | 31 |
| Unable to obtain adequate distal location | 26 |
| Unable to cannulate coronary sinus ostium | 23 |
| Unacceptable pacing thresholds | 21 |
| Dissection/perforation | 17 |
| Unable to access coronary vein | 12 |
| Coronary vein too small | 4 |
| Patient decompensation during implant | 4 |
| Delivery system/tool problems | 4 |
| Patient venous anatomy | 3 |
| Diaphragmatic stimulation | 2 |
| Complete heart block | 1 |

Model 4189 LV Lead Voltage Thresholds

For all patients receiving a Model 4189* LV lead, the mean 6-month voltage threshold measured at 0.5 ms was 1.5V; the two-sided 95% confidence interval was (1.4V, 1.7V). The following table summarizes the voltage thresholds measured at implant, 1-month, 3-month, and 6-month follow-up visits, with two-sided 95% confidence intervals.

Figure 6. Model 4189 Left Ventricular Voltage Threshold Data (@ 0.5 ms)

Models 2187 and 2188 LV Lead Voltage Thresholds

For all patients receiving a Model 2187 or 2188 lead, the mean 6-month voltage threshold measured at 0.5 ms was 1.9V; the two-sided 95% confidence interval was (1.7V, 2.2V). The following table summarizes the voltage thresholds measured at implant, 1-month, 3-month, and 6-month follow-up visits, with two-sided 95% confidence intervals.

Figure 7. Model 2187/2188 Left Ventricular Voltage Threshold Data (@ 0.5 ms)

7.  Secondary Endpoint Results

Patient Survival

Table 24 summarizes survival from all-cause mortality and the number of patient deaths at 3 and 6 months in randomized NYHA Class III and IV patients, comparing control and treatment groups. Survival from all-cause mortality at 
6 months for the treatment group is 92.4%, and for the control group it is 92.0%. There is no difference in survival from all-cause mortality between treatment and control (p=0.72).

Table 24. Survival Estimates for Control vs. Treatment Groups – 
NYHA III/IV Patients

|  | Estimated Survival – 3 Months | Estimated Survival – 6 Months |
| --- | --- | --- |
|  | # of Deaths | Patients At Risk | Estimated Survival | 95% Confidence Interval | # of Deaths | Patients At Risk | Estimated Survival | 95% Confidence Interval |
| Control Group (n=177) | 6 | 171 | 96.6% | 92.6%-98.5% | 14* | 87 | 92.0% | 86.8%-95.2% |
| Treatment Group (n=187) | 7 | 178 | 96.2% | 92.3%-98.2% | 14 | 99 | 92.4% | 87.5%-95.4% |

*One additional patient died after six months, but before their six-month follow-up

Figure 8 depicts all-cause survival for control and treatment groups.

Figure 8. Survival From All-cause Mortality 
During Randomization Period – NYHA Class III and IV

Table 25 summarizes survival from all-cause mortality at 3, 6, 12, and 18 months and number of deaths for all NYHA Class III and IV patients receiving an InSync® ICD system (n=374). Note that at the conclusion of the 6-month follow-up visit, all patients had CRT turned on.

Table 25. Estimated Survival From All Causes of Death – 
NYHA III/IV Patients

| Time Post-implant | Number of Deaths | Estimated Survival | 95% Confidence Interval |
| --- | --- | --- | --- |
| 3 Months | 18 | 95.2% | 92.4%-96.9% |
| 6 Months | 34 | 90.8% | 87.3%-93.3% |
| 12 Months | 52 | 84.3% | 79.8%-87.8% |
| 18 Months | 62 | 76.2% | 69.5%-81.7% |

Figure 9 depicts all-cause survival curve through 18 months for NYHA Class III and IV patients.

Figure 9. Survival From All-cause Mortality, 
NYHA Class III and IV

Cardiopulmonary Exercise

Peak VO2, an indicator of a patient’s exercise capacity, was the secondary effectiveness endpoint specifically identified for NYHA Class III and IV patients. Table 26 and Figure 10 summarize the peak VO2 data. P-values were based on the median paired difference.

Cardiopulmonary exercise testing was performed on a treadmill using a modified Naughton protocol.  Data were analyzed by an independent core lab. Table 27 summarizes the paired baseline and 6-month cardiopulmonary exercise results for NYHA Class III and IV patients and additional cardiopulmonary exercise data collected in the study. P-values were based on the median paired difference.

Table 26. Peak VO2  – NYHA III/IV Patients

|  | Cardiopulmonary Exercise Testing - Results for NYHA Class III and IV Patients |
| --- | --- |
|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P-value |
| Peak VO2 (ml/kg/min) | 118 patients with paired data Baseline: median 13.4 Mean 13.7 + 3.9 6-month: median 13.5 mean 14.0 + 4.0 Median Paired Difference: 0.0 | 118 patients with paired data Baseline: median 13.4 mean 13.5 + 3.5 6-month: median 14.3 mean 14.4 + 3.7 Median Paired Difference: 1.1 | 0.03 |

Figure 10 presents the median change in Peak VO2 for the control and treatment groups at baseline and 6 month follow up.

Figure 10. Peak VO2 Data Summary

Table 27. Cardiopulmonary Exercise Data – NYHA III/IV Patients

|  | Cardiopulmonary Exercise Testing - Results for NYHA Class III and IV Patients |
| --- | --- |
|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P-value |
| Exercise Time (sec) | 119 patients with paired data Baseline: median 494 Mean 518 + 226 6-month: median 435 mean 483+ 265 Median Paired Difference: -9 | 118 patients with paired data Baseline: median 479 mean 494 + 479 6-month: median 526 mean 537+ 219 Median Paired Difference: 56 | 0.0008 |
| Respiratory Exchange Ratio (RER) Max | 112 patients with paired data Baseline: median 1.14 Mean 1.13 + 0.10 6-month: median 1.09 mean 1.07 +0.12 Median Paired Difference: -0.06 | 116 patients with paired data Baseline: median 1.14 mean 1.13+ 0.10 6-month: median 1.12 mean 1.11 + 0.10 Median Paired Difference: -0.02 | 0.01 |
| Ventilatory Exchange (VE) | 112 patients with paired data Baseline: median 47.3 Mean 48.3 + 15.2 6-month: median 44.7 mean 47.6 + 16.4 Median Paired Difference: -1.0 | 115 patients with paired data Baseline: median 47.8 mean 49.6 + 17.9 6-month: median 48.5 mean 51.0 + 16.1 Median Paired Difference: 0.9 | 0.13 |
| VE/VCO2 | 112 patients with paired data Baseline: median 41.0 Mean 42.6 + 10.4 6-month: median 40.9 mean 42.2 + 10.6 Median Paired Difference: -0.6 | 115 patients with paired data Baseline: median 39.8 mean 41.8 + 10.4 6-month: median 38.0 mean 39.9 + 8.9 Median Paired Difference: -1.0 | 0.19 |
| Anaerobic Threshold (AT) | 68 patients with paired data Baseline: median 8.90 Mean 9.11 + 2.56 6-month: median 9.65 mean 10.30 + 3.22 Median Paired Difference: 0.90 | 87 patients with paired data Baseline: median 9.80 mean 9.61 + 2.20 6-month: median 10.30 mean 10.24 + 2.02 Median Paired Difference: 0.70 | 0.41 |

QRS Duration

The change in patients’ QRS duration from baseline to 6 months was analyzed to assess the effects of cardiac resynchronization therapy. Table 28 summarizes the paired baseline and 6-month peak QRS results for NYHA Class III and IV patients. The P-values were based on the median paired difference.

Table 28.  QRS Duration – NYHA III/IV Patients

|  | QRS Duration – Results for NYHA Class III and IV Patients |
| --- | --- |
|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P-value |
| Change in QRS duration from baseline through 6-months (ms) | 119 patients with paired data Baseline: median 160 Mean 162 + 22 6-month: median 160 mean 157 + 37 Median Paired Difference: –6 | 129 patients with paired data Baseline: median 160 mean 163 + 20 6-month: median 150 mean 146 + 28 Median Paired Difference: –20 | 0.01 |

Health Care Utilization (Hospitalization)

Emergency room, outpatient, clinic and hospital admission data ware collected in the study. An independent Adverse Events Review Committee reviewed the hospitalization data and adjudicated hospitalizations as either related or not related to heart failure. The analysis included all admissions of at least 24 hours. Table 29 summarizes the hospitalization (all-cause and heart failure (CHF) related) results for NYHA Class III and IV patients.  The P-values were based on the comparison of the number of patients hospitalized in the two groups. For the Length of Stay P-values were based on median number of days.

Table 29. Health Care Utilization – NYHA III/IV Patients

|  | Health Care Utilization - Results for NYHA Class III and IV Patients |
| --- | --- |
|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P-value |
| Number of patients hospitalized | All-cause: n=176 79 (44.9%) CHF related: n=176 47 (26.7%) | All-cause: n=186 75 (40.3%) CHF related: n=186 39 (21.0%) | 0.40 0.22 |
| Length of Stay (days) | All-cause: 134 hospitalizations median; 7 days mean: 8.9 ( 7.8 CHF related: 70 hospitalizations median; 6 days mean: 8.0 ( 7.0 | All-cause: 127 hospitalizations median; 5 days mean: 7.6 ( 8.2 CHF related: 54 hospitalizations median; 3 days mean: 6.6 ( 7.4 | 0.09 0.04 |

Echocardiographic Parameters

An echocardiographic study was performed to assess the effects of cardiac resynchronization therapy on left ventricular structure and function. Echo data were analyzed by an independent core lab. Table 30 summarizes the echocardiographic results from baseline and 6 months for NYHA Class III and IV patients.  P-values were based on the median paired difference.

Table 30. Echocardiographic Parameters – NYHA III/IV Patients

|  | Echocardiographic Indices – Results for NYHA Class III and IV Patients |
| --- | --- |
|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P-value |
| LV Ejection Fraction (%) | 98 patients with paired data Baseline: median 21.7 mean 23.3 + 6.3 6-month: median 23.7 mean 24.4 + 7.9 Median Paired Difference 1.6 | 97 patients with paired data Baseline: median 23.5 Mean 24.1 + 6.7 6-month: median 25.5 mean 27.6 + 8.9 Median Paired Difference 3.0 | 0.06 |
| Mitral Regurgitation (cm2, jet area) | 68 patients with paired data Baseline: median 7.1 mean 8.5 + 7.6 6-month: median 5.9 mean 7.2 + 6.8 Median Paired Difference –0.5 | 68 patients with paired data Baseline: median 7.0 mean 7.8 + 5.9 6-month: median 4.9 mean 6.7 + 5.6 Median Paired Difference -0.4 | 0.98 |
| Cardiac Index (four chamber view) | 64 patients with paired data Baseline: median 2.45 mean 2.54 + 0.90 6-month: median 2.34 mean 2.48 + 0.88 Median Paired Difference -0.02 | 65 patients with paired data Baseline: 2.38 median mean 2.50 + 0.78 6-month: 2.47 median mean 2.54 + 0.79 Median Paired Difference 0.03 | 0.89 |
| LV Systolic Volume (cm3) | 94 patients with paired data Baseline: median 230 mean 249 + 88 6-month: median 225 mean 241 + 80 Median Paired Difference –8 | 96 patients with paired data Baseline: median 232 mean 242 + 84 6-month: median 212 mean 222 + 94 Median Paired Difference –26 | 0.04 |
| LV Diastolic Volume (cm3) | 94 patients with paired data Baseline: median 302 mean 320 + 96 6-month: median 296 mean 315 + 86 Median Paired Difference -5 | 96 patients with paired data Baseline: median 305 mean 315 + 90 6-month: median 283 mean 298 + 101 Median Paired Difference –25 | 0.046 |
| LV mass (g) | 58 patients with paired data Baseline: median 358 mean 367 + 80 6-month: median 348 mean 365 + 85 Median Paired Difference – 6 | 63 patients with paired data Baseline: median 353 mean 359 + 83 6-month: median 339 mean 354 + 87 Median Paired Difference – 3 | 0.97 |
| Interventricular mechanical delay (IVMD) (ms) | 63 patients with paired data Baseline: median 31 mean 25.4 + 44.2 6-month: median 25 mean 19.1 + 34.4 Median Paired Difference –2 | 79 patients with paired data Baseline: median 35 mean 29.4 + 40.8 6-month: median 20 mean 15.4 + 30.1 Median Paired Difference –18 | 0.12 |
| E Wave /A Wave ratio | 70 patients with paired data Baseline: median 1.22 mean 1.54 + 1.08 6-month: median 1.04 mean 1.45 + 1.03 Median Paired Difference -0.03 | 90 patients with paired data Baseline: median 1.12 mean 1.50 + 1.08 6-month: median 0.99 mean 1.43 + 1.05 Median Paired Difference -0.06 | 0.82 |
| LV Diameter in Systole | 45 patients with paired data Baseline: median 6.8 mean 7.0 + 1.0 6-month: median 6.9 mean 6.7 + 1.0 Median Paired Difference -0.3 | 56 patients with paired data Baseline: median 6.4 mean 6.4 + 1.0 6-month: median 6.3 mean 6.1 + 1.2 Median Paired Difference -0.2 | 0.98 |
| LV Diameter in Diastole | 48 patients with paired data Baseline: median 7.8 mean 10.0 + 10.9 6-month: median 7.6 mean 7.6 + 0.9 Median Paired Difference -0.1 | 58 patients with paired data Baseline: median 7.5 mean 9.6 + 11.6 6-month: median 7.3 mean 7.3 + 1.0 Median Paired Difference -0.3 | 0.85 |

Composite Response

The Heart Failure Clinical Composite Response provides an overall assessment of a patient’s condition.  A patient is defined as “improved” if they decrease NYHA functional class by one or more or if there is a moderate or marked improvement in their global self-assessment score.  A patient is defined as “worsened” if they died, were hospitalized for worsening heart failure.  A patient is defined as “not changed” if the “improved” or “worsened” conditions are not met.

Table 31 summarizes the composite response results for NYHA Class III and IV patients.  The P-value of 0.038 is an overall comparison of the two distributions.

Table 31. Composite Response – NYHA III/IV Patients

|  | Composite Response – Results for NYHA Class III and IV Patients |
| --- | --- |
|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P-value |
| Composite Response Improved Unchanged Worsened | n = 141 40% 26% 33% | n = 150 55% 19% 26% | 0.038 |

At their six-month follow-up patients were asked to rate their global self assessment by answering the following question, “Specifically in reference to your heart failure symptoms, how do you feel today as compared to how you felt before having your InSync® ICD system implanted?”

69% of the treatment patients (n=131) reported at least mild improvement vs. 60% of the control patients (n=121).  The overall p-value comparing the two distributions was 0.68. Note: Patients with global assessment scores are a subset of the patients used in the clinical composite response analysis.

Spontaneous VT/VF Therapy Efficacy

Table 32 summarizes spontaneous VT/VF therapy efficacy of the InSync® ICD in all patients.

Table 32. Spontaneous VT/VF Therapy Efficacy

| Episode Detection Zone | # Patients | Percent Successfully Terminated |
| --- | --- | --- |
| FVT | 14 | 98.5% |
| VF | 41 | 99.2% |
| VT | 50 | 99.2% |
| OVERALL | 78 | 99.1%* |

*Of the episodes that did not satisfy the device’s episode termination criteria, all episodes were  eventually terminated.

Comparison of VT/VF Event Rates

Table 33 compares the VT/VF event rate during the randomization period between NYHA Class III and IV control and treatment groups.  The P-value of 0.38 compares the number of episodes between the control and treatment group.

Table 33. Comparison of VT/VF Event Rates – NYHA III/IV Patients

|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P-value |
| --- | --- | --- | --- |
| Comparison of VT/VF Event Rates | n=177 # patients: 38 (21.6%) Total episodes: 321 Episodes per month: 0.39 | n=187 # patients: 35 (18.8%) Total episodes: 356 Episodes per month: 0.41 | 0.38 |

ATP Therapy Efficacy With Biventricular Pacing

Table 34 summarizes the efficacy of biventricular and RV-only anti-tachycardia pacing (ATP) therapy efficacy for induced and spontaneous ventricular tachycardia (VT) episodes.

Table 34. ATP Therapy Efficacy With Biventricular Pacing

| ATP Site | Rhythm | ATP Success rate | P-value |
| --- | --- | --- | --- |
| RV | Induced VT | 11/20 (55%) | 1.00 |
| RV+LV | Induced VT | 12/20 (60%) |  |
| RV | Spontaneous VT | 247/294 (84%) | <0.0001 |
| RV+LV | Spontaneous VT | 493/523 (94%) |  |
| RV | Spontaneous FVT | 67/74 (91%) | 0.05 |
| RV+LV | Spontaneous FVT | 17/24 (71%) |  |

Plasma Neurohormone Levels

To characterize the effect of cardiac resynchronization, the following plasma neurohormone levels were assessed: brain natriuretic peptide (BNP), big endothelin (Big ET), epinephrine, norepinephrine and dopamine. Neurohormone data were analyzed by an independent core lab.  Table 35 summarizes the neurohormone results from baseline and 6 months for NYHA Class III and IV patients. P-Values were based on the median paired difference.

Table 35. Plasma Neurohormones – NYHA III/IV Patients

|  | Results for NYHA Class III and IV Patients |
| --- | --- |
|  | Control Group (CRT OFF) | Treatment Group (CRT ON) | P-value |
| Brain Natriuretic Peptide (BNP) | 70 patients with paired data Baseline: median 478.5 Mean 889.8 + 1190.0 6-month: median 357.0 mean 799.9 + 1249.6 Median Paired Difference: -2.8 | 83 patients with paired data Baseline: median 450.0 mean 813.0 + 1063.6 6-month: median 320.0 mean 568.9 + 708.2 Median Paired Difference: -38.0 | 0.41 |
| Dopamine | 73 patients with paired data Baseline: median 9.0 Mean 17.2 + 20.9 6-month: median 9.0 mean 19.03 + 26.3 Median Paired Difference: 0.0 | 79 patients with paired data Baseline: median 9.0 mean 13.9 + 8.8 6-month: median 9.0 mean 14.0 + 10.7 Median Paired Difference: 0.0 | 0.56 |
| Norepinephrine | 74 patients with paired data Baseline: median 454.5 Mean 516.9 + 382.3 6-month: median 442.5 mean 478.0 + 289.1 Median Paired Difference: -38.0 | 80 patients with paired data Baseline: median 331.0 mean 376.5 + 263.1 6-month: median 339.5 mean 410.2 + 273.0 Median Paired Difference: 10.5 | 0.24 |
| Epinephrine | 73 patients with paired data Baseline: median 25.0 Mean 53.2 + 109.9 6-month: median 17.0 mean 23.6 + 21.9 Median Paired Difference: -4.0 | 79 patients with paired data Baseline: median 16.0 mean 31.7 + 43.6 6-month: median 17.0 mean 23.2 + 18.5 Median Paired Difference: 0.0 | 0.03 |
| Big Endothelin | 77 patients with paired data Baseline: median 14.3 Mean 26.6 + 34.9 6-month: median 16.4 mean 21.3 + 23.0 Median Paired Difference: –0.5 | 88 patients with paired data Baseline: median 12.3 mean 19.5 + 21.6 6-month: median 11.0 mean 18.1 + 20.0 Median Paired Difference: -0.5 | 0.88 |

Attain LV Lead R-Wave Amplitude Sensing

The sensing performance of the LV leads was assessed by evaluating the R-wave amplitudes from the Model 4189*, 2187 and 2188 leads. Table 36 summarizes the Model 4189* LV lead mean R-wave amplitude data and the two-sided 95 percent confidence intervals on all patients receiving a Model 4189* lead.

Table 36. Model 4189 R-wave Amplitude Data

| Visit | N | Mean (mV) | Two-sided 95% CI |
| --- | --- | --- | --- |
| Implant | 444 | 12.5 ± 4.6 | 12.1-13.0 |
| Pre-hospital Discharge | 454 | 12.6 ± 4.6 | 12.2-13.0 |
| 1 Month | 435 | 13.6 ± 4.5 | 13.1-14.0 |
| 3 Months | 412 | 13.8 ± 5.3 | 13.3-14.3 |
| 6 Months | 317 | 13.7 ± 5.4 | 13.1-14.3 |
| 12 Months | 130 | 14.2 ± 5.7 | 13.2-15.2 |
| 18 Months | 37 | 14.0 ± 5.4 | 12.2-15.8 |

Table 37 summarizes the Model 2187 and 2188 LV lead R-wave amplitude data from all patients receiving a Model 2187 or 2188 lead.

Table 37. Model 2187/2188 R-wave Amplitude Data

| Visit | N | Mean (mV) | Two-sided 95% CI |
| --- | --- | --- | --- |
| Implant | 88 | 12.7 ± 4.9 | 11.7-13.8 |
| Pre-hospital Discharge | 90 | 12.4 ± 4.7 | 11.4-13.3 |
| 1 Month | 85 | 13.3 ± 5.2 | 12.1-14.4 |
| 3 Months | 81 | 13.8 ± 4.8 | 12.7-14.8 |
| 6 Months | 63 | 13.5 ± 5.7 | 12.1-15.0 |
| 12 Months | 30 | 13.9 ± 5.1 | 12.0-15.8 |
| 18 Months | 18 | 12.0 ± 4.2 | 10.0-14.1 |

8.  CS Dissections and Perforations

Table 38 summarizes the adverse events (complications and observations) related to coronary sinus dissections and perforations in all patients. Of the 33 patients with CS dissection/perforation, there were six deaths (82, 263, 384, 415, 436 and 509 days post-implant). None of the deaths was reported as related to the dissection or perforation.  Two of the patients that died received the InSync® ICD system, and four of the patients did not receive the InSync® ICD system (two  received the GEM DR and two received the GEM III DR).

Table 38. CS Dissections and Perforations

| Event | Complications (# pts.) | Observations (# pts.) | Total (# pts.) |
| --- | --- | --- | --- |
| CS dissection | 15 | 9 | 24 |
| Cardiac vein/ CS Perforation | 9 | 0 | 9 |
| Total | 24 | 9 | 33 |

9.  LV Lead Placement

Table 39 summarizes information related to LV lead placement and positioning. The majority of LV leads (72%) were positioned in lateral or posterior-lateral cardiac veins.

Table 39.  Final LV Lead Placement and Position

| Vessel/Location | Basal | Mid | Apical | Not indicated | Total |
| --- | --- | --- | --- | --- | --- |
| Lateral (marginal) cardiac | 6% | 31% | 2% | < 1% | 39% |
| Posterior-lateral cardiac | 4% | 25% | 4% | < 1% | 33% |
| Anterior cardiac | 5% | 10% | 1% | < 1% | 17% |
| Posterior cardiac | < 1% | 3% | 1% | < 1% | 5% |
| Middle cardiac | < 1% | 2% | 1% | < 1% | 4% |
| Great cardiac | < 1% | < 1% | < 1% | < 1% | 1% |
| Not indicated | < 1% | < 1% | < 1% | 1% | 1% |
| Total | 16% | 72% | 10% | 2% |  |

- Confirmation of Biventricular Pacing and ICD Function

To confirm proper biventricular pacing and ICD function, percent ventricular pacing, biventricular capture during cardiopulmonary exercise (CPX), VT/VF detection efficacy and VF detection time were assessed. To assess percent ventricular pacing, the percentage of time that patients were paced was determined from the pace/sense counters of the device at the 6- month follow-up. Electrocardiograms (ECGs) with biventricular pacing ON from the 6-month follow-up visit were compared to 12-lead ECGs at peak heart rate during the 6-month CPX test to confirm that biventricular pacing and capture were maintained.  The incidence of inappropriate VT/VF episodes was analyzed to confirm that there is nothing unique to a CHF ICD population that would increase the risk of inappropriate detection.  Analysis of VF detection times in InSync® ICD treatment vs. control patients was done to demonstrate that VF detection is not compromised by the presence of the LV lead and biventricular pacing.

1.  Percent Ventricular Pacing

The percentage of time patients were paced during their randomized period was determined from each patient’s six-month visit, using the pace/sense counters over the lifetime of the device.  All NYHA Class III and IV randomized patients who did not cross over into the other group (OFF to ON or ON to OFF) are included (116 control and 125 treatment patients).  Save-to-disk files were not available for 6 control and 7 treatment patients.

In Figure 11, the left-hand graph indicates that greater than 90% of the treatment patients received ventricular pacing more than 90% of the time, and conversely, greater than 90% of the control patients received no ventricular pacing more than 90% of the time.  The right-hand graph indicates that more than 85% of the treatment patients had capture margins greater than or equal to 100%.

Figure 11. Biventricular Pacing and Capture Margin

2.  Biventricular Capture During Exercise

NYHA Class III and IV patients with biventricular pacing randomized ON and with paired baseline and 6-month cardiopulmonary exercise (CPX) data (n=118) were included in the analysis. Of these 118 patients, 112 had interpretable 12-lead electrocardiograms (ECGs). ECGs with biventricular pacing ON from the 6-month follow-up visit (just prior to CPX testing) were compared to 12-lead ECGs at peak heart rate during the 6-month CPX test to confirm that biventricular pacing and capture was maintained.

Of the 112 NYHA Class III and IV patients with paired baseline and 6-month CPX data and interpretable 6-month pre- and peak exercise 12-lead ECGs, 83 patients had biventricular pacing at peak exercise. All of these patients maintained biventricular capture at peak exercise. Twenty-seven (27) patients lost biventricular pacing at maximum exercise due to the intrinsic rate exceeding the programmed Upper Tracking Rate (UTR). Two patients did not maintain pacing due to other unique rhythm/programming conditions described in Figure 12. There were no cases of LV-only or RV-only capture during biventricular pacing at any pacing rate.

1UTR programmed to a higher rate than their “normal” UTR setting during the CPX test for 2 patients.

2No biventricular pacing due to PVC initiated Vs-Ar-Vs-Ar… pattern.

3No biventricular pacing due to an intrinsic PR interval less than the programmed SAV interval at peak exercise             (Rate Adaptive AV programmed OFF).

Figure 12. Biventricular Capture During Exercise Testing

3.  VT/VF Detection Efficacy

The incidence of inappropriate VT/VF episodes was similar in the InSync® ICD study population as compared to the Gem DR study population (21.3% vs. 21.9%). The reasons for the inappropriate detections were similar, and resulted from a similar proportion of atrial fibrillation or flutter as compared to 1:1 SVTs. None of the inappropriate detections in the InSync® ICD study were related to biventricular pacing or the LV lead, confirming that there is nothing unique to a CHF ICD population that would increase the risk of inappropriate detection. There were no VT/VF episodes for which therapy was inappropriately withheld. The InSync® ICD with its left ventricular lead and possibility of biventricular pacing did not adversely affect the device’s ability to detect VT/VF or appropriately withhold VT/VF therapy or delay detection time.  47% of the 135 episodes were treated with at least one shock. No deaths, hospitalizations or worsening of heart failure were attributed to inappropriate shocks.

Table 40. Classification of VT/VF Episodes, Gem DR vs. InSync® ICD

|  | GEM DR | InSync® ICD |
| --- | --- | --- |
| Total episodes in VT/VF log | 3945 (278 pts) | 950 (100 pts) |
| Appropriate VT/VF episodes | 3488 (232 pts) | 815 (85 pts) |
| Inappropriate VT/VF episodes: GEE adjusted rate: - Atrial or RV sensing related - Ventricular rate during AF in VF zone - SVT Criteria programming (PR Logic) - Detection algorithm characteristics | 457 (86 pts) (12%) 21.9% 93 (20%) 44 (10%) 68 (15%) 252 (55%) | 135 (31 pts) (14%) 21.3% 28 (21%) 5 (4%) 53 (39%) 49 (36%) |
| Total patients | 933 | 371 |
| Average follow-up time (months) | 3.9 | 7.9 |

4.  VF Detection Time

Analysis of VF detection times in InSync® ICD treatment vs. control patients demonstrate that VF detection is not compromised by the presence of the LV lead and biventricular pacing.

When the number of intervals to detect VF (VF NID) was programmed to 12/16 intervals, the mean VF detection time for the treatment group was 3.8 seconds, and the mean detection time for the control group was 3.4 seconds. These data demonstrate that there was no delay in detection time compared to the theoretical time to detect (P=0.07) described below:

    - VF at a rate of 280 ms x 12 intervals to detect = 3360 ms = 3.4 seconds

When the number of intervals to detect VF (VF NID) was programmed to 18/24 intervals, the mean VF detection time for the treatment group was 4.9 seconds, and the mean detection time for the control group was 5.3 seconds. These data demonstrate that there was no delay in detection time compared to the theoretical time to detect described below:

    - VF at a rate of 280 ms x 18 intervals to detect = 5040 ms = 5.0 seconds

Table 41. VF Detection Times in Treatment and Control Patients 
During the Randomization Period

| Randomization Assignment | VF NID = 12/16 | VF NID = 18/24 |
| --- | --- | --- |
|  | N | Mean Vent. Cycle Length (ms) | Mean Detection Time (sec) | N | Mean Vent. Cycle Length (ms) | Mean Detection Time (sec) |
| Control | 51 | 272 | 3.4 | 14 | 278 | 5.3 |
| Treatment | 42 | 283 | 3.8 | 28 | 256 | 4.9 |
| p-value |  |  | 0.07 |  |  | 0.84 |

- Programmed Parameters

Table 42 shows how Upper Tracking Rate was programmed in NYHA Class III and IV treatment patients at the end of the randomization period (6 months).

Table 42. Programmed Upper Tracking Rate (UTR)

| Programmed Value (ppm) | % of Patients |
| --- | --- |
| <120 120 125-130 136-140 >140 | 5% 49% 11% 33% 2% |

Table 43 demonstrates that the majority of NYHA Class III and IV treatment patients were able to maintain over 90% ventricular pacing regardless of their programmed Upper Tracking Rate.

Table 43. Percent of Ventricular Pacing in Treatment Group 
By Programmed UTR

| % Paced | UTR < 120 (% of patients) | UTR = 120 (% of patients) | UTR > 120 (% of patients) |
| --- | --- | --- | --- |
| <90% |  | 6% | 2% |
| 90-94% |  | 21% | 16% |
| 95-99% | 100% | 70% | 71% |
| 100% |  | 4% | 11% |

E.  Gender Bias Analysis

A sub-group analysis was performed to evaluate the association of gender with the primary effectiveness endpoints (NYHA functional class, quality of life and 6-minute hall walk distance).  Based on this analysis, gender appears to be associated with an improvement in QOL scores, while gender was not associated with NYHA class or 6-minute hall walk distance. The InSync® ICD study enrolled 75 percent males and 25 percent females. There was no selection bias based on gender.

F.  Conclusions Drawn from the Clinical Study

1.  Safety

The InSync® ICD pulse generator met the primary and secondary safety endpoints.  Results were within protocol specified performance criteria for the rate of severe device related adverse events and operative mortality. The study additionally demonstrated that the ICD portion of the InSync® ICD was not adversely affected by the addition of cardiac resynchronization as shown by adequate ventricular fibrillation detection times, antitachycardia pacing (ATP) conversion efficacy, and no increase in the amount of inappropriate ICD therapies.

There was no difference in the mortality, number of hospitalizations, or incidence of malignant ventricular arrhythmias between the control and the treatment group.

2.  Effectiveness

The InSync® ICD cardiac resynchronization system demonstrated a six- month  improvement in quality of life and a reduction in NYHA functional class which were two of the three primary endpoints of the study. The third endpoint, six-minute hall walk, showed no difference between the treatment and the control group. Given this, results from the secondary endpoint of cardiopulmonary exercise testing were used to show a significant improvement in peak VO2 in the treatment group. Additional supporting evidence included an improvement in heart failure status as measured by a composite response assessment.

XI.  Panel Recommendation

FDA’s advisory panel met on March 5, 2002 and voted 6 to 5 for approvable with conditions.  The conditions that were recommended by the Panel included a post-approval study to obtain longer-term information on mortality and performance of the device.  They also requested information regarding the administrative censoring of patients, including documentation of the explicit original intent of the agreed upon sample size of the study. Additional concerns were raised regarding device programming and interaction issues between the ICD and CRT portions of the device.  The Panel requested that questions that were raised on this issue be resolved by submitting additional information to the Agency.  There were significant concerns raised about the safety profile of the Attain Model 4189*.

CDRH Decision

FDA found Medtronic’s facilities in compliance with the Quality System Requirements (21 CFR part 820).

FDA requested the additional information that was discussed at the panel meeting. The approval decision was based on reviewing the entire cohort of class III/IV patients that were enrolled in the study. Because of the subjective nature of the successful primary endpoints of the study, FDA specifically used data from the supporting secondary endpoints to determine a reasonable level of effectiveness of the device. Specifically, data from the cardiopulmonary exercise testing was used to demonstrate a functional improvement in the treatment group. Subgroup analysis did not delineate a specific group of patients that benefits more from this device.

Data were provided to the agency regarding device performance that included ICD/CRT interaction, percentage of time paced as well as other programming issues that were used to further define the limitations of the device.  Lead performance data in the clinical trial did not support reasonable assurance of the safety and effectiveness of the 4189  lead.

The conditions of approval include a 3-year evaluation of mortality and chronic system performance, including adverse clinical events, on 1000 patients to assess the long-term safety of the device. Physicians will be required to undergo training by the sponsor prior to implanting the system.

FDA issued an approval order for P010031 on June 26, 2002. This decision was based on a complete dataset as well as additional supporting evidence regarding device functionality.

XII.  Approval Specifications

Directions for Use:		See labeling.

Hazards to Health from Use of the Device:	See Indications, Contraindications, Warnings, Precautions and Adverse Events in the labeling

# Footnotes

* These data are included for completeness in reporting the results of the clinical study.  The 4189 lead is not currently market approved in the U.S.

* These data are included for completeness in reporting the results of the clinical study.  The 4189 lead is not currently market approved in the U.S.

*  These data are included for completeness in reporting the results of the clinical study.  The 4189 lead is not currently market approved in the U.S.

* These data are included for completeness in reporting the results of the clinical study.  The 4189 lead is not currently market approved in the U.S.

* These data are included for completeness in reporting the results of the clinical study.  The 4189 lead is not currently market approved in the U.S.

* These data are included for completeness in reporting the results of the clinical study.  The 4189 lead is not currently market approved in the U.S.

Weber KT, Janicki JS. “Cardiopulmonary Exercise Testing:  Physiologic Principles And Clinical Applications”. WB Saunders Company, 1986.

Packer et al. J Cardiac Failure 2001;7:176-182.

* These data are included for completeness in reporting the results of the clinical study.  The 4189 lead is not currently market approved in the U.S.

Lee, EW and Dubin, N. Estimation and Sample Size Considerations for Clustered Binary Responses. Statistics in Medicine, Vol 13, 1241-52 (1994).

Wilkoff, BL, et al. Critical Analysis of Dual-Chamber Implantable Cardioverter-Defibrillator Arrhythmia Detection: Results and Technical Considerations.  Circulation, Vol 103, 381-386 (2001).

* These data are included for completeness in reporting the results of the clinical study.  The 4189 lead is not currently market approved in the U.S.

# Text boxes

Not Randomized 3

| II | III/IV | Total |
| --- | --- | --- |
| 5 | 10 | 15 |

Randomized

| II | III/IV | Total |
| --- | --- | --- |
| 191 | 364 | 555 |

Unsuccessful Implants 2

| II | III/IV | Total |
| --- | --- | --- |
| 18 | 50 | 68 |

Successful Implants

| II | III/IV | Total |
| --- | --- | --- |
| 196 | 374 | 570 |

No Implant Attempt 1

| II | III/IV | Total |
| --- | --- | --- |
| 12 | 10 | 22 |

Implants Attempted

| II | III/IV | Total |
| --- | --- | --- |
| 214 | 424 | 638 |

Enrolled Patients

| II | III/IV | Total |
| --- | --- | --- |
| 226 | 434 | 660 |

ON (Treatment)

|  | II | III/IV | Total |
| --- | --- | --- | --- |
| Total | 85 | 187 | 272 |
| At six months |
| Completed 6 mo. follow-up | 82 | 167 | 249 |
| No 6 mo. data received | 1 | 6 | 7 |
| Patient Death | 2 | 14 | 16 |

OFF (Control)

|  | II | III/IV | Total |
| --- | --- | --- | --- |
| Total | 106 | 177 | 283 |
| At six months |
| Completed 6 mo. follow-up | 102 | 158 | 260 |
| No 6 mo. data received | 2 | 4 | 6 |
| Patient Death | 2 | 15 | 17 |