SOP for working with H2 Gas

1.0 Introduction

This document is the standard operating procedure (SOP) for spectroscopic studies of alkene hydrogenation reactions at elevated temperature and pressure.  This SOP covers the experimental plan, safe operation and emergency procedures in the event of an off-normal event.  The system will be used at the APS, Sector 20 BM.  The system is used for x-ray studies of IPA/benzene/H2 solutions.

1.1  Description of the System

A small pressure vessel having an internal volume of 3 milliliters is used to study a chemical reaction using X-ray absorption spectroscopy.  During operation, the contents of this vessel are heated to 100°C and 750 psi with H2 gas.  The volume of this heated material is small.  The cell is loaded and pressurized in the laboratory hood.  The entire apparatus is then transferred to the hutch for in situ reactions studies.  This SOP covers the procedure for loading, filling the cell and emptying the process line after completion of an experiment.

More specifically the apparatus and procedure is described as follows.  The complete apparatus is shown in the attached Figure.  The apparatus consist of four main components, a lecture bottle, an injection reservoir, an accumulator and the x-ray cell.  The entire assembly shown in the Figure is mounted on a 12 inch square pad that can be easily transported between the lab hood and the x-ray hutch.  All of the sample preparation is completed in the laboratory hood and pressurized initially to 200 psi.  The entire assembly is transported to the hutch were the x-ray cell is heated to 100C and pressurized to 750 psi to initiate the reaction.

All of the components are constructed from commercially available pressure fittings (High Pressure Equipment,  Inc) that are rated to at least 10,000 psi.  The entire system will be leak check with an inert gas prior to use.  The possibility of leaks during the experiment in the x-ray hutch are very low.

The total gas volume contained within the apparatus is very low.  The total amount of H2 in the lecture bottle is 2 ft3 after expansion to ambient conditions.  In the unlikely event that this gas rapidly escapes during the experiment, this represents a total volume ratio in the hutch of about 0.1% which is far below the flammability/explosive limit.

1.2 Types of risks

Exposure to hydrogen gas.

Exposure of hot surfaces

Exposure to high pressure equipment

Exposure to organic solvent (10% benzene in IPA)

Special guidelines have been developed to minimize these risks during normal operation or in the event of and off-normal event.  These are discussed in the following sections.

1.3  Qualified Operators for the System

The conditions of operation requires that operators be familiar with all aspects of the equipment and in particular be well-versed in the contents of this document.

# 2.0  GENERAL PROCEDURES

# 2.1  General Practices and Precautions

The system will be used at the Advanced Photon Source of Argonne National Laboratory.

During normal operation of the system, safety glasses are to be worn at all times when operating this equipment including during mixing and loading of the cell.

# 2.2  Safety Features in APS Bldg.

# Safety Shower: 		Located in the chem lab.

Eye Wash:		 	Located in the chem lab.

Fire extinguisher:	In the hallway.

Fume Hood:		 	Located in the chem lab.

Fire Alarm:			In the hallway.

These facilities provide the basis for safe handling of this hydrogenation apparatus.

# 2.3  Precautions for Known Hazards

# 2.3.1  Precautions for Hydrogen Gas

# Standard procedures for handling hydrogen gas should be followed.  All equipment during operation should be approached with care and assumed that it is under pressure until otherwise observed.

# During normal operation the cell will be at 750 psi.  The cell, the tubing and fittings are all designed for operation to 10,000 psi.  There is a safety or rupture disc that is rated at 3,000 psi which is well above the lecture bottle pressure of 1800 psi.

# Prior to filling with H2 gas the entire system will be purge with an inert gas (He, Ar, or N2) and leak checked.  The cell will be filled with H2 gas to 200 psi while the apparatus in the laboratory hood. After transport of the apparatus to the x-ray hutch the spectroscopy cell will be heated to 100°C and filled to 750 psi with H2.   After completion of the experimentation the entire system will be transport back to the laboratory hood and will be purged with an inert gas to remove residual H2.

# 2.3.2  Precautions for Hot/Cold Surfaces

# Standard procedures for handling hot surfaces should be followed.  All equipment during operation should be approached with care and assumed that it is at elevated temperature until otherwise observed.

# During operation, only one component of the system, the spectroscopy cell will be at elevated temperature.  This cell temperature is no hotter than 100°C.  The spectroscopy cell will be covered with a small fiberglass blanket to completely cover all exposed surfaces.

2.3.3  Exposure to High Pressure Equipment

This system consists of high pressure equipment and tubing.  The vessels and piping are rated to 10000 psi, and is protected by a rupture disk rated at 3000 psi. No tightening or loosening should be performed on any fittings or pipes that are above 100psi.

2.3.4 Exposure to organic liquids

Material handling for this system is to be done in accordance with current safe practices.  Isopropyl alcohol will be used in this system as a mixture containing up to 10% benzene or cyclohexene,  1% triethylamine and  less than 1% by weight of a Rh catalyst.  These materials should be handled in a chemical fume hood with gloves and safety glasses.

Only 3 ml of solution (with 0.3 ml benzene) will be used during each experimental acquisition. In the case of a spill, absorb with sand or other non-combustible material. Collect spilled material in appropriate container for disposal. The vapor pressure of benzene is 75 mmHg @ 20 C and the evaporation rate is 5.1 (butyl acetate=1) . Since the quantities are low and the evaporation rate is limiting, the maximum air concentration in the hutch will be well below the exposure limit (5 ppm OSHA STEL 15 minute(s)) in the event of a spill.

During the experimental operation a small amount of liquid (3.0 ml) will be heated to 100°C.   The heated volume is relatively small.

Benzene is potentially fatal on contact with the skin, causes respiratory tract irritation, skin irritation, eye irritation, aspiration hazard, blood damage, central nervous system depression, cancer hazard (in humans).  Do not breathe gas, fumes, vapor, or spray.   In case of contact with eyes, rinse immediately with plenty of water and seek medical advice.  In case of skin contact, wash skin with soap and water for at least 15 minutes

# 2.4  Handling Experimental Solutions

As a part of experimental planning, waste materials should be considered and their production minimized.  The operators of the system will take special precautions so that no liquid waste material is discharged into the environment.  Used IPA/benzene solutions and system rinse material will be stored in properly managed satellite accumulation areas during the project and disposed of as per APS requirements.

3.0 OPERATIONAL PROCEDURES

By following the prescribed Cell Loading, and System Operation checklists, the operators can insure the system is safely operated.  Any significant deviations from the checklist procedures must be noted both on the checklist and in a LRB.

3.1  CELL LOADING AND PURGING

- Don safety gear.  Required gear for this operation are safety glasses and 8 inch Viton gloves.  Two researchers must be present for loading operations.

- All filling operations will be conducted in a laboratory fume hood.

- In a beaker placed in a hood, mix the 5 ml aliquot of isopropyl alcohol containing up to 10% benzene or cyclohexene,  1% triethylamine and  less than 1% by weight of a Rh catalyst  according to the experimental plan.  More preferably, the reagents will be mixed using anaerobic techniques which include, syringing the liquid reagents into a sealed mixing vessel to minimize vapor loss to the environment.

- Using a glass syringe, transfer 3 ml of this IPA solution into  the x-ray cell.

- Attach the x-ray cell to the hydrogenation apparatus.

- Purge the entire system with inert gas prior to filling with H2 and pressurize to 350 psi and check for leaks.

- Fill system with H2 to 200 psi and discharge 6 times to purge inert gas from the cell and solution.

3.2  SYSTEM PREPARATION, OPERATION AND RINSING

- Don safety gear.  Required gear for this operation are safety glasses and 8 inch Viton gloves.  Two operators must be present for operations.

- Transport the hydrogenation system from the laboratory hood to the x-ray hutch.

- Insert heaters with and heat the solution in the x-ray cell to 100°C while gradually increasing the pressure to 750 psi.

- Acquire spectra of the reacting system for 2 hrs or until reaction completion.

- After data collection is complete.  Cool the apparatus.  Transport the hydrogenation apparatus into the laboratory hood.  Decrease the pressure through the discharge valve. Purge the reactor with inert gas and disconnect from the pressure apparatus.  Collect the reacted effluent from this operation in appropriate waste container or sampling vial for future analysis.

4.0  EMERGENCY SHUTDOWN PROCEDURES

Emergency shutdown procedures must be used whenever there is a risk of injury to personnel or equipment posed by the application of the regular shutdown procedures.  This includes shutdowns due to:

	rapid release of pressure (leaks or blown rupture discs);

!	unusual operating conditions;

	building evacuation during a run;

	loss of electrical power;

	other unusual or potentially hazardous circumstances as determined by the operators.

# Emergency Shutdown Procedure

- Turn OFF the heater controllers

- Close the outlet valve on the H2 lecture bottle.