ATDEPARTMENT OF HEALTH AND HUMAN SERVICES

FOOD AND DRUG ADMINISTRATION

CENTER FOR DEVICES AND RADIOLOGIC HEALTH

NATIONAL MAMMOGRAPHY QUALITY ASSURANCE

ADVISORY COMMITTEE

Monday, August 26, 2002

9:00 a.m.

Holiday Inn Gaithersburg

Two Montgomery Village Avenue

Gaithersburg, Maryland

PARTICIPANTS

Maryanne Harvey, Chair

Charles Finder, M.D., Executive Secretary

MEMBERS

James F. Camburn

Nancy J. Ellingson, R.T. (R)(M)

Alisa M. Gilbert

Miles G. Harrison, Jr., M.D.

Jessica W. Henderson, Ph.D.

Debra M. Ikeda, M.D.

Andrew Karellas, Ph.D.

Amy F. Lee, M.D.

Melissa G. Martin, M.S.

Etta D. Pisano, M.D.

Linda S. Pura, R.N., M.P.

Catalina R. Ramos-Hernandez, M.D.

Amy R. Rigsby, R.T. (M)

Donald C. Young, M.D.

C O N T E N T S

Page No.

Opening Remarks:

Maryanne Harvey             	                   4

Conflict of Interest Statement:

Charles Finder, M.D.          	                 4

Introductions                             	               7

Alternative Standards

Charles Finder, M.D.              	             10

Roger Burkhart, Ph.D.               	          11

Open Public Hearing

Ken Crocker                           	        23

Pamela Gormley (by Dr. Finder)          	      33

Overview of MQSA Inspection Findings

and Current Inspection Follow-up Actions:

Michael Divine, M.S.                      	    35

Good Guidance Practices and Directions

for Discussion of the MQSA Guidance under

the Final Regulations:

Charles Finder, M.D.                      	   48

ACRIN Trial of Full Field Digital

Mammography (FFDM):

Etta Pisano, M.D.                  	          121

Potential New Applications of FFDM:

Daniel Kopans, M.D.                  	        169

Inspection Demonstration Project Update:

Charles Gunzburg, M.S.                 	      221

Full Field Digital Mammography

Accreditation and Certification Update:

Kish Chakrabarti, Ph.D.              	        226

Priscilla Butler, M.S.                 	      236

States as Certification Agencies Update:

Roger Burkhart, Ph.D.               	         256

Review of Summary Minutes of August

2001 Meeting                               	            263

Future Meetings                              	          264

P R O C E E D I N G S

Opening Remarks

MS. HARVEY:  Good morning.  The National Mammography Quality Assurance Advisory Committee is called to order.  We have a very busy schedule today so it will be important for us to keep to our schedule as we go.  My name is Maryanne Harvey, for those of you who haven't met me as yet.

Dr. Finder will now give us our conflict of interest statement.

Conflict of Interest Statement

DR. FINDER:  The following announcement addresses conflict-of-interest issues associated with this meeting and is made a part of the record to preclude even the appearance of any impropriety.  To determine if any conflict existed, the agency reviewed the submitted agenda and all financial interests recorded by the committee participants.

The conflict-of-interest statutes prohibit special government employees from participating in matters that could affect their or their employer's financial interest.  However, the agency has determined that participation of certain members, the need for whose services outweighs the potential conflict of interest involved is in the best interest of the government.

Therefore, waivers permitting full participation in general matters that come before the committee have been granted for certain participants because of their financial involvement with facilities that will be subject to FDA's regulations on mammography quality standards with accrediting, certifying or inspecting bodies, with manufacturers of mammography equipment or with their professional affiliations since these organizations could be affected by the committee's deliberations.

These individuals are James Camburn, Nancy Ellingson, Alisa Gilbert, Maryanne Harvey, Melissa Martin, Linda Pura, Amy Rigsby and Drs. Miles Harrison, Jessica Henderson, Catalina Ramos-Hernandez, Debra Ikeda, Andrew Karellas, Daniel Kopans, Amy Lee, Etta Pisano and Donald Young.

Copies of the waivers may be obtained from the agency's Freedom of Information Office, Room 12A-15, of the Parklawn Building.

Several of our members also reported that they received compensation for lectures they have given or will give on mammography-related issues.  However, they have affirmed that these lectures were offered because of their expertise in the subject matter and not because of their membership on the committee.

We would like to note, for the record, that if any discussion of states as certifying bodies was to take place in any meetings of the committee, it would be a general discussion only.  No vote would be taken and no consensus sought.

In the interest of getting as many viewpoints as possible, all SGEs, including state employees, would be allowed to participate in the general discussion so that all viewpoints could be heard.

In the event that the discussions involve any other matters not already on the agenda in which an FDA participant has financial interest, the participants should excuse him or herself from such involvement and the exclusion will be noted for the record.

With respect to all other participants, we ask, in the interest of fairness, that all persons making statements or presentations disclose any current or previous financial involvement with accreditation bodies, states doing mammography inspections under contract to FDA, certifying bodies, mobile units, breast implant imaging, consumer complaints and mammography equipment.

MS. HARVEY:  Thank you, Dr. Finder.

Introductions

Since we have four new members at this meeting--it is nice to see our new members and to see our returning members from last year--I would ask that each one of us give a very short bio so that we can get more acquainted with each other's experience and background.

Dr. Pisano, could I ask you to begin?

DR. PISANO:  I am Dr. Etta Pisano.  I am the Chief of Breast Imaging at the University of North Carolina in Chapel Hill.  I am a radiologist.

DR. YOUNG:  I am Don Young.  I am a radiologist, a professor of clinical radiology at the University of Iowa College of Medicine and practice at the hospital and clinics where I direct the breast imaging and diagnostic center.

DR. RAMOS-HERNANDEZ:  I am Catalina Ramos with the National Breast Cancer Organization.  We are a not-for-profit advocacy and counseling services for patients with breast cancer.

MS. RIGSBY:  I am Amy Rigsby.  I am the Technical Director of the Rose Breast Imaging Center in Houston, Texas.  I am a mammographer.

MS. MARTIN:  I am Melissa Martin.  I am a medical physicist running a consulting practice in Southern California.

DR. IKEDA:  I am Debra Ikeda.  I am Director of Breast Imaging at Stanford University Medical Center.  I am a radiologist.

DR. KARELLAS:  I am Andrew Karellas.  I am a medical physicist.  I have been with the University of Massachusetts as of two weeks ago and now I have moved to join the faculty at Emory University in Atlanta.

DR. HARRISON:  I am Miles Harrison.  I am a surgeon by training.  I am part of the Sinai Surgical Associates in Baltimore, Maryland and one of the designated breast surgeons at the Lapedes Cancer Center which is a Hopkins affiliate.

DR. FINDER:  Dr. Charles Finder.  I am a radiologist working for the Food and Drug Administration and I am also the Executive Secretary of this committee.

MS. HARVEY:  I am Maryanne Harvey.  I am with the New York State Department of Health.  I am a section chief who is responsible for mammography and also the Chairman of this committee.

MS. PURA:  Good morning.  I am Linda Pura.  I am one of the clinical nurse coordinators for the Los Angeles County Breast Cancer Early Detection Program under the Department of Health California, Cancer Detection Section.  I am also the co-founder and present President of the Los Angeles County Susan G. Komer Breast Cancer Foundation.

MR. CAMBURN:  I am Jim Camburn.  I am Chief of the Radiation Safety Section for the State of Michigan.

MS. ELLINGSON:  I am Nancy Ellingson.  I am a radiologic technologist and mammographer.  I am with the American Society of Radiologic Technologists in Albuquerque, New Mexico.  We represent about 100,000 radiologic technologist members.

MS. GILBERT:  I am Alisa Gilbert.  I am a seven-year breast cancer survivor.  I work with Alaska Natives and American Indians.  I am the Director of the National Native Cancer Survivor Support Network.

MS. HENDERSON:  I am Jessica Henderson.  I am an eight-year cancer survivor and I represent the Oregon Breast and Cervical Cancer Coalition.

DR. LEE:  Hi.  I'm Amy Lee.  I am Assistant Professor of Community Medicine at Northeastern Ohio University's College of Medicine and Administrative Director for the Master of Public Health Program also located there.  I am also a physician consultant for the local Breast and Cervical Cancer Program and, in my former life, before academia, I was OB-GYN.

MS. HARVEY:  Thank you.

I will now ask Dr. Finder to talk to us about alternative standards.

Alternative Standards

DR. FINDER:  I just want to give a little bit of background on approval for alternative standards.  For those not familiar with this section of the regulations, FDA may approve an alternative to a quality standard that currently exists under Section 900.12 when the agency determines that, one, the proposed alternative standard will be at least as effective in assuring quality mammography as the standard it proposes to replace and, two, the proposed alternative is too limited in its applicability to justify an amendment to the standard or it offers an unexpected benefit to human health that is so great that the time required for amending the standard would represent an unjustifiable risk to human health and also that the granting of the alternative is in keeping with the purposes of the statute.

Since our last meeting, the division has approved several alternative standards and these will be discussed by Dr. Roger Burkhart.

DR. BURKHART:  I might begin by referring you to one of the documents that you were given in preparation for the meeting, the Modifications and Additions to the Policy Help Guidance System, No. 5.  You will find the new alternative standards included within this document.

The first one, the first new one which we approved last September, is found on Page 57 of the document.  It is entitled The Manufacturer's Software Modification of the Automatic Exposure Control, but really what it applies to is the testing which has to take place after such modifications occur.

Software upgrades or modifications are defined by FDA as being major changes in the system which means that, after they are to take place, there has to be a mammography equipment evaluation conducted of the equipment and any problems that are found have to be corrected before the equipment is used on patients.

Also, the regulations require that this evaluation and the testing which is involved be done by the physicist on site.  But the applicant for this particular alternative requirement made a convincing case that, in this particular situation, we can assure mammography quality if on-site testing is done under conditions of medical physicist oversight.

By medical physicist oversight, what we mean is that the physicist has to be consulted, but it is his or her decision as to whether they actually have to come on site to do the testing or whether somebody else can do the testing and send the results to them for evaluation.

As I said, the applicant for this standard made a convincing case that medical physicist oversight would assure quality in this particular case.  So, for this specific software modification has given an alternative standard, and when it is applied to the units and the models which are listed in the standard, medical physicist oversight is an option for the facility.

The second newly approved alternative standard was reapproved last May as found on page 58, and it, too, refers to the testing, the testing conditions or how the testing is done after the modification takes place.

Like the last one I just mentioned, it also started out as a request related to a specific software upgrade, but in this case, the justification was that the testing which would take place after this particular upgrade, was the same type of testing which is done routinely by the quality control technologist.

So, it was felt that if the quality control technologist is qualified to do this as part of the routine QA program, then, they should be able to do it after the software modification, and so it is to be done in conditions of medical physicist oversight.

But we got to thinking if it applies to this particular upgrade, it could also apply to any with that particular qualification, so we do have the authority to expand on requests, and we did that.

We applied this alternative standard to any upgrade or modification of the computer software in which the testing afterwards is normally done by the quality control technologist.

So, if the manufacturer feels that the modification meets this particular standard, they need to consult with us, confirm with us that we agree with this, and if we do, then they can go ahead and this testing after this modification can be done under conditions of medical physicist oversight.

The third new alternative standard, which begins on page 60, is different in nature.  It deals with the quality assurance program including the quality control testing, full field digital mammography units, again to be more specific, it deals with the time period for which the corrective actions can be taken if the testing reveals that there are problems with the system.

Now, in the case of screen-film systems, we were able, at the time of the regulations, to define two classes of test failures.  There are those test failures which are significant enough that the problem causing them has to be corrected before the piece of equipment is used again on patients.

Then, there are those which for various reasons mostly dealing with the fact that there are compensating methods, we can allow more time for the correction of the test failures for the problem causing the test failures, and the time which was set for that was 30 days.

For screen-film systems, as I say, we can make that distinction, but for full field digital mammography systems, which were still in the research stage at the time the regulations were being developed, we couldn't make such a distinction.

So, for full field digital mammography systems, and any other new modality that might appear in the coming years, we took a conservative public health safety position. We said that any quality control test that has failed, the problem causing the failure has to be corrected before the equipment can be used on patients.

So, this is what the regulations say, and, of course, the expectation is that eventually, if a technology like full field digital mammography, once it becomes fully established, eventually, we would be able to rewrite the regulations and make a similar distinction as we did with screen-film.

But to begin with then, we took this position, but at the same time, we tried to make it clear to the manufacturers and also to any facility that has a full field digital mammography unit that if they feel that there are some of the tests that they can make a case that a 30-day correction period could be allowed for test failures of that test, they could always apply for an alternative requirement.

About two months ago, General Electric made such an application for their Senographe 2000D, full field digital mammography unit, and we approved that alternative on July 14th.

Basically, what this alternative does is divide the quality control tests of the 2000D into three general groups, and an important one as far as terms of the changes which are involved is the third group or the group which is labeled with the letter C.

These were tests, quality control tests of the 2000D, which were equivalent to quality control tests of screen-film systems for which the 30-day correction period was already allowed.

Some of them involved the exact same testing methods using the exact same action limits and even testing components which were identical to the full field digital mammography system as in some of the screen-film systems, so they were than equivalent in that case, they were virtually identical.

The other tests were not quite as identical, but basically, for most of them, the only difference is that the measurements are done off of digital images rather than off of films as it would be in the case in screen-film systems.

So, it seemed logical then that if these tests already were allowed a 30-day correction period when problems are found, then, it would be logical to allow the same 30-day correction period for them with full field digital mammography unit, with the 2000D.

During this 30-day correction period--perhaps I didn't make this quite clear earlier--during this 30-day correction period, the facility continued to use the part of the system which failed the test, they wouldn't have to take it out of service.

The other two groups are the remaining tests, and these are still tests which, if they are failed, the component which failed the test has to be taken out of service until the problem is corrected.

The only reason really for dividing them into two groups is to emphasize something which also exists in the screen-film area, and that is, it is only the unit or the part of the unit which fails the test which has to be taken out of service.

That may mean the whole system in some cases, but it may mean only part of others.  In the case of the 2000D, the A Group tests are tests that the image acquisition part of the system, and so if these tests are failed, the facility may have to stop acquiring new images until the problem is corrected, but as long as the B Group of tests are passed which relate to the interpretation of images, they can continue to interpret old images.

The opposite is true, if a test in the B Group is failed, the facility may have to stop with the interpretation of images, but as long as the A Group of tests are passed, they can continue to acquire and store images for each interpretation.

So, then this alternative requirement applies only to the Senographe 2000D.  As I mentioned, other manufacturers and facilities which own full field digital mammography units or the manufacturers, have the option of applying for a similar alternative requirement and we will consider it and determine whether or not it can be accepted.

These, then, are the three new alternative requirements that we have accepted since the last NMQAAC meeting.  I would be happy to answer or try to answer any questions that you might have on them.

DR. PISANO:  What do the other manufacturers have to do, do they have to apply for the same thing to get, because obviously, this only applied to GE?

DR. BURKHART:  It only applies to GE as you would expect as GE was the applicant.  They would have to go through a similar process, and their lists of tests, depending upon their system, will be different.  They might have different tests, more or less, and are different, that could be given a 30-day correction period.

DR. PISANO:  Just as a comment, some of the things listed under Item C were probably added to the QC manuals of GE at least, and the other companies, because of the MQSA requirements for film.  So, my prediction, I actually have two other units myself, I have GE Fischer and a Fuji system at UNC, is that all these tests listed under Item C are also required by the other manufacturers.

So, it might expedite things a bit if a more general statement could be made about these particular tests rather than every company, I don't know what the process is, but we would be grateful, those of us who are using the equipment, if there could be a more expedited process for this.

I hate to shut down a machine just because one of these things that wouldn't shut down my GE unit, you know, if my Fischer unit had one of these problems, you know, before Fischer had the chance to go in and apply for it.

So, it would be nice if the FDA could make a more general statement about this, just as a user of the device.

MS. HARVEY:  One more reminder.  Before we speak, let us give our names for the record.

DR. PISANO:  Oh, sorry.  I am Etta Pisano from UNC/Chapel Hill.  It is just a suggestion of making this more general.

DR. BURKHART:  Roger Burkhart again.  We thought about that at the time we were looking at the GE application.  As I mentioned, we did expand the second alternative requirement to be broader than the original.

We decided at this stage, though, that since there is still I guess not much of a consensus in general on the testing with the different models, that it would be best to take each one in turn, but once the ground is broken, obviously, you know, it makes it much easier for the people coming along behind if the tests are really the same, it would only be for any unique things that they might have that we have to look at further.

The process I might mention, the process for looking at alternative requirements, actually also is described in a general way, and, in fact, it is right after the listing of the three alternative requirements, the new ones, on page 62, and it basically involves when a request comes in, a staff member is assigned to evaluate it.

Sometimes we see right upfront that more information is needed, so the staff member will ask for it, but if it looks fairly complete, then, we form a committee to look at it and evaluate it.

The committee tries to come up with a consensus, a recommendation being either to approve or disapprove or to ask more information, and then it goes to the branch chief of the Accreditation and Certification Branch, which is the branch responsible for this, to agree or disagree, and then it goes to the division director for final decision.

So, the process is not really complicated and, as I said, once the ground has been broken in that area, it can go fairly fast.

The other point I guess I should mention, too, is that a facility can apply, as well as a manufacturer, so you do not have to wait for the manufacturer to take action to make the case.

MS. HARVEY:  Any comments from our health physicists?  Ms. Martin or Dr. Karellas.

MS. MARTIN:  These standards are fine with me.

DR. KARELLAS:  I agree with Melissa.  The only thing I am a little apprehensive is that physicists just have to watch closer since different manufacturers have different requirements, so which is okay, that they can do it.  Physicists practicing out there are very much into that.  It just will add a little bit on their time for it. That is my only concern, but I am sure they can do it.

MS. HARVEY:  Thank you.

DR. BURKHART:  Thank you.

Open Public Hearing

MS. HARVEY:  Now, we are moving into the open public hearing aspect of our meeting today.

We have comments on quality control for full field digital mammography from Ken Crocker, who is Director of Marketing, Product Planning for the Fischer Imaging Corporation.  Welcome.

MR. CROCKER:  Thank you.  My name is Ken Crocker and I very much appreciate this opportunity to address the committee on what I think is becoming a more important topic as time goes on.

As you all are probably becoming more aware, digital mammography has reached a greater degree of acceptance throughout the U.S., and I think is becoming more of a standard operating practice with probably well over 300 systems already in place throughout the country.

[Slide.]

There is I think a few issues developing as this acceptance has increased, and I thought it would be important to bring to the attention of this committee some of the issues that I think apply to not only the manufacturers, but to the accrediting bodies, the FDA, and, of course, most importantly, the actual users of the equipment, as well.

We are in a situation right now that is rather unique, because when the regulations were really originally developed, of course, digital mammography was basically only a gleam in the eye of most people, but, in fact, we have reached that point that things need to happen now for it.

Because of that, the original regulation only stated that the Operator's Manual from the manufacturer should be followed as the appropriate quality control procedure.  That is going to result in delay in providing oversight to users of the full field digital mammography system.

While it may be true that the FDA, as part of the PMA process, does review the Operator's Manual, and, of course, I think they do a fine job of that because they are the same people that are reviewing the proposed quality standards that certifying and accrediting bodies would propose, so I think we should be confident that they are doing a good job in that area, but nonetheless, it provides only a limited amount of oversight.

Once the PMA has been approved, the manual is in use, FDA will look at quality control charts after six months of use from a facility, but that is kind of the end of the process right there, and I think we want to get to a point where we have uniformity and standards.

So, basically, I will show you what the proposal is, but the issues today, there is a lack of uniformity because you do have each manufacturer proposing their quality standard, and there is limited oversight because it is primarily that review that happens as part of the PMA process.

Dr. Burkhart described the approval of alternate standards process.  Certainly, you know, I think it has its place, but overall it would only be a stopgap measure in this particular instance since we are looking at a complete new set of standards for digital mammography, and I don't think we could rely strictly on that to address all of the needs.

[Slide.]

So, why does this issue linger?  Obviously, hopefully, it hasn't lingered too long, but right now for us, as manufacturers, there is not a tremendous amount of incentive to standardize.

Once we get through that process, which is challenging unto itself, we feel pretty confident that we have produced a reasonable quality control approach, and unless there is some really undue needs, we would prefer to just keep things running, because we want to be able to meet all of our customers' needs.

The MQSA Reauthorization Act of 2002 is not going to substantially change the landscape, I don't believe, at least in the last versions I have seen of it, that it addressed any of the issues that we are talking about here related to full field digital mammography.

The approval of alternative standards doesn't address the needs of accrediting bodies, as well.  The accrediting bodies need to be able to get more involved with this process of controlling digital mammography.

[Slide.]

The proposal is to charge the FDA and accrediting bodies with development of these uniform standards, and to encourage their cooperation.  I know we will be hearing from Dr. Chakrabarti this afternoon on that, and I know there have been starts into this area, but I think there needs to be more urgency applied given how rapidly the acceptance is taking place with digital mammography.

We need to charge the FDA and the accrediting bodies to seek guidance from industry.  Certainly, individual manufacturers are interested and willing to participate, and then NEMA, which, of course, is an industry group, also has started efforts in that area, as well, and would be more than happy to participate in doing that.

Our third point would be that FDA should allow accrediting bodies to accept the Manufacturers Manual for accreditation on an interim basis, and this would allow the transfer of responsibility over to the accrediting bodies, so that they could start becoming involved in this process. As of today, you know, they don't really have any sort of regulatory capabilities in that area.

[Slide.]

Charging the accrediting bodies with simplification and uniformity as a longer term goal, would also then follow with that, but at least we would have that immediate knowledge that the Operators Manual is available and accepted by the accrediting bodies and the FDA.

Eliminate the requirement to maintain a film-screen system.  As I mentioned, the review of the manuals is already carried out by a very respected group of individuals at the FDA, that are also responsible for the oversight of film-screen certifying and accrediting efforts.

If they are capable of doing that, I would expect they would be applying the same degree of rigor to digital mammography, and believe, in fact, that they have, so this requirement of maintaining a film-screen system, I think imposes an unnecessary burden both on users of the systems, as well as the entire community.

Lastly, getting to Dr. Pisano's point about why don't we just do this for all manufacturers, this synchronizing of tests that allow 30 days for corrective action, certainly, we agree and support that position.

Tests that are as common as repeat analysis, I know from our standpoint we tried to make things as simple, straightforward, and consistent as possible for users of our system so that they wouldn't have to completely rethink the process of how they do a repeat analysis, for example.

I think that the FDA should be charged with developing that uniform standard rather than waiting for either a user or a manufacturer to come to them.  We actually did go to the FDA over a year ago with a similar request to what had been made for the GE system, but it is a rather challenging process to get through.

I mean I am not blaming anybody for what happened there, but I would just say that from my standpoint, I think it would be better for all of us if the FDA could take a proactive stance rather than a reactive stance to requests for these kinds of changes.

That is all I have prepared.  I just want to thank you again for the opportunity to address the committee and appreciate your efforts to provide the best quality mammography for our community.

MS. HARVEY:  Thank you, Mr. Crocker.

Any questions?  Yes.

MR. CAMBURN:  Jim Camburn from the State of Michigan.  I think I have one question for you related to one of the things you commented on, eliminating the requirement to maintain a film-screen system.

Are you suggesting, then, that the facility would not have any film-screen unit at the site where they would have a full field digital machine?

MR. CROCKER:  Yes.  Let me explain a little bit about what the requirement is today.  The requirement today is that there is at least one film-screen system within a particular FDA jurisdiction that is under the supervision of a particular radiologist who has responsibility.

They don't have to be at physically the same location, you could have one at a hospital and then you could have a digital at an off-site facility, and as long as there was one film-screen present at one of those two locations under the jurisdiction of a particular supervising radiologist, that would be acceptable.

But we see it all the time now that there are situations where different groups of physicians want to become involved with digital mammography, they have the experience, they are willing to do the quality control that has been approved and recommended by the FDA, but they do not want to invest in having a film-screen system, as well.

In fact, they have no intentions of using the film-screen system, but because of the way the regulations are today, they will go out, they will buy a film-screen system, they will do the absolute minimum to maintain the accreditation of that system or certification of that system, and therefore, I don't think it is really accomplishing much.

I think we are better off letting them focus on the quality control of the digital system that they really intend to use.

MR. CAMBURN:  We see this from maybe a slightly different perspective because we have a number of facilities that have one digital unit and one film-screen unit, and they seem to use them differently, at least some facilities do.  The digital full field mammography machine has a relatively small image receptor compared to the larger film size that you can get with film-screen imaging.

What they do, they will--average size patients might fit fine with the digital image receptor, but larger patients would require two exposures for each projection, and it kind of doubles the patient dose in the area that the x-ray beam overlaps.

So, from a radiation dose point of view, isn't it better to have the ability to do both types of imaging?

MR. CROCKER:  I certainly appreciate and agree with what you are saying.  With the full field digital mammography from Fischer, it has a larger field of view, and therefore, in fact, the larger field of view is 21 by 29 cm, so the percent of the population that would require a multiple stitching together of images is no greater than what would be required under a film-screen system.

So, for our particular equipment, we don't see that problem, but I certainly can understand where you might be concerned about that from a radiation dose standpoint with some other systems that are available in the marketplace.

MS. HARVEY:  Any other questions?

DR. PISANO:  I just have a follow-up comment.  I actually think from a public health viewpoint, in terms of getting digital out to remote areas where the images could be beamed back to a central site for interpretation, it makes a lot of sense to not require the film mammography, because you are probably limiting access to remote areas, if they are going to use digital, they also have to have a film mammogram unit.

So, I would agree with his comments that he made, that we would like to move this process along, I would like to see it moved along, so that it is more standard and that the film mammography isn't required.

MS. HARVEY:  Thank you.

Dr. Finder.

DR. FINDER:  We have one comment that came in, and the person who submitted it would like me to read it into the record.  It is a written statement from Pamela Gormley, who is a mammography supervisor at Epic Imaging in Oregon.

Her statement is as follows:

The following is a mammography item that I believe the FDA needs to expedite the changes on.  We have had two of the new FFDM GE 2000D digital mammography units since October 2000.  However, the FDA says we still have to have a film-screen unit on the premises, plugged in and ready to use, even though a film-screen is outdated technology.

This is approaching two years.  This is wasting both their resources and space for mammography.  We would have replaced that unit with newer technology if the FDA allowed it.  We also have to maintain a film processor that we don't want or need.

All of the quality control tests that we do on the digital units show that they have much better detail on the phantom image and on patients, with one-third less radiation per view than the former state-of-the-art film-screen system that we have.

Our film-screen combo is the detailed Fuji AD-M fine screen with Fuji AD-M film, dedicated Kodak M-35A processor with White Mountain chemicals, 135-second processing at 95 degrees, using GE Senographe DMR bi-metal tube mammogram machine maintained by GE service, but it cannot begin to compare with what we see with the digital system.

Please get this changed immediately, so that we can provide the best medical care to our patients without wasting money.

MS. HARVEY:  Thank you.

DR. FINDER:  I would like to add that we are going to have some more talk about this entire issue later on in the afternoon.

MS. HARVEY:  We are a little ahead of schedule, so if Michael Divine is prepared, we will move on to the open committee discussion.

Michael is going to talk to us on Overview of MQSA Inspection Findings  and Current Inspection Follow-up Actions.

Overview of MQSA Inspection Findings and

Current Inspection Follow-up Actions

MR. DIVINE:  My name is Mike Divine and I work in the Inspection and Compliance Branch in the Division of Mammography Quality and Radiation Programs.

[Slide.]

My talk today is, appropriately enough, on inspections and compliance.

[Slide.]

I will be going over a summary of problems that we have found during our annual inspections and also an overview of the various actions FDA might take when facilities have serious problems or failed to correct these problems.

[Slide.]

For the inspection data, my talk will cover the last two complete fiscal years for FDA data plus most of the current fiscal year which will end on September 30th.  While the data for this year is not complete, I think we have enough data for comparison purposes.

[Slide.]

While most people here today are probably familiar with our inspection levels, I thought a slide was needed for those who might not be familiar with them.

Level 1 is the most serious and could result in FDA action if not corrected.

Level 2 is less serious, but still significant enough that a facility is required to respond to FDA with their corrective action.

Level 3 findings are considered minor.

[Slide.]

As you can see from this first slide, facilities continue to improve and the overall rate of problems has been declining, which is very good news.

While this slide only shows two full fiscal years plus most of a third, if we extended these data back to 1995, when we started inspections, the trend would be even more pronounced.

[Slide.]

This slide shows Level 1 problems with personnel. While the chart shows a jump for some categories in 2001, the small numbers compared to the overall percentage of inspections doesn't indicate that this is a real problem. For the medical physicist, the number of violations has all but vanished.

I would mention at this point that this data represents inspections of approximately 9,500 facilities.

[Slide.]

Processor QC problems continue to be a source of problems, but these numbers are also going down.  The same is true for missing phantom QC data.

[Slide.]

On this slide, as opposed to the previous slides which showed data for the facility QC testing, these data come from our inspector testing.  The number of violations for phantom image is very small, as are data for processor speed.  Fog values are somewhat higher, although these numbers have been declining.

[Slide.]

For the medical physicist surveys and mammography equipment evaluations, the most common problems are overdue surveys, surveys missing specific tests or data, and failures to do evaluations on x-ray units and processors. These numbers are also declining.

[Slide.]

For interpreting physicians, the number of facilities cited for initial training or experience remains low.  For continuing education and experience, the numbers are greater, but are declining.

[Slide.]

For radiologic technologists, we see a similar trend.  The jump in continuing experience is probably due to this first thing checked in 2001.  We expect these numbers will decline in the years to come.

[Slide.]

For medical physicists, the small number of facilities cited exaggerates the difference between the years.  The missing bars for 2002, under the initial requirements, is due to no facilities being cited, as with the technologists, checking on the continuing experience is a relatively new assessment.

[Slide.]

For medical records, a lack of an appropriate assessment category on mammography reports dominates the problems, however, we have seen a substantial drop in the numbers in just three years.

[Slide.]

This chart shows some other requirements we check during inspections.  The problems with x-ray units has dropped to almost nothing.  For our first inspections with complaint and infection control procedures, the drop in the number of facilities with these problems has dropped dramatically.

[Slide.]

For the medical outcomes audit, only a small number of facilities still have problems.  The last three sets of bars here reflect requirements only being checked in the last two years.  As with some of the other cases like this, we expect these numbers will go down with time.

[Slide.]

This last slide from our inspection data shows a number of facilities that had at least one problem during their inspection for not having complete documentation for their personnel.

[Slide.]

Moving away from the inspection, the next few slides will focus on the various options FDA has when facilities have continuing problems complying with our regulations.

[Slide.]

These types of actions include a follow-up inspection, additional mammography review, patient and physician notification, which is actually a follow-up action in case the additional mammography review shows problems, a directed plan of correction, civil money penalties, suspension or revocation of a facility's certificate, an injunction, which is actually a court order that would shut the facility down.

[Slide.]

When facilities fail to meet specific requirements, we may need to reinspect the facility to see if it has corrective problems.  Most of the time, these inspections only focus on areas where the facility has failed in the past.

[Slide.]

Additional mammography review is a review of mammograms and/or mammography reports to investigate previous or ongoing clinical problems at the facility.  The purpose of the AMR is to look for serious problems where patients and physicians need to be notified.  If there was a serious risk, there could be a possible patient and physician notification.

[Slide.]

For additional mammography review, we generally select certain types of issues that we think we want to do an AMR.  One we do which is the most common although it has been significantly declining the last few years is we find a phantom image problem that is at Level 1 during an inspection, we will do an AMR.

We could do one for an interpreting physician that would fail to be qualified.  Clinical image quality problems would be an obvious one.  If there was an overall failure in the quality assurance program at the facility, that could trigger one, and we have done a few for fraudulent recordkeeping situations.

[Slide.]

The extent of an AMR could range from a few films to a larger sample.  Our most common reason for AMR, as I mentioned, is Level 1 phantom failure.  A larger sample is usually needed if a smaller AMR shows serious problems or the problems at the facility to make a smaller review inappropriate.

[Slide.]

When an AMR shows serious problems, FDA would send the facility a letter requiring the patient and physician notification.  These letters outline options referring physicians and patients have, such as getting their mammograms reread by another interpreting physician or getting a new mammogram.  The letters are written in plain language, avoids using complicated jargon with patients.

[Slide.]

A directed plan of correction is a regulatory action FDA may take that imposes additional requirements on the facility.  The goal of the DPC is to force the facility to perform mammography in compliance and allow FDA to easily monitor this performance.

Under a DPC, the facility is usually required to send FDA copies of records on a monthly basis and are subject to additional inspections to check on their performance.

[Slide.]

For more serious problems, FDA may suspend a facility's certificate.  Once a certificate has been suspended, the facility can no longer perform mammography. In most cases, facilities are usually given a hearing prior to the suspension, however, FDA may suspend prior to a hearing if there is a serious risk to human health or other substantial violations.

[Slide.]

A last list of the remaining options that FDA has is rather than shutting a facility down, FDA may opt for charging a facility civil money penalties, and this could be up to $10,000 per violation or per day.

We could also revoke a facility's certificate, which is equivalent to suspension, however, once a certificate has been revoked, the owner or operator of the facility cannot own a mammography facility or operate a mammography facility for at least two years after the revocation.

[Slide.]

Lastly, if everything else fails and we feel that we have to go to court, we have the ability to use an injunction which actually closes the facility down through a court order.

[Slide.]

In closing, this table shows the number of times FDA has taken specific regulatory actions based on problems occurring at facilities.

That concludes my talk.

MS. HARVEY:  Thank you.

DR. PISANO:  I have a question.

MS. HARVEY:  Yes, Dr. Pisano.

DR. PISANO:  Is this in the history of enforcement of MQSA or is this one year?

MR. DIVINE:  This is the entire program.

DR. PISANO:  And what is the denominator like 900-plus facilities per year times 8 years, something like that?

MR. DIVINE:  Well, we inspect about 9,500 facilities.

DR. PISANO:  That is what I meant, 9,500 times about 8 or 10 years?

MR. DIVINE:  Yes.

DR. PISANO:  So, it is 95,000 facilities, and these are the numbers, something like that, is that right?

MR. DIVINE:  Yes.

DR. FINDER:  I would want to add that this represents actions taken by FDA.  This does not include actions taken by the State, and in several cases or many cases, the State has taken action before we have, and in that case, we don't pursue it any further, so it is not the total number of facilities that ran into problems.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  You mentioned something about that the equipment-associated problems are something like very few or next to nothing, which is very encouraging, but I would like to comment for the public and for the lay press, because often we read about that there are no problems with the equipment or it has nothing to do with the equipment.

The reasons that inspectors find very few problems with the equipment is that equipment is very well maintained.  We find problems with the equipment all the time routinely.  Almost on a weekly or monthly basis, technologists will walk in and will find problems with a processor, on occasion with the automatic exposure control, they typically call service or physicist depending on the situation, and the problems are taken care of.

So, this is why you don't see the problems.  I am sure you know that, but the public perhaps doesn't understand that.

MR. DIVINE:   That is a good point.  We only go in once a year to do the inspection, and when we look at the equipment, basically, we find that there was a problem, but it has been fixed.  It certainly is not something that shows up during the inspection.

I would also point out that as the years have gone by, a lot of equipment that had problems and couldn't be maintained has been replaced or repaired to where it can meet the requirements.

MS. HARVEY:  Dr. Lee.

DR. LEE:  Amy Lee.  I was wondering if you ever analyzed your data for specific trends, like geographical areas that tended to have more violations or specific kinds of equipment, and if you have, have you noted any kinds of trends or clusters.

MR. DIVINE:  I am not aware if we have done any geographic types of analyses.

MS. HARVEY:  When we look at the Level 1 phantom image violations, would you talk to us a little bit about the scores that that might represent, what is passing and what scores would be considered to be a serious violation, is it triggered immediately after below 10?

MR. DIVINE:  The criteria we use for phantom image, we have two, Level 1 and Level 2.  Level 2 is where it fails at the accreditation body's limit, which all the accreditation bodies use the same values, which are 4-5ers, 3 speck groups or 3 masses.  If any of the objects go below any of those, it's at least a Level 2.

Now, our criteria for Level 1 is if it goes below 3, 2, or 2, which is one unit below the criteria.  So, we do have a certain number of Level 2 phantom failures, and those are higher than a Level 1, but even those are not very high.

MS. HARVEY:  Equipment has become much better at resolution over the years.  Is there a debate about raising the image score?

MR. DIVINE:  I am not aware of one.  I have heard some people mention that, but there has been no urge for us or, as far as I know, the accreditation bodies to raise the values, but it is possible that there has been, I am not aware of it.

MS. HARVEY:  Dr. Young.

DR. YOUNG:  Don Young.  Have you compared your data with the States that are accrediting bodies and certifying bodies that had the data required relative to the inspections?

MR. DIVINE:  Not that I am aware of.

MS. HARVEY:  Dr. Young, do you have any data?

DR. YOUNG:  No, I don't personally.

MS. HARVEY:  Any other questions, comments?

Thank you.

MR. DIVINE:  Thank you.

MS. HARVEY:  It's time for a break.  It's about 5 minutes of 10:00, perhaps 15 minutes, back at 10 minutes after 10:00.  Thank you.

[Break.]

MS. HARVEY:  Dr. Finder will provide us with information on Good Guidance Practices and Directions for Discussions on MQSA Guidance under the Final Regulations.

Dr. Finder.

Good Guidance Practices and Directions for

Discussion of the MQSA Guidance under the

Final Regulations

DR. FINDER:  Before we begin our discussion of final regulation guidance, I would like to briefly explain the procedures that FDA is following as it develops new guidance.

In response to public comment regarding the use of guidance documents, FDA held an open meeting on April 26, 1996, and on February 27, 1997, they published a federal notice outlining the steps the agency needed to take prior to issuing guidance.

In brief, it stated the following.  Guidance has to be developed in an open manner that permitted input from the general public and the regulated industry.  In most cases, new or controversial guidance had to allow for such input prior to its implementation.

While the statutes and their associated regulations were binding and enforceable, guidance was to represent a way or ways of meeting the regulations, but other ways would be acceptable as long as they met the requirements of the underlying regulations or statute.

Before we begin our discussions, I would like to emphasize the following.  We are here to discuss the proposed guidance, not the underlying regulations.  The regulations have already gone through their own extensive approval process and while they are subject to future change, the purpose of today's meeting is to address the proposed guidance.

When you hear or see words like shall require or must, they refer to the underlying regulation, whereas, the words should, may, or recommend refer to the guidance.  I also want to add that since the beginning of the program, we have issued a large amount of guidance to help facilities meet the underlying regulations.

This material, this guidance has been compiled into what we call the "policy guidance help system," which is a computerized search engine that is now available on the Internet to aid facilities in their compliance with the regulations.

There is probably about anywhere from 5- to 700 pages worth of guidance encapsulated in that search engine and what we are in the process of doing right now is going through all that guidance to update and revise it.

One of the documents that you have, which is Modification Document No. 5, is the first in that series where we are actually going page by page through all the previously issued guidance to update and add material as appropriate.

So, with that said, I think that probably the first item that we would like to talk about is the issue of the agency automatic exposure control.

With that said, I guess we are done with AEC.

[Laughter.]

DR. FINDER:  Let me give a little bit of background.  We did send out a letter to the committee for them to look at prior to the meeting, and basically, this raised several issues about testing of the automatic exposure controls in some of the newer equipment that have multiple different configurations and submodes.

If anybody would like to start the discussion on that, I would appreciate it, otherwise, we are going to have a lot of time between now and lunch.

MS. HARVEY:  Yes, Dr. Karellas.

DR. KARELLAS:  At least I would like to start in one area of the AEC issues.  There are certain systems that they may have various modes and medical physicists may be evaluating modes that they may not be actually used by the facility.

My own view is that there should be no need to test every available mode of a complex AEC system if the facility does not intend to put it to use, and a facility should decide as to what they use, and that should be tested.

Now, I understand that in real life, a facility will start with something and perhaps a month later, they will decide that they need to use another mode, and that will happen.  Although it may be not a problem for a physicist to test these two or three modes and have that, but if they are far more complex than that, and there are too many combinations, it may be unrealistic to be testing all these modes.

Then, the physicist could come back and reevaluate the system a few months later if that had to be, but I am not suggesting that the physicist should evaluate the AEC every time every minor modification is made, and the way it is used or some very minor repair.

I am saying that it should be tested only if there is a very substantial departure from what the system was initially tested.

MS. HARVEY:  Ms. Martin.

MS. MARTIN:  This is Melissa Martin.

I am the other medical physicist on this panel.  Obviously, what we are discussing affects what Andrew and I do the most.  As a consulting physicist, just to put this in perspective, I obviously provide the medical physics services for, at this point, around 150 facilities, which covers around 250 mammography units on an annual basis.

We have many of the high-level, multi-mode, multi-target, multi-filter units in our practice.  We have made great strides to test what I have considered all the clinical modes used for each one of these units when we go on site the first time.

In Southern California, I cover an area I call Southern California.  I cover sites that are about 300 miles away.  It is to my benefit and the facility's benefit to make these measurements when I go out there initially.  That is why, as Andrew said, I try to cover all what is going to be called clinically useful or possible clinically useful combinations.

For those that use, as an example, the GE DMRs, I don't have any facilities that use the dose mode on a 2-cm breast, so it makes no sense to require the physicist to test the dose mode for a 2-cm breast.

i think we do need to set an understanding here of what is clinically useful and what the physicist would be expected to test.  Some of the newer units, the low radium 4's, it is not even possible to test at the low kVp's or the high kVp's for thin breasts, because the grid doesn't have time to move.

So, if you technically looked at some of what is proposed, it can't be.  The other factor is how much time are we taking a facility down.  We do impact the access to people to get mammography.  If we go in to perform our measurements, we are typically in a room somewhere between 4 and 6 hours.

That is 4 to 6 hours that room is out of service and available for serving patients, and I think we have to be aware and very careful not to set measurement criteria that is not clinically relevant, but which will also add cost to the facility and decrease the amount of time that the patients can be examined.

DR. KARELLAS:  Melissa put that very nicely.  I think the vast majority of medical physicists feel that way.

MS. HARVEY:  What percentage of the time that you are testing the equipment do you think the AEC testing would involve?  Is it a major part of the testing?

MS. MARTIN:  At least 25 percent currently.

MS. HARVEY:  And there are no surrogates, there aren't any simple tests that we can go to that would be representative of larger--

MS. MARTIN:  I am saying it's 25 percent to do what I have been considering the clinical modes.  If one of the discussions was pursued here that I had to test every kVp in every mode, you are adding at least 2 to 3 hours of testing, so you are roughly adding somewhere between 3- and $500 of additional cost and another 2 to 3 hours of time out of the room.

MS. HARVEY:  Plus your time.

MS. MARTIN:  Yes.

MS. HARVEY:  Any other comments about AEC testing?

DR. FINDER:  This is Dr. Finder again.  I think you framed the issue.  Now, we have got to get down to some of the specifics, and we did ask a couple of questions in the document that went out.  We would kind of like some guidance on what you think that reasonable testing is under the various scenarios.

The first issue that we talked about here is, as we have said, the current requirement is that a mode or configuration needs to be tested prior to clinical use. There are two different areas where that could occur.  One is at the initial evaluation called the mammography equipment evaluation, and the other is during the annual physics survey.

The requirements in the regulations are slightly different for those two types of testing.  One of the issues that has been brought up is, is it enough to test just all the modes and configurations at the initial mammography equipment evaluation, and then do something lesser of not possibly all the clinical modes at the annual survey.

This is one of the issues that we would appreciate the committee's guidance on.

MS. MARTIN:  I will just respond and again this is where I am coming from.  I basically use Alternative Test No. 2, which conforms to the ACR's suggested forms that are available in the latest QC Manual for the physicists.

That test I perform annually on all the modes with every machine.  I personally have not skipped any of them, so I don't have any feel for what percentage of people typically only test in the contact mode.  I always test the mag mode.

DR. KARELLAS:  I always test the mag mode and most people I know test the mag mode, because we go pretty much by the ACR guidelines, so we use that as a guide, and we may make one or two additional measurements for other things that we feel might be necessary.  That is what we go by.

MS. HARVEY:  Does that help you, Dr. Finder?  No?

DR. FINDER:  Well, it partially addresses some of the issues, but we also have the concept of these units that have multiple different AECs.  As has been brought up, some of these AECs are not used over the entire 2 to 6 cm range.

The regulations, however, say that the AECs have to be tested over that range, and we have some situations where a facility may say, well, we never use the 2 cm range for this type of submode of AEC, but we do use it at 4 and 6.  Well, how is that going to be tested?

We also have the issue of a facility that says they are going to use one submode at 2, another submode at 4, another submode at 6, what is the appropriate testing under those types of scenarios, do you just look at those three individual submodes at those levels, or do you require each one of those submodes to be tested at 2, 4, and 6?

These are some of the questions that have come up, and how do you deal with those kind of situations.

MS. MARTIN:  What we have made the choices, if they are clinically using it, in other words, again, go back to the AOP contrast, the GE DMR has three different modes, well, actually several different modes, but three automatic modes AOP, which is contrast, standard, and dose.

Typically, what Dr. Finder is saying is a 2- or 4-cm breast would be examined in the contrast mode, 6 could be either contrast or standard.  I typically test 2 and 4, I don't test 6 under contrast unless the facility says that is what they use.  A standard, I typically do 4 and 6, I don't do 2 unless the facility says that is what they use.

I really do think you have to look at what the facility is using, but if those modes are set up and cross with each other, that basically is your test.  I think it is crucial that each mode be tested at least on some thickness to verify that you are tracking between modes.  I don't find it mandatory that you test every thickness on every mode.

MS. HARVEY:  So, if you are testing a unit and you look at either the 2, 4, or 6, and it is working properly, can you presume that it would be working properly at the other two thicknesses that you are not testing, or do you actually have to do each thickness?

MS. MARTIN:  I have not found a problem.  If it is working properly in the mode for which it is pretty much designed to work in, I have not found a problem with it tracking between the other modes, but certainly I am not the only one performing these measurements.

Dr. Karellas.

DR. KARELLAS:  We actually test at higher thicknesses than 6, but it is not a requirement, but we just do it because we want to see how the machine works, and we just draw a line there, so if it deviates, and sometimes they do, that then we know at least that we have bracketed for the requirement.

Usually, the deviation may be very nominally below or above what the requirement is, but we know that the deviation is not really against any regulations, state or federal, but we are aware of it.  If we see something and we need to adjust it, at least we know.

MS. MARTIN:  One of the more crucial items as far as time testing is for those units that are now out there, that have individual detectors, there is one manufacturer that has seven or eight separate detectors, independent detectors.

I think the more crucial time thing is I have found it and what I have been doing is doing the full set of tests on one of the detectors and then cross-checking all the other detectors for the 4 cm breast.  I have found that to be sufficient.  I have not been testing all eight detectors for every target filter for every one.

I think again, as long as you are making a reasonable attempt to verify at typically the 4 cm thickness, that your detectors cross with each other, that should be considered an acceptable test.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  Isn't it true--this is a question for the physicists--isn't it true that if there were a problem with one of the AECs, for example, in that eight system, it would be obvious in the clinical images?

Wouldn't it be that it would be either too light or too dark over a certain region of the breast, so that it is not something that is likely to create real clinical problems?  In other words, the radiologists, the readers, and the local physicist, if there is one, would be able to spot it very quickly?

MS. MARTIN:  You are going to spot it very quickly, right.

DR. KARELLAS:  It is true that eventually, it will become obvious when an astute radiologist or technologist will discover it, however, the concern is that there will be certain studies that will be done, and the patient will be gone, and images may be suboptimal, so eventually, it will be found, it will not go for very long, but it is just that there may be something compromised, perhaps not of great significance, but certainly mammography may not be done at the state-of-the-art level.

MS. HARVEY:  I have another question.  Do you generally find a problem on the initial testing, or is this a problem that occurs over time, that you see that the detectors go out of whack on your annual when you come back, do you see problems at the annual testing, or do you see more problems at the initial testing?

MS. MARTIN:  The initial testing is usually good. I think it is absolutely crucial, though, that it be tested, at least sampled, again cross-check some way for each one of those initially.

My experience is the installations are usually done very well now.  I wouldn't say that was necessarily the case two years ago, being the installers have gotten much better, and I think part of that is because they are trying to make the criteria that has been set.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  The question that I have is what we should be doing when we have a brand-new machine and it has all these multiple modes, and we have not discovered how we are going to be using that machine, and the site needs to accept it.

If you accept only two or three or four modes, and then six months later, you discover that you want to use all the others, and you find something that doesn't work very well, then, it is more difficult to go back to say that that was not done properly in the beginning, especially post warranty.

So, it is somewhat of an issue as to whether we need to test absolutely everything upon installation, but that can be a very frustrating experience because you are testing something that may be so far out of the real application.

So, I still maintain that upon installation, the site should define certain modes of use and perhaps one or two or three above and beyond that based on how this machine should be used, and perhaps after the first or second year, you could perhaps narrow it down a little narrower to say that we are never going to use these modes, we are only restricted at just to these three modes.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  I have another question for the physicists.  Isn't it the case also that your phantom testing on a weekly basis would suggest there is an AEC problem, you would see your OD changing over time even with the same settings?

I am trying to get a feel for how dangerous this is for patients.  My sense of this is that it is not very because of these two things, these clinical images will chance and the phantom imaging will change, so even if you don't check every mode, even if they use it one or two, you are going to find it somehow.  But I would like to hear your comments on that.

MS. MARTIN:  Again, as long as you cross-check the modes, I don't think you are going to have a problem.  I have not found it a necessity to check the complete thickness for every single mode, and, yes, the idea of if you wanted to extrapolate it for those instruments that have eight detectors, do you want eight phantom images every week to verify that you have consistency.  You could take it to the nth degree and make the same requirement, and I don't think any of us want to go there.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  I agree with Dr. Pisano.  First of all, I do not feel it is dangerous and I feel very strongly that weekly phantom tests are very good, and technologists, in fact, in most situations, it is technologists, that they call us about problems, and they are very vigilant about image quality.

On the down side is that if something ever might happen that shows okay on the 4.3 cm phantom, but it doesn't track very well when the thickness is 6 cm, that is possible, but at least in our experience, most of the time technologists call us and they say there is something wrong with my AEC, and interestingly, you may have tested it three months ago and everything was fine, and they alerted us.

So, what radiologists and technologists see every day is extremely critical.

DR. FINDER:  One point that I would want to bring up in terms of the phantom testing, the phantom is one image taken under the clinical conditions for that thickness of breast, so you will not be checking or necessarily have any idea how the other submodes might operate in terms of the AEC.

In fact, depending on how a facility sets up its protocols, if it does, their standard patients, say, in the typical AEC mode, and not the full auto mode, you won't have any idea what is going on with the full auto mode in terms of looking at the phantom.

But these are the issues that we would like discussed and try to come to some kind of consensus.

MS. HARVEY:  Mr. Camburn.

MR. CAMBURN:  Maybe I can just relay some of the information that we get from our inspectors from time to time about this, especially in terms of your fourth question that you ask here about testing submodes 1, 2, and 3, when submode 1 might only be used with a 2 cm thick breast and submode 2 with a 4 mm, and so on.

What we find from time to time is the technologists will inadvertently use the wrong submode especially if you have a number of technologists working with the same equipment, they don't all seem to be on the same page at the same time, plus there are patients whose size fall between 2, 4, and 6, and the technologist makes a judgment and may sometimes judge to use a different submode than what may have been initially assigned for that thickness.

So, we kind of like to see, although a reasonable amount of testing, that the submodes are all tested.

MS. MARTIN:  That is why I made the comment earlier.  I think each of the submodes needs to be tested at least at some thickness to show that it tracks, but if all the submodes are tested--and I would add to Andrew's comment, we always test an 8 cm breast because at least in Southern California, we have several women that fall in that category--so, I think the 8 cm breast is absolutely critical to be tested.

Now, what I have found is for many of the 8 cm breasts, the 8 cm phantoms, it is necessary to adjust the density on some units to achieve the optimum density, and that option is nice to have, and I think that is part of the physicist's responsibility to give the facility a technique that will bring their large breasts into the same density range as their average breast.  That is part of working with the facility.

As long as I have tested those modes over some part of the thickness and they all meet the tracking, I have not found a problem with testing every mode for every thickness.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  I guess I want to reiterate Melissa Martin's point about the cost to facilities of having machines down for longer periods of time than are needed.

I think the reason why I kept going back to the point of patient safety is I think that maybe one could make a case that at acceptance, all the modes and all the thicknesses should be tested and then maybe after acceptance, then, only the ones that are clinically used should be tested routinely.

In that way, I really think you are probably doing the maximum at the beginning and then you are not going to hurt patients or I want to also echo Andrew Karellas' point that the technologists are really right on top of those, when the AEC drifts out of calibration, we know about it pretty quickly, so I don't see that there is a practical real problem.

It is more because the regulations say it, you have to figure out how to do it problem.  In reality, we are on top of this AEC, and we know when it is not working properly in clinically relevant modes.  So, I feel like we should probably try to make it as supportive of the regulations, but not hugely time-consuming for the facilities.

DR. KARELLAS:  I think we should also realize that the way we are testing it, at the various modes and the various thicknesses, we are using Lucite, which is really a structureless material as far as x-rays are concerned, and the sensors, that is not what they see.  They see very inhomogeneous density.

So, we can be testing some of these things forever and never really reaching perfection as far as matching it to the anatomy, and we must realize that.  There is a point that when we go above and beyond, we get diminishing returns.  We just do not get much better image quality.

I do not want to de-emphasize the importance of the proper exposure.  There is no question that with film-screen, the correct exposure is one of the most critical aspects of a good mammogram, but I think it can be done without going far above what we are doing today on testing the AEC.

DR. FINDER:  I just have an attempt at clarification here.  Suppose a situation occurs where in the beginning, a full testing of all the equipment modes as best as possible was done for the evaluation.

Then, the facility decides that they are only going to use, let's say, two of the submodes, contrast and let's say dose, whatever, and then you do your annual survey testing those submodes, but sometime after your survey they decide that they want to use a third submode.

In your opinion, would that require you coming back to retest before they could use it on patients?

MS. MARTIN:  Not if you could actually give a cross-check, and I think that is where you would fall into the medical physicist oversight.  What my advice to a facility would be is have the technologist on site shoot a phantom in both modes, and if they cross-check with each other, so you could calculate a dosage, and the dosages are reasonable and the technologist is trained to read out that phantom image, and so is the radiologist, if I get feedback that that is acceptable, that would be fine with me.

Frankly, as far as testing all those modes, what you find is for 6 and 8 cm, standard and dose pull the same thing anyway, so you really only have to test contrast and standard, because after about 6 or 8 cm, they are all going to pull the high kV and high filters anyway, there is not that much difference in them.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  This is one point that I am not sure I am in total agreement with Melissa Martin.  Perhaps I don't understand or perhaps she has conducted some experiment on cross-modes, and there are some data that we should look at.

If, for example, we have tested something all in the contrast auto mode and somebody all of a sudden switches to the dose mode, we do have a very different situation in the equipment, and I am not certain that the system would behave the way that we would want to.

Now, I wouldn't be surprised if it does happen with certain machines and when you test it across all modes and somehow everything just clears through and everything is fine, but if I get a call as a medical physicist, and they ask me is it okay to do that, then, I would have to ask them to conduct the measurement.

But then they are performing a measurement that I should be performing, and we run into a gray area although theoretically, it is possible that somebody can send you three, four images under these conditions that you prescribe and measure the optical density and generate a report and say that, yes, it actually does conform versus going there on a visit and checking it on your own.

MS. HARVEY:  We have Question No. 2.  I refer you to the document on AEC testing, the second page, at the bottom, which has to do with, "Since some or all of the AEC configurations may share key components or algorithms, is it reasonable to assume that the failure of one configuration immediately makes the other suspect unless the cause of failure in one configuration can be isolated as unique to that mode.  In that case, only the manual mode could be used as back-up until repairs have been made.

"An example of an isolated configuration failure would be a system that incorporates separate AEC detectors for different image receptor sizes.  If one detector fails and can be identified as the cause of failure, then the continued use of the AEC with other image receptor would be appropriate."

MS. MARTIN:  I think that comes back to Dr. Pisano's point a while ago.  If you have one of these instruments with eight detectors, and suddenly you find out that one of these is off, it doesn't necessarily mean all of them are off if you can verify that you move it to a different unit and it performs fine, then, obviously, you make a note and post something that says one unit is not usable.

The same think would come down to the Siemens unit.  You could very well have the large bucky fail or the small bucky fail, but you wouldn't necessarily fail both of them.  Obviously, that can happen.

Frankly, I don't have any facilities that, as the physicist, I allow them to use the manual mode of exposure. If my AEC failed, they are down.  There is no way we use a manual technique for anything, and I think that is your bigger--the idea that you are going to allow screen-film mammography in today's world to be performed with manual techniques is out of date.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  I totally agree with Ms. Martin. There is no way that I could think of mammography going on in a facility, going on in an manual mode.  If the AEC fails, they are down, period.

I would be interested to know whether anybody really would do that, but we don't, and I think most of my colleagues would not allow that to continue, and I think most technologists would stop.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  Just to clarify, because I don't understand, may have misunderstood, you are not saying, however, like for the Siemens unit, where there are two separate AECs, if one of them was down, the big image detector, but the smaller, you would still go ahead and allow imaging with the smaller detector?

MS. MARTIN:  Correct.

DR. PISANO:  Okay, because that is the way we do it at our place, and it seems appropriate to me.

MS. HARVEY:  Is this a frequent problem, that sensors go down?

MS. MARTIN:  No, I don't hear it that often.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  We have had several problems with AEC on several units.  In some cases, we identify them, and in some cases, the technologists would call us.  When we say "frequent," it is not too frequent, but if you have 10, 20 or 25 mammography units, one of them, in two or three or four years, something will come up.

MS. HARVEY:  Rather infrequent.

Dr. Pisano.

DR. PISANO:  This is a question for the physicists.  In my experience, the AECs tend to drift a little as opposed to totally failing, so you notice on your phantom that the OD is changing, either going up or going down.  That is what you notice as opposed to it's not working at all and the clinical images are really terrible, it's really just a very gradual, is that correct, is that what generally happens?

MS. MARTIN:  That has certainly been my experience, and that is why you do QC every week, and that is why you track those phantom images, and that is why you have PMIs on the machines, is to bring them back into your desired range.

MS. HARVEY:  Any other comments on Question No. 2?

Actually, I think we have sort of answered No. 3 since we have talked about manual mode.

"In the event of AEC failure, the manual mode may be used for up to 30 days while the AEC is being repaired." I sense that that is not what the panel is recommending.

MS. BUTLER:  Could I ask a question from the floor?

MS. HARVEY:  Certainly.  I recognize Ms. Butler from the floor.

MS. BUTLER:  This is Penny Butler from ACR.

I would just like to ask for clarification on the document that was provided.  I think it may assist the discussion that is going on.

AEC failure, what exactly does that mean?  Does that mean that it fails the physicist test or does it mean that it just doesn't work, because I think a clarification on that point may sort of influence how the discussion goes.

The other question is what is the definition of manual mode, because in the current guidance that is out there, there was a discussion of if the AEC performance fails one of the performance tests in full auto, it would be appropriate to temporarily use the fixed kVp AEC mode in order to continue operating.

I would like FDA's interpretation of how this plays into this discussion.

MS. HARVEY:  Thank you.

DR. FINDER:  I guess it plays into the discussion in the sense of as we are trying to get a handle on the fact that these units have multiple AEC modes, and if you can figure out that only one or two or three of these modes are affected by whatever problems is causing it to fail a test or to cause problems, but the other remaining AEC modes are not, then, obviously, you could continue to use those other AEC modes.

If all those fail, the regulations do allow a manual technique for up to 30 days.  Again, that is the way the regulations were written, taking into account the guidance that were received at the time the regs were written.

But the facility certainly has flexibility in terms of if they have a functioning AEC that is within the limits, and they can have confidence that it is, they can use that if their other AEC modes fail, for example, the full auto mode fails in some manner, they could use the fixed kVp AEC mode and continue on that basis.

But you do raise a good question of how do you know what it fails, what is the definition of failure. Obviously, there are multiple definitions here, and I think that they each raise their own issues.

There is the failure that occurs during the physics testing, and then there is the failure that occurs clinically when somebody suspects that there is a problem and what do they do in that case.

Generally, what obviously we would recommend is if they believe that there is a problem, they get their physicist and take a look and see what really is going on, so that they do have a better understanding of what is failing and what isn't.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  In our experience, failure may be gradual, as Dr. Pisano described, you see some drifting and the technologists may catch that before anybody else.

The other mode of failure is when, on the annual testing, that it does not track with thickness, and we will notice that it is slightly off, and the other mode of failure is when you get a call from the technologists and they tell you that something is just very erratic.

This is not unusual, that they will tell you that it works the first 30 minutes in the morning, or if I take an exposure 15 minutes after I turn the machine on, it doesn't work very well, and then it sort of behaves somewhat better.  That is somewhat of an erratic mode, and the technologists pick it up.

On the other part about the AEC, this is an automatic mode, so switching to more conventional AEC, as Ms. Butler indicated, fixed kVp, automatic exposure control, this is a form of an automatic mode versus going all manual.

MS. HARVEY:  Ms. Martin.

MS. MARTIN:  I would totally agree with what has been said, that would be the first option, if one of the full auto, auto mode fails, you would go to the next level down, which is the manual section of the kV and target and filter.

Again, I would come back to that is why initially, we do check, at least cross-check for the phantom with all the modes and make sure all the modes are functioning properly, which certainly allows the facility that option, and if they lose the auto-auto mode, they can very well use the phototimed mode, and that has already been checked out by the physicist and it is ready to go, and it won't take them down.

I don't consider the single auto mode as a manual technique.  I was thinking of manual as totally manual where that technologist is setting the exposure.

DR. FINDER:  That is the correct interpretation of that.  AEC mode, the fixed kVp is an AEC mode, it is not a manual mode.

MS. MARTIN:  Yes.

DR. FINDER:  Let me also ask this, follow up with this.  You do the cross-testing both on the initial equipment evaluation and during each of the annual physics surveys?

MS. MARTIN:  Yes, I do.

DR. FINDER:  If a physicist didn't do that, would you say that if they hadn't tested it during the annual physics survey, at least at the 4 cm cross-check level, that if the facility wanted to switch to one of those modes and it hadn't been tested, the physicist would have to come back out and do that testing before it could be used clinically?

MS. MARTIN:  If they don't have the data available to calculate a dose and image quality, some type of check I would think has to be made.  I would think you would have to at least shoot a phantom.

Not all facilities have the 2, 4, 6 cm Lucite to test, so you are either going to have to have a physicist on-site or some acceptable procedure previously outlined that the physicist is willing to accept.

That could be the medical physicist oversight is what I am coming back to.  If you have checked it out, so that you say in the auto-auto mode typically pull a 26 kV, and your mAs is 143, and you shoot it in phototime at 26, and it shoots 145, that should be perfectly acceptable, but you do need that cross-check done before you are going to use it on a patient.

DR. FINDER:  Just to clarify things, right now the regulations, as interpreted, as written, require that the physicist come on-site.  What is being mentioned here is the possibility of doing this kind of remotely through physicist oversight, which would be a modification of what we have right now.

That is what you are proposing or recommending or suggesting?

MS. MARTIN:  Yes, I am suggesting.

DR. KARELLAS:  I don't think this is unreasonable for providing a set of data for the physicists under specific conditions if it is needed remotely to advise the facility on something like that, on cross-checking, however, it raises the question whether the physicist should be doing one back-up mode on the annual inspection, because there are many physicists that they will do only one.

They will ask the facility, what do you use, and they will use contrast auto all the time, and they will evaluate the contrast auto, and that's it.

If the contrast auto does not work, then, you don't have any data on the other mode, so you don't have a back mode, so the option is to either have an evaluation done there, and you can tell them go ahead, you can switch to the other mode, we have the data and your other modes would work.

But the question is how do you know now that the other modes would work?  If one mode doesn't, how can you assume that the other modes would work.  In all fairness to the patient, we do not know.  So, somebody has to do something at that point.

Now, in the more real world, a technologist is going to tell me on the other side, I have been using contrast auto all the time, I am not going to switch now to select the kVp.  With three or four technologists doing that, they are going to get all confused, so chances are they are going to tell me I am calling Service right now and we are stopping.

I think they would be very unwilling to just go and do all kinds of things because they would be afraid that they would be doing the wrong thing.

MS. MARTIN:  I guess maybe I have got technologists that would have no problems with that.  I think it strictly depends on the facility, and I think that has to be part of the medical physicist's understanding and agreement with that facility is when they are going to be called and what they allow the technologists to do.

MS. HARVEY:  Have we completed?

DR. FINDER:  I just want to clarify, in the fourth one, where I think we had already gone over this, about the testing of the 2, 4, and 6 cm.

I just want to clarify in my own mind the consensus or at least some of the comments were that you would test the 2, 4, and 6, but only at the submodes that were used at those levels.  Is that correct?

MS. MARTIN:  That would be my interpretation.  I think that is the suggestion, that is certainly the training the technologists are given when they are given their clinical training, that it is never suggested that they use the dose mode for a 2 cm fatty breast.  That is part of any technologist's understanding is the appropriate mode to select for the type of breast being examined.

DR. FINDER:  Just again to clarify, let's say the contrast is used at the 2, and the standard was used at the 4, and then all of a sudden they wanted to use the standard at the 2, any additional testing required or no?

MS. MARTIN:  It would depend on whether it is a new unit or a reevaluation unit.  If I have checked it at 4, and it crosses its contrast at 2 and 4, and the standard crosses at 4, if they want to shoot standard at 2 and the techniques are reasonable, it is probably going to work.

It hasn't been my experience that that would necessarily be a problem.

MS. HARVEY:  Any other comments?

DR. FINDER:  I just wanted to check.  Anybody from the FDA side have any questions that they would like asked, or any items?  Speak now or forever hold your peace.

MS. HARVEY:  Or any other members of the audience?

I refer you to the Modification Document No. 5, Guidance for Industry and FDA.  There have been quite a few changes in this document, which has been out for a while now for use by individuals.

DR. FINDER:  I just want to again bring what this document represents.  It basically took the guidance that we had already issued on these items, and what we are trying to do is update and modify what needs to be changed.  So, that is why you are looking at a lot of issues that deal with the same type of topics, such as accreditation and certification.

If anybody has any comments about the changes, these actually have already been published and are out to the public.  They are up on our web site.  If anybody has any items that they would like to discuss, now is a good time to do it.

I would also mention that there is quite a long, new section dealing with full field digital units and their certification, and explaining what we are doing presently with those types of units.

MS. HARVEY:  Does anyone see anything?  Yes, Dr. Pisano.

DR. PISANO:  I just wanted to talk a little bit about the digital requirement, the digital pages 32 through 37 or so, 40, I guess.  This ties in with what I am going to talk about a little bit this afternoon, so I don't know if you want me to wait and talk then or you want me to talk now.

MS. HARVEY:  Well, give me a hint.

DR. PISANO:  There is a currently active clinical trial going on for which I am the PI, called the American College of Radiology Imaging or Digital Mammographic Imaging Screening Trial, otherwise known as D-MIST, and we actually have a fair amount of data at this point about what tests, you know, we have been doing the manufacturers' recommendation as is required under MQSA, under this law, plus we have data on other tests, plus we have been doing it for quite a while.  At this point, we have 19 open centers and the trial has been open since October of 2001.

So, we have a lot of information, and I would encourage the FDA to move forward on kind of looking at the data that exists and trying to perhaps pare down the requirements over what is in the manufacturers' documents or their user manual, whatever it is, whatever you want to call it.

I am concerned about spending a lot of time doing tests because we have always done them on film, they may not be appropriate for digital, and it is just that time is money and if we don't need to do it, we probably shouldn't have to do it.

The reason I bring this up, it is going to be presented publicly at RSNA, the Radiologic Site of North American meeting this November, and Martin Yaffe, out of the University of Toronto, is the PI of the quality control piece of the trial.

I feel that once this presentation takes place, there is going to be more pressure on FDA to kind of respond and maybe cut down the requirements, so I would like to see us kind of be proactive.  I want to echo what Ken Crocker said from Fischer this morning in his public announcement.

I would like to see us kind of have more specifics kind of and detailed recommendations for QC for digital as opposed to just what the manufacturers recommend.  I think there are many reasons.  It is in the interest of patients, I think, just because of the amount of time that we spend doing it does cost the facilities money.  It is obviously in the interests of the facilities and people like me who run facilities to kind of try to keep the requirements to a minimum.

Understandably, they have kind of mushroomed into a big set of requirements because the companies just didn't want to leave anything out that the FDA might want them to put in, but I think we now have really kind of--I am not prepared to discuss today in my talk what things should be cut out, but I know that we will have, you know, this thing hardly ever drifts or once in a blue moon drifts or never.

So, I just want to encourage the FDA to kind of perhaps talk to ACR.  I am talking about the American College of Radiology Imaging Network, not the ACR mammography presentation program, now there is two separate entities, about what is available, perhaps even to hear what Martin has to say before the talk in November.

DR. FINDER:  Yes, we would appreciate any access we could get to that information as soon as possible.

DR. PISANO:  There is actually a public meeting in Washington, D.C.  Actually, it is in Arlington, Virginia, at the Ritz Carlton in Pentagon City in October of the American College of Radiology Imaging Network where I am sure we will all get a glimpse of the information.  I can't say who is going to give a polished RSNA presentation at that meeting, but I am sure we are going to hear about this at that meeting.

So, I will inquire of the ACR Imaging Network folks if you all can be invited to listen.  It is a public meeting, so you certainly are welcome to come to the public sessions and perhaps meet with Martin there and the other physicists.

All the physicists for all the sites will be present or at least they are all supposed to be present at that meeting, so I would expect a very useful amount of information, you know, you will be able to get a lot of interesting and useful information at that setting, so I would encourage you to attend.

DR. FINDER:  I do want to kind of put this into perspective.  The regulations dealing with full field digital, when they were written, obviously, there weren't any digital units that have been approved yet.  We took as has been stated before the conservative approach and said that without any data, we would rely on the manufacturer of the equipment to establish a quality control system that would be adequate for their unit.

I think the idea has always been that as more information became available and the ability to kind of standardize the quality control for these units was developed, that that is what would happen, but until we get enough data available, it is going to be difficult and as we just heard, there is some data that is going to become available soon and as soon as it is, we are going to certainly want to take a look at it and see if we can progress along that frontier.

Another issue that has to be kept in mind is that some of these digital units are quite different from each other and that the quality controls that might apply to one unit may not apply to another.  That also is an issue and the ACRIN study will be dealing with a lot of different units and hopefully will be able to provide us with enough information, so that we can start formulating the ideas for a standardized quality control system.

I am sure that this committee is going to be directly involved when that information becomes available and guiding us in terms of what we would require.

DR. PISANO:  Just to follow up on that point, we will have data from D-MIST, as you mentioned, for manufacturers, so we will be able to compare the need for different tests for each manufacturer, a very rigorous quality control program centrally monitored also, which is one of the strong features of it.

It is a little stronger than what MQSA does because FDA's inspections are an annual snapshot of what happens.  This is literally being monitored by central physicists every week, so we can watch for it because it is so important in the trial to be sure we have the highest quality images, because we don't want people to question our results at the end as has happened in other clinical trials.

We want to really be sure, and so this is being monitored very, very carefully by physicists every week, so I feel that we are going to have about as comprehensive data on this topic as you can get.

MS. HARVEY:  Very good.  Thank you.

Dr. Karellas.

DR. KARELLAS:  There is no question that they did some mammography units very different from each other, and it will require different modules for physicists and technologists who are involved in that.

I find it somewhat more difficult when I deal with the manual that comes from the manufacturers although it is welcome that they have that, but I believe medical physicists will find it a lot easier if they have to deal with modules from an accrediting body, such as the American College of Radiology, and I believe that it is a lot easier to communicate, a lot easier to ask a question.

We can always call ACR for what we want or hopefully, they feel an obligation that they have to do it. Frankly, the companies are excellent and they will give you what you want.

I find it somewhat more difficult to have to call each company to find each person, if I have a question or the particular book or the version, the accrediting body provides more of a centralized area that I can deal with, who feel absolutely obligated that they have to provide that information that I need.

So, if we work in that direction, and if manufacturers cooperate to provide the information they feel that is very important, that their systems need to be tested on, and that can be worked into a document that parallels the existing ACR guidelines that we have.

Let's not forget that quality in mammography, at least from the technical point of view, has been achieved to a large degree because of the uniformity that we are able to achieve in quality assurance through these manuals that radiologists, technologists, and physicists have, and we refer to them all the time.

MS. HARVEY:  Thank you.

Any other areas for comments?

DR. FINDER:  If nobody else has any questions about the guidance document per se, there are some other guidance issues that have come up, I just wanted to bring them to the committee's attention.

MS. HARVEY:  Hold one second.

Ms. Gilbert.

MS. GILBERT:  I do have a question.  This is Alisa Gilbert.

In the documentation here, under Breast Implants, I know that there is not a whole lot written in this, and for a lot of the native patients that I work with, this is a new procedure that is being offered now, but I also know that there is not a whole lot of information when it comes to following screening after that on the remaining breast or even doing mammography with implant or not even implant, but other types of procedures have been done on tram flaps or some information like that.

I would like to see some documentation or additional information in this body written on that, as well, on the procedure and protocol that should be developed for that, just to include it.  I know that it wasn't even included in this, it just said implant, but there is other procedures that are now being taken that haven't been.

Walking into the procedure and asking for follow-up, it is completely an unknown, and I just think that that might be something that might be brought to the attention.

DR. PISANO:  Can I ask a question?  I am a little confused about what you are talking about, and I just want a clarification.  Are you talking about women who have had implants placed post-mastectomy for reconstruction purposes?

MS. GILBERT:  Yes, and the follow-up for that.  I know that in here, it just says, on page 63, there is a question, "Is there a specific amount of training or number of mammograms of breast implant patients that the technologist must perform under direction supervision prior to performing these studies?"

It is requiring that 40 hours initial training for that procedure.

DR. PISANO:  I just have another question.  I am not sure what extra procedure you may be referring to for patients who are post-reconstruction.  I am not exactly sure.  You said there are new procedures being offered, and I am not sure what you are referring to.

MS. GILBERT:  I just see the notation here alternative requirements or breast implants.  Implants aren't the only procedures that are being done for women that have received mastectomies, and I was just kind of interested to know if that is going to be something that is going to be another addition to this, for women that do have just like one--maybe I am not posing my question or my concern clearly.

DR. FINDER:  I think at the time that these regulations were written, there were two major issues that were being discussed, and I think that the idea of the implants here was more the cosmetic implant use.

I think you bring up a very good point about additional training for patients that have undergone surgery and the correct procedures on how to do those examinations.

That is an issue we may be able to deal with in some way through guidance as a recommendation, but in terms of the regulations, what we have got is what we have got, but I do think that we do have the potential to expound a little bit on the guidance and deal with some of these other issues that you do bring up, and I think that is a possibility, and if not directly in our guidance, then, referring people to other sources where they can get the correct procedures to do some of these patients, so I think that is a very good idea.

MR. GOODE:  I am Claude Goode from the State of California.

I would like to relate a horror story.  I had a patient that complained of a breast compression that had an implant, and the compression ruptured the implant.  The horror story was that the woman literally was screaming at the technologist to release her, and she would not release the compression.  Evidently there was significant compression applied although the machine was within normal operating standards.

This patient has been left to suffer.  There were no basic regulations that we could enforce, and I do believe that there is a need for FDA or someone to at least come forth with some standards for the mammography of patients with breast implants.  This needs to be discussed at length and in depth, and presented somehow for the technologists, the physician who is not present, and this does present a major problem.

I would just like to relate that horror story.

MS. HARVEY:  Dr. Harrison.

DR. HARRISON:  Miles Harrison.  May I ask a similar question?  This is someone with breast implants for cosmetic purposes, this is not post-mastectomy?

MR. GOODE:  That is correct.

MS. HARVEY:  Dr. Ikeda.

DR. IKEDA:  This is Debra Ikeda from Stanford University.

I think we are talking about two different things here, I would like to clarify.  That is a terrible story.  I would like to first address the question about the post-mastectomy patient and the tram flap patient.

I think if I understand correctly, you were discussing patients who have undergone a mastectomy and have either had latissimus dorsi flap reconstruction or a tram flap reconstruction with autologous material from the abdomen, or patients, for example, who have a mastectomy and may have a small amount of residual tissue, and the recommendations for imaging that.

There are various amounts of scientific literature for stating that either it is not recommended, for example, there are some articles in which patients have had mastectomies and then looked, they have tried to find out if there is a breast cancer recurrence, because that is what everybody is concerned about, and there is varying data on that.

Many places state that you should not be doing those patients routinely for screening, but if there is a lump, then, special views are often used for that, and every patient is so different, that I am not really sure that if there is a problem, that you can actually say that this is the right view to do or that is.  Oftentimes, we have to come up with special views to address that specific patient's problem.

So, it is important when the technologist is initially trained in their 40 hours, for example, I think guidance is--correct me if I am incorrect--but guidance states that the technologist must learn all of the views, as well as patients who have implants, and they must be doing patients with implants under supervision, so they do them correctly.

I think we are talking about two different things.

DR. FINDER:  Let me just correct one thing.  The regulations do require that technologists who qualify under the final regulations have some training in doing patients with breast implants.  There is no requirement that they have to do a number of breast implant patients, though.

DR. HARRISON:  I am in complete agreement with Dr. Ikeda, we are clearly talking about two separate populations here, and actually, the standard of care in our setting is such that we don't do imaging routinely of people post mastectomy with either autologous reconstruction or implants.

I guess I need to ask the question, are you referring to the training of the technologist to, in fact, be able to do mammography on women who have had cosmetic breast implants, are you addressing the training?

MS. GILBERT:  I guess I am addressing both.  I know in the native population, like Alaska Native specifically, that follow-up isn't recommended after that.

DR. HARRISON:  Follow-up is not recommended?

MS. GILBERT:  It's just unknown, it is just one of those unknown procedures, and because of these procedures, it really is now, just in the last year, that it is being offered for native patients to have the reconstruction services and those provisions.

I know that there is just not a whole lot of clarity when it comes to that, in that specific population.

DR. HARRISON:  It is very cloudy, because presumably, in that setting, in that clinical setting, you have retained all the natural breast tissue, you simply have implants placed, so that is someone who clearly would need to continue to be screened age specific.  I am still confused.

MS. HARVEY:  Dr. Kopans.

DR. KOPANS:  Dan Kopans, Boston, Massachusetts.

I just wanted to reinforce what you just said, Dr. Harrison, that the standard of care is to not do routine mammographic imaging of women who have had mastectomies on the mastectomy side.

There are only two papers that I am aware of looking at women who have had tram flaps, and there is some debate about whether it is valuable when they have had trams, but women who have implants on the mastectomy side, if it has been a simple mastectomy or modified radical mastectomy, there are no data supporting doing mammography of the implant side.

As I recall, in the discussions years ago about implants, as Dr. Finder was pointing out, really, the implant discussion had to do with augmentation, and not reconstruction.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  I just want to make a comment in reference to the story from California.  I think that was obviously, a tragic event, and horrible thing for the patient, but I feel like there really is a lot of guidance from professional societies both for technologists and radiologists, and probably physicists, too, about the imaging of breast implants.

This particular technologist obviously made an error, it sounds like to me from the story, but without knowing more about it, it is hard to be sure, but I don't think we should leap from one story to we need stronger regulatory oversight on this particular issue.

There is actually a fair amount of information out there for technologists and radiologists already for implants for cosmetic purposes, and the same, in my opinion, goes for the issue of the reconstructed breast.  There is a lot of information out there.

We teach people what Dan and Debbie have already said, which is that you should not be imaging these people routinely after a mastectomy.  It is too confusing and we find more false positives than negatives, so I feel like the regulatory guidance is sufficient, as written, myself.

DR. HARRISON:  Last comment.  As many of you know, in the situation with mastectomies, you can have upwards of 5 to 6 percent of breast tissue still there after mastectomy, so that imaging of those areas, if there are clinical concerns, as has been pointed out, is specific.

Very clearly, if something is palpated, if there is concern about those areas, those areas can be imaged, but routine imaging is not expected.

DR. IKEDA:  I agree and I think that if the patient has a problem that is a medical decision and it is a physician decision with the patient rather than I think something that should be seen by regulatory oversight.

I think there is plenty about imaging the breast with FDA regulations.

MS. HARVEY:  Any more comments about that?

DR. FINDER:  If we are finished with comments about Modification Document No. 5, there are some issues that have come up recently referable to use of assessment categories, and I just wanted to kind of bring up the issue before the committee because I understand that the BIRADS Committee, the people that came up with assessment categories many years ago and which we have basically pretty much adopted into our regulations, is reviewing the current use of the assessment categories.

I just want to bring up some issues that we have become aware of referable to this.  Obviously, the idea behind the use of assessment categories was to get some standardization into the medical report, so that anybody who is reading these reports would have a pretty good idea of what was going on.

I am sure that radiologists on the panel can tell you about the two- and three-, and four-page reports that used to be issued, and by the time you finished reading them, you didn't know where you were or what anybody was trying to say.

In order to alleviate that problem, the use of assessment categories was mandated.  Right now there are basically six assessment categories that can be used in the reports, and my understanding is, is that the BIRADS is looking at some of these because we have encountered situations where one of these assessment categories may not apply in some of the areas that we have been aware of.  For example, people have raised concerns about using the currently allowed assessment categories when you are evaluating a postoperative breast.

There really is no great assessment category that fits that type of a situation.  It is kind of hard to say that a patient that has just been diagnosed with breast cancer has a negative or a benign looking breast when you know that they have got a diagnosis of malignancy.

Other situations that have caused some concern are dealing with the male breast.  Not a lot of patients get these, not a lot of male patients end up getting mammography, but when they do, they are also covered under the regulations, and their reports also require an assessment category.

Other areas of concern that we have heard from the community deal with the use of the incomplete assessment category for situations where they are waiting for comparison films.  Again, no great way to define it in terms of the current assessment categories.

Other issues deal with Assessment Category 3 are probably benign.  Some people don't like the recommendation that follows in BIRADS where all probably benign lesions are supposed to have a recommendation of follow up rather than in some cases where they want to recommend biopsy.

Now, that is not a problem for the MQSA regulations because we don't tie our recommendations directly to the assessment categories, but that is an issue that we keep hearing about.

Another issue that has come up deals with the possibility of using multiple assessment categories.  In our requirement that if you are going to do that, let's say you have two or three lesions on the mammogram, you have to assign an overall assessment category, which is based on the worst situation.

The theory behind that or the thinking behind that was that we had situations where there were multiple lesions on the mammogram, the report discussed all of them.  Unfortunately, they discussed the benign ones first and buried the malignant one in the middle of everything, and they were missed.

So, the idea here was to require an overall assessment category that was the worst one, so that people would know not to miss the cancer.

Every solution creates its own problems, and there is at least a theoretical possibility under our system where you have to list one as the overall assessment category.  Let's say there is a suspicious lesion and there is one that requires additional workup and incomplete.

If you only have one assessment category, the overall assessment category of the suspicious one, is it possible that somebody will forget about the other one that needs further evaluation.

So, these are some of the issues that we have encountered.  We would like the committee's discussion on this a little bit, taking into account that the BIRADS group is looking at modifying their assessment categories to deal with some of these issues.

DR. IKEDA:  The first issue of the patient with the known cancer can be an issue for the current BIRADS codes in that the way that cancer is being treated now is much different than it was even five years ago.  For example, a patient who has a 5 cm breast cancer in the past would get a mastectomy, whereas, now many of these patients are getting neoadjuvant chemotherapy.

For those of you unfamiliar with this, this is the patient gets chemotherapy before the tumor is removed, and so the reason for imaging and for doing the chemotherapy is to watch the tumor shrink, so, for example, the patient may have a mammogram with a suspicious tumor, gets a BIRADS 5, highly suggestive of malignancy.

Six months later, after four rounds of chemotherapy or three rounds of chemotherapy, what the oncologist wants to know is has the tumor shrunk.

They know there is a cancer there, for example, and so we are in the unfortunate position of saying yes, there is a tumor there, and then they get another BIRADS 5 suspicious for malignancy, and everybody knows it is there.

I think that there was movement afoot of perhaps adding another code, BIRAD 6, known cancer present, to address this and to reflect the changing of how breast cancer is being treated, and so that these patients aren't being scared to death with a letter saying, hey, there is a cancer, they think they have a new cancer, and then they are frightened, and that's horrible.

I think that that would be something reasonable if that was approved by the BIRADS Committee.

MS. HARVEY:  Ms. Pura.

MS. PURA:  Linda Pura.

Dr. Finder, I concur with you on the Category Zero, not only is the client who is very confused and very scared, but our primary care clinicians also interpret this in very strange ways, so I would like to ask for a clarification on Category Zero on the BIRAD scale.

DR. PISANO:  I guess I can provide it.  Category Zero is intended to be, at least when a patient is in the process of being worked up.  It is not intended to be the final assessment category.

So, we, at UNC, only use it when we do screening mammography to say we read our screens in batches and we interpret them.  Basically, the patients either get a zero or a 1 or a 2.  That is just the way we have decided to do it, and I am sure this varies from some people use 4's and 5's in the screening setting, too, but we actually bring everybody back in for the complete workup, make sure they didn't have a biopsy there causing the spiculated mass, for example, talking to them one on one.

So, I think in most settings, zero means we still have to do more tests before we are sure whether something needs a biopsy or not, but perhaps it is being used differently than that in some centers, I don't know.

MS. HARVEY:  Dr. Kopans.

DR. KOPANS:  I am on the BIRADS Committee and there is discussion about Category Zero in terms of use if someone is waiting for old films.  I think it is going to come out that that is an inappropriate use of Category Zero. As Dr. Pisano just said, the categories were set up to you are evaluating the study now and what is your final assessment now.

If you need additional evaluation, that is a Category Zero.  If you are waiting for old films, you interpret the study as if you don't have the old films, and Category Zero therefore is not appropriate unless there is something that needs additional evaluation.

So, the people who are saying they don't know how to use it, they are not using it properly if they are using Category Zero that way.

MS. HARVEY:  Thank you.

Yes, Dr. Pisano.

DR. PISANO:  I just want to make another comment about it.  In our center, it is sometimes difficult because the two breasts have different assessment categories, and we are supposed to dictate one final assessment category, and we have been cited actually on inspection, because we put this is a 5 on the right breast and a 3 on the left breast.

In fact, that is kind of a good communication skill.  When you were talking before about burying information in the report, we would like the final word on each breast to be in the conclusion.  You know, you don't want them to forget about the 3 in the other breast, the one that has to be followed in six months just because you found a cancer in the other breast.

So, I would like to see some flexibility about that built into the advice to clinicians, and I am looking forward to the new BIRAD scale because we do have all these other problems you have mentioned.

DR. FINDER:  The issue that you bring up has been brought up before about the use of assessment categories for each breast.  That kind of shifts the problem down one level because if you still have two lesions in one breast, you still have the same issue.

The question really I guess is, is should these assessment categories be for each lesion rather than overall, and if so, it would require a change in the regulations obviously or at least our implementation of them, but as I said, we did have the situation where multiple lesions were identified this way, and unfortunately, the malignant one was buried in the middle of a whole bunch of other benign processes.

It is one of those things where no matter what you do, you create another problem somewhere else.

MS. HARVEY:  Dr. Kopans.

DR. KOPANS:  The issue again I think of multiple assessment categories, the report is still there, and every physician is supposed to read the entire report, not just the final assessment categories.  The final assessment category is supposed to be a summary, and I think the BIRADS Committee would come down on the fact that it is the most important, which is what Dr. Finder already said.

If you are just going to repeat the whole report in the assessment categories, it becomes a little redundant and silly, I think.

DR. FINDER:  Dr. Kopans, what is your feeling or do you know what the BIRADS group is thinking about in terms of assessment categories for each breast?

DR. KOPANS:  To be honest, I hadn't actually heard that as an issue.  It may be, they are reviewing a whole bunch of things at this point.  Again, BIRADS was designed to simplify the report, and the issue of what BIRADS does versus what the FDA requires, I suspect probably are two different things.

I don't think that the committee is going to change to an assessment for each breast, but I don't think that has been decided.

DR. PISANO:  When I made my comment, I was really referring to FDA regulations because the BIRADS Committee, I don't think anybody from ACR would object if, in the interest of clarity, you said in your conclusion, you know, there is a cancer in the right breast, but there is also something in the left breast that needs to be followed at six months, it is probably benign.

That just fits the way we do mammography, and I agree with Dan in reading the whole report, but sometimes you do need to call out the most important information.  I am concerned because we were cited because we did that at an FDA inspection, so I feel like that is good medicine to bring out the most salient points in the conclusion, and yet if it conflicts with--we actually had to change our practice, so we didn't put it in our conclusion, which seems kind of silly to me.

DR. FINDER:  Let me just talk about that a little bit.  We do have guidance actually that addresses the issue of the calling or labeling each breast with its own assessment category.  It is allowed under the regulations, but what we also say is if you are going to do that, then, you have to have one overall assessment category.

But you certainly are free under the regulations to do it that way, but if you are, then, you have to have one overall assessment category.  That is the way the regulations are pretty much written and interpreted.

One thing I do want to mention is one of the reasons we came up to that conclusion didn't have to do so much with the assessment categories as the fact that at least in the beginning, there were facilities that when they divided the reports up into right and left breast, they were counting those as two studies, and they were using those to meet continuing and initial requirements by dividing up the mammogram into two studies that way.

So, in order to stop that, we said that it is per patient, not by breast, and that is how we ended up where we were with the assessment category, so it all ties into that. There usually is a method to how we come up with these things.

DR. KOPANS:  Another way of working around this if there are multiple final assessments that are important, you can put just before the FDA statement, "Please note there are multiple important findings, the most significant is," and then the--

DR. FINDER:  Right.  There is a lot of flexibility in terms of what the report can contain and how you write it.  We don't mandate the makeup or the format of the report, but what we do say is there has to be an overall assessment category.  That is the one thing that we pretty much do require.

MS. HARVEY:  Thank you.

Any more comments?  Yes, Dr. Ikeda.

DR. IKEDA:  I am a little concerned about adding a second assessment for the second breast.  I am concerned about missing the most--I thought the intention of FDA was that the most suspicious finding not be missed and be acted upon, and so if you have, instead of 10,000 reports, you have 20,000 going out, I am concerned about that patient having her cancer missed.

So, as far as the regulatory recommendation, I know that you can put down the multiple assessments, but I think the overall assessment for the one patient is a good idea and to keep it simple, so she gets treated for her cancer.

MS. HARVEY:  Dr. Young.

DR. YOUNG:  Don Young.  I would emphasize that also.  I still have referring physicians call me and say what do you mean by "probably benign."  Adding another layer, assessment layer or two or three is going to confuse the patients and the referring physicians.

We cover these things in the body of the report, as has been brought out earlier, and you can phrase your report in such a way that it mandates see above, paragraph 2, if you get into two paragraphs.  I try to avoid two paragraphs.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  Actually, it may be my suggestion was being interpreted differently than what I was really suggesting.  I am not suggesting that we increase the number of requirements in the conclusion.

I am suggesting that we not be cited if we do mention the opposite breast in the conclusion besides the one with the most suspicious finding.  I mean that is what we were cited for, and I wish that the guidelines to the inspectors would be clearer, that it is allowed to use what you just said, Charlie, a few minutes ago, about the report has some flexibility.

I feel like we should have some purview over good clinical practice in this area.  There are times when you know the referring physician, you know how they read their reports.  You know that you need to emphasize both things, and I feel like that is a doctor's decision, we shouldn't be told not to do it if it's good clinical practice.

So, that is the point I was making, not that we should start doing it on every one, that is the last thing I want is more paperwork and more time spent by me.  I want the ability to do it if I have to.

MS. HARVEY:  Any other comments?  No?  Thank you.

I think we have come to a natural stopping point for lunch.  We will return at 1 o'clock.  Thank you.

[Whereupon, at 11:45 a.m., the proceedings were recessed, to be resumed at 1:00 p.m.]

AFTERNOON PROCEEDINGS

[1:00 p.m.]

MS. HARVEY:  Dr. Finder will speak first on our topic of Full Field Digital Mammography.

DR. FINDER:  I just wanted to expound a little bit about some of the things that were said earlier this morning about the way that full field digital units are currently handled and the need to have them hooked up to a film-screen unit.

Just to give you a little bit of background and history on how we ended up where we are, when the first full field digital unit was approved by FDA for commercial use in early 2000, there was no accreditation body that existed in order to accredit those units.

Now, we deal under the Mammography Quality Standards Act, and that Act requires that these units be or that the facility be accredited and certified.  The only way we could handle this under the situation was to tie a full field digital unit to an otherwise already accredited and certified film-screen unit.

That is the way that we have been handling it in the interim until an accreditation body is approved.  It is not the optimal situation, but in addition to keeping us compliant with the law, it also gives us some assurance that the people that are involved with the full field digital unit also have some experience with film-screen and are capable of becoming fully accredited with the film-screen unit, so that does give us a sense of security for these people to use the FFDM unit.

So, as has been stated in the morning, yes, we do require that currently, because there is no accreditation body out there, that if a facility wants to have a full field digital unit, it in some manner has to be tied to an accredited and certified film-screen unit.

In order to make this as viable as possible, we have expanded out our ability to tie that unit to another facility, and has been stated in the morning, it isn't necessary for the two units to be actually physically in the same place.  We have allowed some leeway there.

The alternative, if we didn't do this, was basically to say that since there is no accreditation body and no other method to deal with this, that you can't use it at all, and we didn't find that to be acceptable, so we came up with this method that kept us within the legal constraints, also gives assurance that the people that were using this were capable of otherwise accrediting or capable of performing quality that would get them accredited for their film-screen unit.

So, I just wanted to clarify that before we go on any further.

MS. HARVEY:  Thank you.

Our next presentation this afternoon is Dr. Etta Pisano.  She is going to talk to us about the ACRIN Trial of Full Field Digital Mammography.

Dr. Pisano.

ACRIN Trial of Full Field Digital Mammography

DR. PISANO:  Thank you.

[Slide.]

As I mentioned earlier, the National Cancer Institute has funded a trial, which is a screening trial for asymptomatic women, through the American College of Radiology Imaging Network, which some of you may or may not know is a new cooperative group.

The trial is slated to cost about $27.5 million. It is supposed to last three years total, and it is underway as of last October.

[Slide.]

A screening trial, as I mentioned, is one for asymptomatic women, women who do not have lumps or discharge, women who would normally present to the centers where the trial is open for screening mammography.

[Slide.]

Women undergo both digital and film mammography as part of the trial, and then the images, the screen-film mammogram, and the digital mammogram, all four views, of course, the same standard four views are taken for each exam are read independently, so that one reader reads the film exam and another reader reads the digital exam.

The same reader does not read all the film exams, and the same reader does not read all the digital exams.  The readers are split evenly between the two conditions, so that if you are reading on patient A, reader 1 will read the film and reader 2 will read the digital, and then for patient B, they read in the opposite conditions across all the images acquired at their site, so that it should be about evenly split across all the readers at the site.

There are a variable number of readers at the different sites.  Some sites only have two readers and some sites have as many as five to seven readers.  It depends on the number of people that are qualified to read digital mammograms.

If one exam is positive, the patient undergoes a workup for that exam.  If both exams are positive in the same area that the workup occurs, and if they are positive in different areas, workup occurs for all lesions found no matter whether both exams are positive whether only one exam is positive, and workup occurs as per usual clinical protocols meaning if normally at that site, additional views and ultrasound are obtained after a positive mammogram, that is what will happen after the positive film and/or digital mammogram.

A negative mammogram is treated the same way a negative mammogram is treated in any practice with routine imaging and clinical follow-up, meaning at a year for this study.

We do not have a lower bound for age, we don't have an upper or a lower bound for age.  We are allowing the sites to determine which patients they normally screen, so if a site starts screening at 35, the woman is eligible for screening at that site, then, she is eligible for the trial.

Some sites don't open their screening to women until they are 40.  Some women vary the first age of screening with risk, so some women are eligible to be screened starting at 25.  If they would normally be screened at that site, then, they are eligible for the trial.

[Slide.]

Our outcome measures are going to be--obviously, we don't have these yet--are going to be ROC, assessment of the two technologies with the primary outcome measure being area under the ROC curve, which reflects both sensitivity and specificity.

We are also going to measure positive and negative predictive values.  We are both going to perform an assessment of digital mammography as a modality and for each individual manufacturer, because we can do that.  We are going to have lots of patients in each condition.

[Slide.]

I mentioned this already.  It is consecutive women.  All women are approached, not all participate, of course.

[Slide.]

The total numbers are somewhat staggering - 49,500 women over a two-year period will be enrolled.

I mentioned the trial will last for three years. That is because we had to build in a year of follow-up.  At this point, there are 19 centers open to accrual.  Actually, this week we are opening our 20th center.  There are 29 centers that are in the process of becoming open, in other words, nine more still waiting to be opened.

We have 6 to 11 centers for each digital mammography manufacturer that has equipment available to be tested at this point.  Trex is a subdivision of Lorad, so I have Trex there, but you can substitute Lorad for that.

[Slide.]

These are the names of the sites.  These are for the Fischer equipment.

There are a few minor changes on my slides compared to what is in the handout.  That is because I sent it in a couple weeks ago and then I found some errors.  So, what is up on these slides is correct, so if you notice changes, that is why they are changed, is because I corrected my slides after I sent it in.

So, these are the Fischer sites, and you can see that we have a variety of types of sites.  We have academic practices.  This is a private practice here in D.C., Washington Radiology Associates, as this is LaGrange Hospital in Chicago.

[Slide.]

These are the Fuji sites.  Mainly, these are academic sites, these are all academic sites.

[Slide.]

The GE sites.  We have more GE sites than any other machine type, primarily because GE was the first machine type that was FDA approved, so we do have more machines in that category than in any other category.

Originally, when we first opened the trial, we envisioned balancing by machine type, but we quickly realized that we needed more patients than we thought we were going to need--well, the same number of patients, but we couldn't get as many patients per site as we were hoping to get, as quickly as possible, so we added sites, and so we ended up with this unbalanced per-machine type.

[Slide.]

These are the Lorad sites.  Here is the private group, Monmouth Hospital in New Jersey.  I think there were several in the GE sites that are private groups, as well, so we do have a range of types of settings.

[Slide.]

There will be 1,800-plus women enrolled at each center.  It is going to vary, of course, from center to center, because our goal is to accrue for two years, and some of the sites don't even come open until after a year.  We started out with the idea that we would have 19 or 20 sites and that there would be about 4 or 5 per machine type.

We quickly realized that the sites could not enroll at the rate we were hoping, so we then added another 10 sites, so that we are in the process of adding those extra sites right now.  So, the numbers in terms of per-machine or per-site is a little bit flexible.

Some sites, the ones that came up early, will have higher numbers, some sites, the ones that came up later, will have lower numbers per site, so we are going to hope to accrue the 49,500 women in a two-year period.

When we first opened, we thought we could do it in 18 months, so we have had two adjustments made just from the realities of accruing this many patients per day.  It was supposed to be 6 to 8 women per day getting two kinds of mammograms.  You can do the math and you can figure out how hard that would be for sites.

They are doing it, they are doing quite well, but it is a lot more challenging than we originally thought.  The interpretations are the least of it, getting consent, filling in forms, all the things that go with running a clinical trial, it is very, very labor intensive, so we have adjusted very rapidly to this information that sites are not able to accrue as fast as we had hoped they would.

So, this is the current plan, which is two years of accrual.  We started in October of 2001.  We are hoping to stop accrual in October of 2003, more than a year from now.  After that, we will have a year of follow-up, which you must have a year of follow-up for negative mammograms, at least a year, to declare the patient cancer free.

In fact, probably that could be debated.  You might need as much as three or four years to really know a patient didn't have cancer when she had her mammogram, but for purposes of this study, we are defining a woman's mammogram as a true negative after one year of follow-up if she does not have a diagnosis of cancer.  Actually, the exact range for follow-up is 10 to 15 months.

So, if a woman presents within 15 months of that mammogram, the entry mammogram with the cancer, we will consider the original mammogram a false negative.  I should say--I didn't make that quite clear.

Anytime up to 15 months, the 10-month period is when we must have, we are mandating a follow-up mammogram of some sort.  The earliest after the original mammogram that can count as being appropriate is at 10 months.

[Slide.]

Patients are excluded if they have a lump or they have a bloody or clear nipple discharge.  If they have implants, they are excluded.

Implants require extra views of the breast, and we didn't think it was appropriate to include them because they would require four views per breast times 2, so a lot of images, and we thought it was too much radiation and we wouldn't have enough numbers in that category to really get statistical validity or statistical power, so we excluded them rather than including them.

Anyone who is pregnant or thinks she may be pregnant or if a woman can't undergo follow-up mammography. We are not requiring that they come back to the same institutions for their follow-up mammograms because we wanted to open the trial to as many people as we could, but they at least have to provide the follow-up mammograms for review by study radiologists.

[Slide.]

As I mentioned, we are doing two standard views of each breast, and there is some tiling needed sometimes, women sometimes have larger breasts and the detector size for films and for digital, so if a woman is too large for the detector, for just one view in each projection, she might have to have inners and outers, uppers and lowers, just as we do with film mammography in standard practice right now, and the same technologist does both exams, and so she is supposed to do both exams the same way.

If she has done inners and outers on the film, she will do inners and outers on the digital.  We are not doing any extra diagnostic views during the screening part of the trial.  You can do diagnostic workup with the digital system if it is FDA approved, but some of the systems are not, at least one of the systems now, just the Fuji system, is not yet FDA approved, so they cannot do diagnostic workup using a non-FDA-approved system.

In addition, there are no extra views done during the screening study.  That is just not what a screening study is.  So, we are studying just the screening exam in this study.

[Slide.]

I mentioned that the technologist is the same.  She has to be eligible to take mammograms under MQSA.  She is supposed to be using the same degree of compression and the same angle with both systems.  That is definitely the case for the Fuji systems since it is an image plate system where the film is just replaced with the plate.

It is more challenging when you have to move the patient between rooms, but they are supposed to try to replicate the exams as much as possible.

In addition, we are randomizing the order of acquisition.  We were worried about the technologist being able to kind of tweak her technique a little, if she saw the film mammogram, she always did the film first and then did the digital, she would then know that it was a dense-breasted woman or something about the patient, the way she compressed or something, and make the other exam better.

So, that is why we are randomizing the order of acquisition.  That is assigned centrally.  When you register the patient for the study, you are assigned a randomization, so you either perform digital first or film first for all 49,500 women in the study, that will happen.

We are trying to standardize the dose between the two systems and if the AEC is available, try to use it in the same manner for both systems.

[Slide.]

As I mentioned, there are two radiologists per patient, one to interpret each exam, one for digital, one for screen-film.  They are not supposed to talk to each other.  They are not supposed to be aware of each other's interpretation.

We have been careful at the sites to take any little hints they might get about what the other person thought.  Every place has their own way of making sure the patients are called and so we have to inform them if they have a positive mammogram.  We have to pay a lot of attention to that.

We go there to make sure that the second radiologist can't figure out what the first thought.  We have no fellows or residents in the room during the interpretation of these exams.

They are certainly free to be taught on the images after the interpretations are completed, but not until then, not until the radiologist has entered his or her final interpretation, because we were afraid of increasing the sensitivity of the study if even a second person was reading, we don't want that to happen across the study, so it is just reader for the final reading for each set of images.

[Slide.]

As I mentioned, we are splitting the radiologists. I am a reader for this study, so I am reading equally in film and digital, as is every other reader.  We started out with a requirement that there would only be five readers per site, and we soon found, because of people retiring and quitting, and doing all the other things, that we couldn't stick with this requirement, so this has been dropped, there is no more than five readers allowed at each site.

We have an executive committee that meets weekly by conference call, and it consists of myself, Ed Hendrick, who is the co-PI of the study, and four radiologists, who are one for each machine type, who has experience with each machine type, plus the ACR personnel, plus Martin Yaffe, who runs our QC program for this.

We meet weekly and we discuss issues such as this, and five readers, it soon became apparent that some of the sites started with five readers and then lost people, and we didn't lost penalize them, so we now require that all readers read equally in each condition, and we try to have every reader read more than 50 mammograms, which even that is difficult to definitely require because of the fact that we have had people start reading and say I can't stand doing this, I am not doing it anymore, and so you can't really penalize the site, you can't throw those patients out because a reader only wanted to read 20 cases, so we are keeping those readings in the study.

[Slide.]

I mentioned already that we have no trainees, just MQSA-qualified or would be, some are Canadians, so would be qualified if they were living in this country, or practicing in this country, physicians.

We enter all data on-line, into a web site, where the data forms are kept, so it's on-line data entry for a lot of the data of the trial.

There also are worksheets that radiologists can fill out because you can imagine.  I am a busy person, I don't want to sit there at the computer all day entering data, so we have forms that a person can fill out, instead of going on line and entering directly themselves, and then it can be entered later by someone.

You can enter on line, but you have some flexibility about that.  As I mentioned, if either exam is abnormal, the workup should take place according to the standard clinical protocols, and even if only one exam is abnormal, even if you think the other exam kind of clears the first exam, you are supposed to work up the lesion, because, you know, you might be interpreting the digital mammogram differently or the film mammogram differently given the first study.  So, it is best to just go ahead as if it is a regular clinical workup on that patient, not refer to the other exam.

[Slide.]

We use usual equipment and practices at the different sites.  At most, especially when we first opened the trial, most of the cases were worked up with screen-film mammography, but as I mentioned, as they become FDA approved, we have allowed people to be worked up with digital mammography.  Often senographe is used.  Some sites are using MRI and other imaging tests.

We are keeping track of all the downstream costs of the positive exams, and of course, biopsy is one of the most important outcome measures of the trial.  That is how we find out who has cancer or not for the positive mammograms, the BIRADS 4 and 5 lesions.

[Slide.]

I already mentioned this.  You can consult widely with other experts if that is your normal clinical practice for the workup.  We don't control the workup at the sites, so that if you want to show it to three or four other radiologists before you make a decision to biopsy, if that is your normal clinical practice, that is not controlled in the protocol.  We want to know the truth about these patients, so we are allowing people to do whatever they normally do, and that might include consultation with many others.

[Slide.]

We use the standard BIRAD scale, but that is not the scale we are using for the primary outcome measure in the trial.  We are collecting that for every patient for both digital and film.

We are also collecting the probability of malignancy on a 7-point scale--I think I have it up in the next few slides--and a call-back scale, which is the probability that you think that the patient needs to be called back with, how worried are you basically.

[Slide.]

Here is the probability of malignancy scale.  It is just words, as you can see.  We also ask for the actual percent probability of malignancy as a separate item, which I don't know what good that is going to do us for sure, but we definitely have collected that.  There was a lot of discussion of whether we just use the 7-point scale or the 100-point scale.  We are collecting both.

This is the scale that you must choose on every single patient, you must give one of these 7 points.  You give it for individual findings, and then you do it for the overall assessment, similar to what FDA requires, so you are supposed to classify each individual finding and then classify the whole mammogram.

So, this is the scale for that, definitely not malignant, all the way to definitely malignant.  You can see there are gradations in between.

[Slide.]

The call-back scale is a similar idea.  We are trying to get an idea of the level of uncertainty that radiologists have in deciding whether the patient should come back or not.

So, you might call something BIRADS 1--I am just making this up, I am not thinking of a particular case--you might decide to call something BIRADS 1 or maybe 2, but you think to yourself, well, maybe there is some evidence the patient should be called back for diagnostic workup, but I am not going to call them back, I think it is below my threshold, so you might give that a 2.

In other words, you are noting there is something there, but something that is below your threshold.  The only things you have to give more information on is the things you actually do call back.

So, you can see all the way down from no evidence to overwhelming evidence the patient should be called back.

[Slide.]

We have, of course, follow-up, because that will help us determine truth about the vast majority of patients who have negative mammograms.  Patients are supposed to keep in touch with us.

We have a system where we are sending out newsletters and keeping in touch with the patients.  The local sites have the contact information for the patients, and they can contact the patients by phone and mail.

Obviously, we would like them to all come in for follow-up mammography.  If they don't come in, we have a way of calling them and finding out whether they have been diagnosed with breast cancer or not.

In addition, we plan, for the patients who are lost to follow-up, make sure we check their medical record at the hospital where they are imaged and search tumor registries, and things like that.  I am hoping we don't have huge numbers of women lost to follow-up because it is included in the consent form that we expect them to come back in a year and we are doing our best to keep up with them, but some women will be lost to follow-up, of course.

We are going to code follow-up information on the mammograms.  We need to know whether there are new findings on the follow-up mammogram.  Basically, we are going to collect the BIRAD status on the follow-up mammograms primarily because that will help us determine the truth about that first year's mammogram.

Obviously, if they have had any biopsies in the meantime, we are going to collect information on that, as well.  We have central over-read of all pathology.

It says 9 to 15.  We have had a huge debate about this.  What I said earlier was 10 to 15 months, and I believe, I mean I wish I could remember really straightforwardly right now, but I think we actually changed it to 11 to 15 recently, so I would have to check on that, but the reason is because you want to be--obviously, if a patient has breast cancer anytime up to 15 months, we are going to call the first one a false negative.

The question is when do you allow yourself to call it a true negative, how soon follow-up can you have to call it a true negative, so there has been a lot of discussion about that, and my recollection now is that it is 11 to 15 months, so this is slide is not correct.

[Slide.]

I mentioned this earlier, that there is the team of physicists headed by Martin Yaffe at the University of Toronto, and then helped by Ed Hendrick who is the co-PI on the study.  He is at Northwestern University.  They are overseeing quality control at all these centers.

They are doing the MQSA requirements for the screen-film units, and they are also sending all the digital data and the screen-film data up to Toronto where it is looked at centrally, as well.

In addition, we had a very comprehensive acceptance testing protocol.  Some of these machines are not FDA approved yet.  So, we are making sure that they are functioning at a minimum standard, what we have determined in advance prior to the entry of patients into the study.

We have had no real glaring problems with quality, no reasons to shut sites down or anything like that during the--let's see, 10 months we have been open at this point.

[Slide.]

We are also doing a cost effectiveness analysis, which is, as I mentioned, we are keeping track of all the downstream tests, and if someone has both a positive digital and a positive film, then, you can't really pull out which test caused which, you know, which downstream cost, but you will know that both would have caused those downstream costs.

One of the hypotheses is that digital actually has fewer false positives.  That has been shown in the other large screening trial that has been performed by John Lewin and Ed Hendrick and Carl Dorsey, which has been published already.  They did find a statistically significant lower call-back rate and lower false positive rate, lower false positive and lower biopsy rate.

So, that has been shown before with the GE system in their big study, so we are hypothesizing that will be shown again in our study, so we are trying to figure out what that means in terms of medical costs.

We are asking the radiologists to record what additional tests are done on these patients, and Dr. Anna Tosteson, who is a health care economist at Dartmouth Medical School, will be doing modeling to look at the cost effectiveness issue.

[Slide.]

In addition, we have a patient telephone survey that will be done on a subset of patients in the trial.  That was developed by Dr. Danny Fryback of the University of Wisconsin, which we expect fewer false positives, so we are trying to find out how important that is to patients by interviewing patients who had false positive mammograms versus the control group.

Those are being done centrally, but through telephone surveys, so it is a very short, few-item questionnaire, and it is being done out of UNC.

[Slide.]

All pathology reports will be reviewed by one of two breast pathology experts.  In addition, the specimens of any breast tissue that is obtained on patients during the course of the follow-up period will be reviewed by the breast pathologist, these same two breast pathologists, one of the two of them.

In addition, if there is a disagreement between a local pathologist and a central pathologist, then, the other central pathologists will be called in to give a tiebreaker vote, and hopefully, there won't be a three-way disagreement, but we may even have one of those, so we will have to find a fourth person to kind of help us.

I don't have any data about how often this is happening.  I did just see the data from R.5 on this, which was a multicenter clinical trial of stereotactic and ultrasound guided core biopsy for nonpalpable lesions, and the rate of disagreement in that study was only 4 percent, which is very good actually for a kappa number.  It is 96 percent, 0.96 kappa is extremely high for any technology.

So, I am hoping we have numbers like that for this study.

[Slide.]

What I have described just now was what we are doing to achieve our primary aims and I also described some of the secondary aims because I talked about quality of life and cost effectiveness, but what I just described are the main aims of the study or how we are achieving the main aims of the study.

We are also planning additional studies during year 3, after all the patients have been acquired, all the images have been acquired, we are going to run reader studies to assess other aspects of digital mammographic performance, diagnostic performance.

For example, a big issue is softcopy display versus printed film display, does the technology perform equivalently in both conditions.  Just as an aside, I didn't say this, but we have assigned reader methods for each system.  For example, GE, all readers are reading in softcopy.  For Fuji, all readers are reading in hardcopy.

For Fischer, all readers are reading in both conditions, both soft and hard copy, and for Lorad, right now they are reading in hardcopy, but they may be switching to softcopy.  The point is that we have standardized it across sites, so we really can't answer this question, softcopy versus printed film, within the study.

Really, it wouldn't be appropriate because each system has its own softcopy system to really make a statement across digital mammography for soft versus hard copy, so what we are going to do is run reader studies in the third year where we compare digital softcopy to digital hardcopy for each machine type, so that each new manufacturer will give us their best version of their softcopy system and will run reader studies comparing soft to hardcopy for each manufacturer to determine the diagnostic accuracy.

In addition,  I am going to skip over the second bullet for a second because it is kind of an arcane statistical thing that I think the statisticians will be happy we did them, it will take me a few minutes, so we are going to look at breast density and its effect on diagnostic accuracy and compare digital film for that, because the hypothesis is that obviously, people believe digital may be better than film in dense-breasted women, and so we are going to have a lot of data on that by the end of this trial.

You know, this trial is costing $28 million and doesn't even include all the systems that exist in the world for digital mammography.

I am personally aware of two other systems that didn't make it in time for this trial, there may be others out there, as well, and we can't keep running $28 million studies every time we have a new system, so the statisticians and those of us who do clinical trials are very interested, isn't there some other way we could do this without having to run a full-fledged clinical trial, so we are actually going to look at the use of a reader study in predicting, and I know the FDA statisticians at least over in the Device Approval Branch are going to be interested in the result of this study, is there any way to predict the results of the full-fledged trial with carefully controlled reader studies, can we estimate the diagnostic accuracy.

We are going to do this study with that percent prevalence in the reader sets to see if any particular mix, case mix is most predictive of the outcome.

I talked about this already as the secondary aim. Each unit may have different performance versus film-screen mammography.  It is possible that one unit really performs extremely well and better than film.  It is possible that another unit performs less well than film, so we really need to divide them up.

We have the whole technology kind of lumped together in the overall aim, but we will divide them up and look at them that way, as well.

Obviously, we have much less power for this aim than we do for the overall aim of comparing digital to film, but we do what we can and obviously, we will have a lot of power for the GE because they have a lot of systems out there being tested in the trial.

We will have less power for the other manufacturers because we only have six machines.  In addition, all these other characteristics will be tested, digital versus film.  Obviously, we have less power against any of these than we do against the main objectives - age, lesion type, pathologic diagnosis, et cetera.

[Slide.]

We also have a lot of technical aims in this trial.  We have a big quality control program going on, collecting a huge amount of data on a daily, weekly, monthly basis.  We are hopefully going to have a pretty good handle at the end of the trial on the effect of spatial and contrast resolution on diagnostic accuracy, if there is such an effect.

We have a range of image detectors in the trial on the Lorad system that was available until recently, which is getting discontinued, which is the one that is FDA approved at this point, has a 41-micron pixel size, and the GE system has 100-micron pixel size.

In addition, we have a range of contrast resolutions, 10 to 14 bits, and so we will have some idea, we hope, about the importance of those factors.  We are going to have a good idea of whether image quality varies tremendously across sites and we hopefully will know about dose, as well, and we are going to know about the range over time, whether things change a lot, drift a lot, so far there haven't been major surprises, but we will have lots and lots of information about that.

[Slide.]

I just wanted to show you a few pictures.  It may be too bright in here to really appreciate this.  This is an image from the University of Pennsylvania, and there is really two lesions here.  There is a big, obvious spiculated mass, and then right next to it is a little smaller, indistinct, irregular mass, which both were cancers.  I assume one was invasive ductal and the other was invasive lobular, but I am not sure about that.

[Slide.]

Here is another example of an infiltrating ductal carcinoma with DCIS, which was on the Fischer unit.  The first image I showed you was from a GE system.  This one is at the University of Toronto.

So, I am happy to answer any questions.  We have a lot of times still, I only took 40 minutes.  Does anybody have any questions about the trial?

MS. HARVEY:  Dr. Harrison.

DR. HARRISON:  You mentioned in getting pathology read that you were hoping that there was not a necessity to break a three-way tie, you either have malignancy or you don't.  How can there be a three-way tie?

DR. PISANO:  Well, we have a 5-point scale from pathology.  There is malignancy, there are obviously malignant, invasive ductal cancers, there are obviously benign lesions like fibroadenomas and things like that, but there are kind of gray lesions like--

DR. HARRISON:  It was rhetorical.

DR. PISANO:  Oh, I know you know the answer.  You want me to clarify.

DR. HARRISON:  Right.

DR. PISANO:  LCIS and the lobular neoplasias, lobular carcinoma in situ, the other lobular neoplasia lesions.  There can be quite a bit of disagreement between pathologists around those kind of high-risk lesions.

That is where we found disagreement in the R.5 data is right at that point.  We did very well overall, but that group had a higher rate of disagreement, and then DCIS, of course, sometimes there is disagreement in classifying a lesion as ductal carcinoma in situ, so we do have a range of possibilities.

Hopefully, the pathologists will agree most of the time, but we don't want it to be too high a rate.  I don't expect a very high rate.  Frankly, the R.5 data was better than we expected.

It was earlier studies that indicated as much as 7 to 10 percent disagreement rate between pathologists with open biopsy.  So, we were pleasantly surprised to see it was smaller than that for R.5, and presumably that suggests a training effect over time.  Maybe people are learning, they are not so unfamiliar with these border cases.

MS. HARVEY:  Dr. Lee.

DR. LEE:  I was wondering if you had a sense of out of the women who are eligible for the study, what percentage actually are participating.

DR. PISANO:  I don't really know for sure.  I can tell you about my site.  We actually have two machines going in two different buildings, and we have more women, we have lots of women who want to participate.  It is mainly the rate limiter is not the number of women who want to participate, it is more just physically, how many mammograms can you do a day.

You know, you are using these machines for clinical work, and you have to kind of fit in the research into the clinical day, and there is only so much you can do. We are taking all comers who want to do it, but you can't keep working until midnight.

So, that is really not the issue that patients aren't willing to do it.  We have lots of patients willing to do it.  That has not been our problem at all.  It is more just physically being able to handle that much per day.  It is very, very labor intensive.

MS. HARVEY:  Dr. Ramos.

DR. RAMOS-HERNANDEZ:  Because of the disparities among different ethnicities and the new tendency for African-Americans and Latinos to be diagnosed at younger ages, in the general information that you are getting, some sense, or are you collecting specifically age and ethnicity, so when we get results, we have a sense, see if this is more accurate for specific groups?

DR. PISANO:  Yes, we are collecting age and ethnicity, and, in fact, we have been monitoring very carefully our African-American and Hispanic and Asian enrollment in the trial.

We noticed in the first six months of the trial that we had fewer than we would like both in African-American and Asian and Hispanic women.  We had appropriate percentages of white women and Asian women.

So, we actually have targeted interventions directed toward the minority populations that we have put in place at specific sites.  What we did was we looked at the demographics at the sites that had more African-American and Hispanic women, and we have asked those sites to target those populations in specific ways.

We have noticed already that we are getting reports from the Data Safety and Monitoring Board and the Statistical Center, we get monthly reports, and the ethnicity report we got in April, we have already seen a correction, so we have gotten the numbers back up to where we want them to be or I should say up to where we want them to be already from where they were, which was not acceptable or below what we thought it should be in April just by getting a couple sites to make outreach into the community.

Actually, Washington Radiology is one of the sites that we targeted, and Denver, the University of Colorado, both those sites.  Washington Radiology and Denver both had large numbers of Hispanic patients, and so we improved the enrollment for Hispanic patients by basically having translation of consent and available person at the site who can speak Spanish, things like that, simple things like that.

In addition, the NCI has a communications kind of guru, someone who could help clinical trials recruit patients and target certain populations, and they have been quite helpful in coming up with strategies.  So, we have dealt with that issue.  I don't know whether there will be a difference, but we will look at it, the ethnicity of the patients as one of the factors affecting the outcome of the trial, we will look for it.

DR. RAMOS-HERNANDEZ:  Since the beginning of the trial, did you set up specific percentages of special populations that you wanted to achieve?

DR. PISANO:  What we did was we looked at the U.S. population and we are trying to reproduce the percent of the population studied that reflects what Hispanic, African-American, Asian, white, are in the population.

We probably are not going to succeed in doing that exactly, but we are going to come as close as possible to it as we can with these interventions that we have created or are undertaking.

MS. HARVEY:  Dr. Harrison.

DR. HARRISON:  Along the same lines, historically, African-American participation in clinical trials has been very low, down around 1 percent.  What are the numbers that you are seeing right now?

DR. PISANO:  Well, we are seeing much better than that, but I don't remember the exact figure off the top of my head, but my memory is it is closer to 10 or 15 percent African-American, maybe even higher than that.  I really don't remember the exact numbers, but we are coming close to the percent in the population at the different sites.

For example, Alaska, Native American, we don't have any sites that are really in places where there are Native Americans, so we can't possibly achieve the percent that really exists in the population of that particular subset of the population.  We are doing what we can.

I will say at my site, we are definitely getting a percent of women who exist in our population.  At UNC, it is about 30 percent African-American, and we are getting about 30 percent African-Americans at our site, but I don't remember the exact numbers at the rest of the sites overall. I would have brought the figures, I mean I can provide them if you are interested, but I don't have them with me today.

DR. HARRISON:  Are you allowing people with previous cancer diagnoses to be screened?

DR. PISANO:  The answer is if a woman has had a mastectomy, and she would normally get screening mammography for the opposite breast, then, she is eligible to participate in the trial.  Some sites do not image women as screening mammograms who have ever had a cancer diagnosis, so those sites, those women are never screened--

DR. HARRISON:  Right, they are always defined as diagnostic at that point.

DR. PISANO:  They are always diagnostic.  Now, women who have had lumpectomy and radiation or without radiation are not eligible because those women, we define them as needing diagnostic mammograms as opposed to screening mammograms, but women who have had mastectomy can participate if that center performs those patients as screening mammograms.

DR. HARRISON:  Right, because presumably, the contralateral breast does not have architectural disturbances that we have created.

DR. PISANO:  Exactly.

MS. GILBERT:  I know that you did address this earlier.  What are some of the strategies that you are using to recruit these women?  49,000 women are a lot of women. Can you just give me some idea of what you are doing?

DR. PISANO:  Well, the centers that are, first, the centers that are participating in this trial are all high-volume centers.  We don't have anybody who does five mammograms a day or anything like that, so we are not talking about centers that have no population to recruit from.

The primary thing we do is publicize the trial at the centers where there already are screening mammograms happening in large numbers per day.  At my site, it is over 70 screening mammograms per day.  So recruiting 15 of those women, you know, 10 to 15, some number like that every day has been no problem.

On top of that, there has been local publicity at the local sites.  I mean it is kind of a tradeoff.  You don't want to distort the population at the site away from the screening population.  You know, it is a little bit of a mixed blessing getting other women in who don't normally come to that site for screening.

Certainly, you want people to know about the trial, but perhaps people will come to get screened who wouldn't ordinarily or should not be screened.  So, we want to make sure that women who will come to the center for screening, we want to be able to make sure they come back in a year for follow-up, so we publicized it locally.

Each center has brochures they can distribute to local primary care physicians or to women's groups locally. We send it out.  At UNC, for example, they sent it to local PTAs, we have sent these brochures out, we have done all sorts of things like that.

I am not sure how many patients for sure come in because of that, but we do have some publicity.  There is a web site for the trial, which is a very attractive, nice place to come visit, and it gives very explicit instructions to patients on how to get a mammogram appointment at the site if they are interested in participating, or to women, I should say, these are screening women, they are not really patients.

If women want to come in to the centers for screening, they can come in based on the instructions at that site, but occasionally, patients, you know, we have to be vigilant that patients who aren't eligible for screening might get into the trial.  We don't want that to happen with this trial, as it has happened in other clinical trials.

We want patients to really be asymptomatic patients or women--I am getting my terminology mixed up--but we want women to come in who have no problems, and if they are symptomatic, they shouldn't be part of the trial.

We have done a lot of stuff, but, as I said, you have to be careful that you don't almost overpublicize it to the point where people who ordinarily aren't getting screened are coming in, maybe 20-year-olds.  I did have a 20-year-old come up to me after I gave a talk on it and asked me how she could be part of it, and I had to tell her, sorry, you are too young to be screened.

MS. GILBERT:  Is there an age range that you are looking at?

DR. PISANO:  Well, as I said, the centers screen a population of women, and there are some women who are eligible for screening below the age of 40, and so those women, if they come normally to that center for screening mammography, for example, if you are at high risk for breast cancer, your mother had breast cancer at age 40, at UNC, we start screening those women annually at age 30.

So, those women come to UNC for routine appointments now, they show up on our schedule for screening mammography now, they are recruited to the trial, we don't exclude them.

Obviously, we don't want to have too many women that are at low risk, and some women are worried about their breasts and start screening on their own at a young age, and we are kind of leaving that up to the local sites to make the decision would they normally screen these patients.

If they would normally screen these patients, they are eligible for the trial.  We don't really have a lower age cutoff.

There is a long discussion about that.  We talked about cutting it off at 40, but there are really high-risk women who wanted to participate in the trial, so we decided we would err on the side of including them as opposed to cutting them off because of their age.

MS. HARVEY:  Ms. Rigsby.

MS. RIGSBY:  Do you charge the ladies that come in for the screening?

DR. PISANO:  We do charge them for film mammography.  They get the digital mammogram for free, so it is really their insurance, most of them have insurance, some don't.  The ones who don't, they have to pay for their film mammogram out of their pocket or whatever way they normally pay for it.  We don't have, believe it or not, $27.5 million, enough to pay for the full mammogram, too.  It's a huge undertaking with this many thousands of women.

MS. HARVEY:  This is a phenomenal amount of data that you are collecting.  When will we expect to see the first reports of your findings?

DR. PISANO:  If we wait until all the follow-up is in, that would be October--it will probably be December, you know, given a few extra months there, the 15 months as opposed to the 12 months, it will be December of 2003.  That is until all the data is in from follow-up.

Obviously, we will have some preliminary idea before then, at least the statisticians will.  My guess is sometime in 2003, there will be something, but I don't know exactly what month that would be.

MS. HARVEY:  Thank you.  Any other questions?  Dr. Karellas.

DR. KARELLAS:  I have two questions.  If a patient asks for her original films because she is going somewhere, moving somewhere, how do you handle that?

DR. PISANO:  In the consent form for this study, we ask their permission to keep the films centrally, so the women relinquish their films to the study.  That is so that we can do these reader studies during the third year of the trial.

If they need the films back for clinical purposes, which is quite possible, they can be requested, and we ask them to send them back after they are finished with them. So, the patients aren't keeping their own films right now, they are all being stored.

The digital images, a copy of them is being stored in ACR Imaging Network Headquarters in Philadelphia, and the original film mammograms are being stored in ACR Imaging Network Headquarters in Philadelphia.

I had recently a patient who came in with a problem about three months after her film mammogram, and we got it by FedEx, her old films, within 24 hours.  That happened to me at UNC, so I know it is working.  It has worked for that one case anyway.

That is an experience we have had in the past, that we can get the cases pretty quickly back if need be.

DR. KARELLAS:  Have you thought of how you and all the sites would handle if a percentage of these patients want a copy of their digital mammograms at some point?

The reason I am asking you that is that I am finding it very difficult to deliver digital mammograms to people who want them, and it is okay when they want one or two or three a year, but it is very difficult to envision if we have 100 patients in a week ask for their digital data, it is very time-consuming.

DR. PISANO:  I don't know if I have any information from within the trial, but I will talk from my own personal experience as someone who has digital mammography in my practice.  I have not found this to be a big problem where patients are asking for their images a lot.

As I mentioned, we have two digital units doing clinical work plus we have another one participating in the trial, so I have not had a real problem with patients requesting their images.

What we have done when they have requested, we have a printer and we just print out another copy and hand it to them.  If everybody were asking for it, it would be exorbitant, but we haven't had that many patients, just as most patients don't want to keep their film mammograms, we haven't had that many people request that, so it hasn't been a problem in our practice.

DR. KARELLAS:  The reason I am asking about that, it is very easy if they request it at the time of examination.  It is always more difficult of they come back a year later, because the images are not just all neatly packed in one place on patients, so there can be various CDs or various locations, so that is something for the sites to think about as to how this could be handled in the future.

DR. PISANO:  That is a good point.  I haven't had anybody ask me for the mammogram except at the visit where we did the mammogram, but you are right, if it is even a week or two later, and it has already been archived, it is a little bit of a challenge to get that case back and printed, but it wouldn't be impossible.

Of course, you could always hand them the one you have in the jacket, knowing you could print another one in a few days or something while you search for it.

We don't have a good pack solution at UNC, I wish I did, but I don't.

MS. HARVEY:  Thank you.  Excellent talk.

We have a few minutes while we prepare for Dr. Kopans' computer to get set up, which might be a good opportunity for us to discuss any questions that any of the board members might have up to this point or any other issues that you would like to talk about.  There is lots to think about with the digital.  It is going to be the wave of the future.

Dr. Harrison.

DR. HARRISON:  When, in fact, a suspected malignancy is found in these clinical trials, the clinical handling of that is done by the local clinicians?

DR. PISANO:  I should have mentioned that in the talk probably.  Each site handles the identification of abnormal mammograms and the biopsies and everything else as per their usual clinical protocol.

For example, at my site, we wait until both mammograms are interpreted, and we have a nurse who contacts the patients and tells them about the positive exams, so that person makes the phone call.

Instead of just saying you had an abnormal screening mammogram, you need to come back for additional views, she says you had an abnormal digital or you had an abnormal film or you had both, an abnormal digital and film, and you need additional views.

Then, the letters they get include both, I mean everything has been adapted for the trial.  The primary care physician gets a dictation that says, in our case, I am talking about UNC--each site had to do what their IRBs would let them do or wanted them to do--so, at UNC, first, they get a report, the primary care physicians get a report that says film mammogram report with a signature, digital mammogram report with a report and a signature, and somewhere in the report, put on by a secretary at the top, is this patient participated in digital, she received both a digital and film mammogram, so that the primary care physician gets both reports, knows what is going on.

When we get the patient back for workup, the radiologist dictates into the final report, this patient had both, was part of this study, had an abnormal whatever, and we did the workup of this, this way, we did the workup of that, that way, and the final interpretation category.

So, we have adapted everything to this trial.

DR. HARRISON:  One corollary question.  Do you certify a letter to the patient?

DR. PISANO:  No, we don't send our letters certified mail right now, we just send them regular mail.

DR. RAMOS-HERNANDEZ:  Is there anything in the study or what would you do with people that actually couldn't afford to pay the screening mammography, but they are diagnosed, they get something, an abnormality, and they are going to get help to pay?

DR. PISANO:  I am sorry, I am not sure I understand your question.

DR. RAMOS-HERNANDEZ:  Is there any support for payment when finding an abnormal mammogram?

DR. PISANO:  So, if the abnormal mammogram occurs, is there any way to pay for the downstream costs?

DR. RAMOS-HERNANDEZ:  Yes.

DR. PISANO:  There is no money allocated at present in the trial to pay for downstream test costs.  That is part of the consent process, that the patient would be responsible for those downstream test costs.

That is the way we have dealt with it, is just inform the patient upfront that she may have additional costs if she is not insured.  We have had no insurance company refuse to pay for the downstream costs that I am aware of.

MS. HARVEY:  Thank you.

Dr. Kopans will talk to us about tomosynthesis.

DR. KOPANS:  And a few other things.  I come from Boston.  I have to admit I am a little anxious to not be standing with my back to the wall, which is the way we are taught, so if something is coming from behind, I wish you would let me know and have me duck.

Potential New Applications of FFDM

[Slide.]

DR. KOPANS:  I suspect that the committee is well versed in digital mammography, but before I get into tomosynthesis, which is really the main thrust of my comments this afternoon, I just wanted to briefly kind of review just so that everybody understands digital mammography and why there are multiple advantages to digital that are going to be coming along.

I am going to actually speak about two other aspects of digital other than tomosynthesis, but I will do that briefly.

[Slide.]

Basically, a digital mammogram, the way I think of it is a spreadsheet where you have got rows and columns, and each row can be labeled with a number and each column can be labeled with a number, and then the x-rays that are collected at each cell, if you will, or pixel, can be labeled with a number.

So, for example, 60 photons--I am just making these up--but 60 x-ray photons come through and are collected in this cell, then, we know that this is cell 5-5, and there are 60 photons, so we can define the location and the number of x-rays at that point in space by numbers, and hence, the term digital, so that the whole image can be described as rows, columns, and numbers of x-ray photons.

Then, we can manipulate that image because we can assign different levels of gray or color if we really want to, to help whoever is looking at the image appreciate things that may not be obvious with one sort of presentation.

So, a digital mammogram is just an image made up of numbers, and, of course, that allows the computer also to be brought to bear.

[Slide.]

I suspect the committee knows the basic advantages of digital mammography - higher contrast resolution, greater dynamic range than the film-screen systems, because you are not exposing the image to meet the parameters of the detector, for example, the film-screen, you have to have specific exposure values, otherwise, when you process the image, it won't be a satisfactory image.

With digital, you are actually uncoupling the detector from the display, and you can adjust the image afterward in any one of an infinite number of display characteristics, but that way you can image the breast, so that you get the imaging characteristics, and you don't have to worry about the display.

You can manipulate the image.  We use digital for real-time procedure guidance, I think you probably know that, with stereotactic devices.  The areas where I think digital is really going to have an impact over and above conventional film-screen mammography are in the areas that I have listed below.

I am not going to talk about telemammography, but I heard you discussing issues of some underserved women, and certainly being able to transmit images from areas where there may be remote access, digital mammography will permit that.

We have actually sent images between New York and Boston, the short way, 22,000 miles out to a satellite and 22,000 miles back again.  We did that over 70,000 times without a single loss of information.  So, that is doable. You could be in some remote area in, say, the western part of the country and transmit images anywhere in almost real-time.

[Slide.]

We use digital now, of course, for procedure guidance.  I am not going to spend any time on this.

[Slide.]

There are multiple ways of doing digital imaging. This is the image of a prototype slot-scanning device that we actually had back in the 1980s.

[Slide.]

It was way before its time, and unfortunately, the company didn't have the staying power, but the Fischer system works in a similar way with x-rays coming through a pre-breast slot and then a post-breast collimator, and then detected as a detector.

This is the system we had in the 1980s, but it is similar to the Fischer where the detector moves, it is a linear detector and it moves across as the exposure is made.  This has the advantage of being very good at scatter rejection.

[Slide.]

This is just for historical purposes.  I think probably the first very high-resolution digital study, again done in the 1980s with the American Science and Engineering system, this is 9 line pairs/millimeter of a high-grade ductal carcinoma in situ.

[Slide.]

Just one of the quick things you can do with digital is if you don't like white dots on a black background, you can just push a button and it will be black dots on a white background.

Although this is more than a parlor trick, sometimes just changing the image presentation can make things more visible to the human eye than other manipulations, so that sometimes just breaking up our pattern recognition allows us to see things that maybe we wouldn't have appreciated in the conventional way.

With digital, this is just a push of a button.

[Slide.]

There are other technologies.  This is the Fuji approach, which is a stored phosphor system where a plate is stimulated by the radiation to a higher energy level, and then that is read with a laser.  I am not going to go into the details of the different technologies.

[Slide.]

This is just a mastectomy study we did a number of years ago with the Fuji system.

[Slide.]

There are charge coupled devices, which basically convert the x-ray photons into light photons, and then the light is channeled down to a charge coupled device, which converts the light into an electrical signal, and that then goes to the computer.

[Slide.]

This is actually the system that the Fischer unit uses in somewhat of a linear array.  The initial Trex, Lorad, Bennett, I don't know how many other companies, had this system, and then as I think everyone knows now, Lorad has converted to a selenium system, which I will mention again in a second.

[Slide.]

These generally have to be put together in some form of mosaic, because the charge coupled devices aren't big enough to cover the entire breast.

[Slide.]

This was the old Lorad system with 12.  Each one of these is a CCD camera with fiberoptics in the screen that converts the x-ray photons.

[Slide.]

This is just from Lorad, a slide from Lorad using selenium to convert the x-ray photon directly to electrical signal without going through a light conversion approach.

[Slide.]

Again from Lorad.

[Slide.]

This is from General Electric, which is their system converts the x-ray photon to light and then, in an electronic array behind the light conversion level, they convert the light into electrical signal, and this is actually the reverse of what we do with most of our computer displays where there is literally a wire to every point in the area.

The difference between this and the selenium detector is that the x-ray photon is converted directly to electrical signal in the selenium detector, and in this, there is a light conversion factor first.

[Slide.]

Again, you can imagine somebody having to stretch these little, tiny wires to each one of these points, so the manufacturing of these systems is tricky, but I gather the companies are doing pretty well with it.

[Slide.]

You have probably seen digital images.  Notice how different the ACR phantom looks on a digital image than an x-ray film-screen image.  The answer is it doesn't.

[Slide.]

This is just so you can see the ACR phantom really is obsolete for digital images.  You can see just about everything in the ACR phantom with a digital system, and I gather there are new phantoms.  I know there are phantoms under construction, I don't know if any of them have been accepted yet, but I guess this committee is going to have to face those kind of issues as to how to test the digital systems.

[Slide.]

Again, you can see the major difference between film-screen and digital is no difference in terms of if you want to print the digital image or present the digital image the exact same way as the film-screen image, you can do that.

[Slide.]

The advantage, of course, with the digital image is that you can manipulate the image, and you are not just stuck with what you have got in terms of with a film-screen image, what you get is what you get.  You get one shot.  If the film isn't exposed properly, you have to do it again.

With the digital image, you can manipulate the image, and another advantage for those of us who are always looking for magnifying glasses in our divisions, someone in this country is collecting magnifying glasses because I bet if you asked every radiologist in the audience, they are all disappearing from our departments.  I don't know where they go.

But the joke is going to be on that person because we won't need magnifying glasses.  With digital, you just make the image larger, and that is a major advantage just by itself.

[Slide.]

Now, digital mammography, and D-MIST is going to I think confirm what most of us who have looked at digital and obviously are experienced with film-screen mammography know, and that is that digital mammograms are as good as film-screen mammograms.

Are they better?  Just by themselves, I don't know.  I think there are advantages, and D-MIST may be able to tease that out, we will have to see, but what really makes digital better than film-screen are the things that you can do with digital that you just can't do with film-screen.

I am just going to mention three of them today and finish with the one that we are most excited about, which is tomosynthesis.

[Slide.]

There is a technique called dual energy subtraction.  I hope everyone on the committee knows that one of the things that we look for on mammograms are calcium deposits.

In every mammogram there is probably at least one calcium deposit, so calcium, by itself, doesn't mean a whole lot, but it is patterns of calcium deposition that worry us, and ductal carcinoma in situ frequently will cause the deposition of calcium whether it is from dead cancer cells or actual secretion of the calcium.

One of the things that we look for on mammograms are so-called clustered calcifications.  Well, they are hard to see because these calcium deposits are under a millimeter in size, virtually all the time in cancers, and down around 500 microns, that is about a half a millimeter particles, and trying to see them against the background of the very heterogeneous breast parenchyma can be difficult.

One of the things that the computer-assisted detection units, one of the things that they do well is to find these white spots on the mammogram.

There is another approach which is dual energy, which we think will allow us to see them very easily.  What is dual energy?  Well, you take a high-energy mammogram, and at a high-energy/high kV, the difference between soft tissue attenuation and calcium attenuation is not as great as when you do a low kV image, where the calcium attenuation becomes significantly higher.

With a computer, if you have taken these basically simultaneously, you can adjust the soft tissue attenuation numbers, so that you change the image, so that the soft tissue in the high kV and low kV image match up exactly, and you can then subtract those, so that there is no soft tissue image left.

Because the calcium doesn't subtract exactly with these two changes, then, the calcifications will show up as a separate signal.

[Slide.]

This is just schematically how this would work.  You take a low kV image, you take a high kV image, subtract the two.

[Slide.]

If it works, the tumor calcifications will stand out and the breast will disappear.

[Slide.]

This is from an article in the Journal of Radiographics from many years ago where they did this in a chest x-ray.  Just to show you, this is the heart and the mediastinum, and the ribs have been subtracted out.

[Slide.]

That is not quite as impressive as when they subtracted out the heart and mediastinum and left the ribs, and this was done without a scalpel, which is kind of impressive.

This is sort of what you can do potentially with dual energy subtraction, and this is a fairly crude subtracted study we think that we can do better.

[Slide.]

This is a breast tissue sample with high grade ductal carcinoma in situ, that we did a dual energy image with just a standard General Electric digital mammography system, and this isn't just window and leveling, this is actually subtracting out the soft tissue and leaving only the calcium.

[Slide.]

Here is another one with calcifications from a cancer, and you can imagine if, on a mammogram, when you push the button, you saw nothing but where the calcifications are, that would make the radiologist's search a lot easier for finding calcifications.

This is not available on digital systems yet, so you don't have to worry about it yet as the MQSA Committee, but this is one of the things that you can do with digital system that really would be next to impossible with a conventional film-screen system.

[Slide.]

Another sort of corollary to this is digital contrast subtraction angiography.  This is a technique that we use all the time in the rest of the body, and now that we have digital detectors for the breast, investigators are starting to look at it in the breast.

This is a way of looking at only the vascularity. I think most of you know that breast cancers or cancers in general, solid tumors can only grow to a certain size and then they need a blood supply.

They can't just continue to grow unless they call in a blood supply, and that is called angiogenesis, a neo-, a new vascularity that's developed, and this, of course, is one of the areas that people are looking at to try and kill cancers if you can get rid of the blood supply, then, maybe you can kill the cancer, but it is also the way we confine cancers.

[Slide.]

In magnetic resonance imaging, here, on a CT scan, this is back from the 1980s when Chang in Kansas pointed out that cancers will enhance with iodinated contrast on CT. This is a pre-contrast scan and a post-contrast scan, you can see the cancer has taken up the contrast, and is much more visible with contrast.

[Slide.]

This is on magnetic resonance imaging, actually, a little cancer that we found in a study.  Actually, we were studying her left breast because she had a known cancer there, and we found an unexpected and unsuspecting cancer in the other breast, that was found because it enhanced with contrast because of these new blood vessels that had formed, and this was not visible on a mammogram, ultrasound, we could only see it on MRI and CT.

[Slide.]

Here is another case, not that one, but a similar one, where we found it on MRI and then located it on CT because of the contrast enhancement.

[Slide.]

Well, this is a potentially powerful tool if we could do it in mammographic, with mammographic resolution. The CT scan doesn't have nearly the resolution that an x-ray mammogram has, and so if we could take the advantage of seeing small blood vessels at high resolution, we may have a very powerful tool.

[Slide.]

So, again schematically, we would do a pre-contrast mammogram--this is before injecting any contrast--then, intravenous administration of iodinated contrast.  This is the same material that is used every day in departments for CT scans and just about anything that we do that needs what are called contrast agents.  It is generally iodine.

Then, you subtract the pre- and post-contrast, and the only thing that is left after subtracting is the vascularity, and in this nice schematic, it works beautifully.

[Slide.]

This is an animal model that we did a number of years ago.  This is bone.  There is a tumor that has been implanted right here.  It is a little hard to see because we haven't done the subtraction.

[Slide.]

When you do the subtraction, now you can actually see the nest of blood vessels that has developed around this tumor.  This was done with a digital mammography detector, so that the spatial resolution is quite exciting.

[Slide.]

Other folks are now looking at this in the human breast.  Martin Yaffe and his group up at the University of Toronto, using the Fischer system, has shown contrast enhancement of a cancer, and just as we can with magnetic resonance, they can track the contrast as it flows through the tumor, so as the contrast gets to the tumor, there is a sudden increase in contrast in the tumor.

Some of the contrast actually leaks out, so it  stays in the tumor, so the tumor contrast stays high, but some of it leaks out, and this is a fairly characteristic curve for a cancer, which this turned out to be.

We have seen this on magnetic resonance.  We should be able to do this less expensively and at higher resolution with iodine and digital mammography.

[Slide.]

John Lewin has also done some work in this and loaned me these slides, where he is dealing with the issue of registering these images.

One of the big problems that the radiologists and physicists in the audience know about is to do accurate subtraction, where you subtract one image from another, they have to line up perfectly, otherwise, you get all kinds of artifacts.

John has taken another approach where he takes a high-energy image and then a low-energy image, and subtracts them, and that can be done very quickly, just like the calcium that I just showed you.

That leaves the iodine, which blocks x-rays at much greater rate than soft tissue, so then you can have the iodine image, if you will, of the cancer left behind.

[Slide.]

Here is just another case that John gave me, sort of a subtle lesion, and then much more obvious following the contrast.

[Slide.]

Now, what I was sort of asked to talk about today is a technique that we are developing at Mass. General, and I should issue the disclosure that I am, number one, biased because this was our idea for the breast and we hold patents at Mass. General on it, so that I am not totally objective, although I will try to be as objective as I can here.

Three-dimensional digital mammography, I think is going to be the first technique that really sets digital mammography off from conventional film-screen, and there are a lot of advantages that I will show you.

[Slide.]

What do I mean by 3-dimensional mammography?  Well, there is a technique that was actually written about back in the late 60s, early 70s, called tomosynthesis.  What is tomosynthesis?

In conventional x-ray--and we will talk about mammography, which is an x-ray technology--the breast is between the x-ray beam and the detector, and x-rays come through the entire structure of the breast, and something that is in this plane, superimposes on something that is in this plane, and they project on top of one another on the mammogram.

One of the challenges that radiologists face is to try to separate the structures that are overlapping, and one of the reasons we take views from the side and from top to bottom is to help the radiologist understand the 3-dimensional location of structures, so that we are not fooled by overlapping structures.

I should tell you, though, that in our practice, and Ed Sickles, I was talking with Ed, and he has the same experience at the University of California at San Francisco, about 25 percent of the women that we recall because we think there is a problem on their mammogram, are recalled because of overlapping structures.  It turns out to be nothing, but the radiologist, just looking at the two films, can't be sure of that, and so the patients are recalled.

[Slide.]

Now, this just an example of a very subtle tumor. It is not so subtle because we have got a spot compression panel on it, and I wanted you to be able to see it.  Here, you can see a spiculated, fairly high-density mass that is partially hidden because of the rest of the breast, the normal breast structures.

[Slide.]

Now, once this is taken out at surgery, this is following needle localization, it is easy to see.  Why is it easy to see?  Well, first of all, it is a thinner structure, and so there is less scatter, and so on, but we have taken away what is called the structure noise of the breast, which is what is in front of it and behind it on the mammogram.

[Slide.]

I liken that to looking for a birch tree in a pine forest, couldn't find a birch tree in a pine forest, but this is a deciduous, but it is close anyhow.  Here is the birch tree.  It is pretty easy, everyone can see it.  It is white, it stands out against the green background.

But as you start walking along--I had to go all the way to Norway to do this, by the way, it was a tough experience, but I did--as you walk along, it starts getting hidden, and after a while you really can just barely see the birch tree, if at all.

[Slide.]

Here, my diagram of a birch tree in a pine forest, it is very hard to see the birch tree even if it were just one row back behind these pine trees.  Well, how does this compare to the breast?

The breast, you all know about dense breast tissue.  Dense breast tissue are the fibroglandular structures, principally the fibrous structures although the glandular tissue is also dense, that are highly attenuating of the x-ray beam, and cancers have very similar x-ray attenuation to fibroconnective tissue.

So, it is sometimes very hard and often--not often--but not infrequently impossible to differentiate a cancer from the normal tissue, and this is the reason why cancers don't show up on mammogram.  They are hidden by the dense tissues.

This can happen at any age, and the density of the breast, just as a quick aside, has nothing to do with how firm it feels on clinical breast exam.  The density is an x-ray attenuation phenomenon, you can only tell density by doing a mammogram, and the density changes gradually with increasing age.  It is nothing that happens at menopause or age 50, as we have been led to believe.

Women at age 30, about 90 percent of them have dense breast tissue, and that changes by about 1 to 2 percent with each increasing year of age, so that the percent of women with fatty breast increases.  Fat is like window glass or having a birch tree out in the middle of a field.  The density of the breast is like the pine trees.

[Slide.]

So, if we could just get rid of the pine trees without having the environmentalists against us here, we would be able to see the birch trees.

[Slide.]

So, the trick really is to make slices through the breast, and I think many of you are probably familiar with CT scans, which make slices through the breast tissue--again, the radiologists are all well familiar with cross-sectional imaging--and what that does is it gets rid of what is in front and what is in back, so that you can see the birch trees.

[Slide.]

Well, how can you do this in the breast?  In the days when I was very young, we used to move the x-ray tube in one direction, and the detector, in this case it was film, in the other direction, and they would move in tandem, and you can probably imagine that as you move the x-ray tube in one direction and the detector in the other direction, only the things that are at the fulcrum of the motion will not look like they are moving, everything else will look blurred.  I will come back to that in a minute.

But this is the old way of doing conventional tomography, move the tube in one direction, the detector in the other, and you could blur everything but the plane that you were interested in.  If we were doing this in the breast, it would be in the breast.

The problem with this is for every slice that you want to make through the tissue, you have to repeat this sequence at full dose.  So, to go through the breast, say, get 60 slices through the breast, would require 60 mammographic exposures.  Clearly, we are not going to do that, so this approach to making slices through the breast isn't going to work.

[Slide.]

It turns out that using a computer, you can move the x-ray tube through an arc, and what we do in our system is move the x-ray tube through an arc of 50 degrees.  That is just what we started with.  We think that there are better arcs to choose.

We are moving it through 50 degrees, and as we move it, it is stopping and taking a picture 11 times from 11 different angles.  We can do that in 7 seconds.  The breast stays stationary and the detector doesn't move, so the breast and the detector are stationary, just like a conventional mammogram.

It takes a little bit longer.  It takes 7 seconds as opposed to conventional mammogram is anywhere from half a second to a couple of seconds, and we take 11 images.  You say, well, gee, that must be a lot of dose.

Each individual image is a fraction of the total x-ray dose of a film mammogram or, for that matter, a digital mammogram at the same dose.  So, the dose for these 11 images adds up to about 1.5 film-screen mammogram exposures, so the dose is actually less than the tubular mammogram.

[Slide.]

Now, why are we doing this?  Well, if you look at again our objects that are in the breast here, when we take a conventional x-ray, they are lined up and they are hard to tell apart.

As we move the x-ray tube through an angle, you can see--and I will do this back again, so you can see it--that at different angles, something in one plane moves differently than something in the other plane.

Again, we move through the angles, and if you watch, the structures start moving apart.  That is called parallax, and if you close one eye and point to an object on the wall, and then change, your eyes closed, the first eye, and open the other eye, you will see the object seems to shift.  That is because your eyes are looking at it from a different angle, and it is that shift that is called parallax.

[Slide.]

We can take advantage of that with tomosynthesis to make slices through the breast, because we can then take these images that we have put together, and line them up, and this is why we need digital systems because you really can't do this certainly efficiently with film images.  You need a computer image, it is much easier.

[Slide.]

Here, in the synthesized images, now the spiculated, the irregular abnormality, or anything in that plane will register on all the images.  The other ones won't because they are in a different location.  They will misregister and you will have blurring by the misregistration.

Now, let's say we wanted to see the plane that had this irregular shape in it.  Well, we just shift the images and add them again, and now what is in the plane of the irregular shape adds up and is sharply defined, and everything else is misregistered.  That is called tomosynthesis, that was named tomosynthesis.

[Slide.]

Well, we have taken it several steps beyond now.  We don't do that so-called shift and add, but we accomplish the same thing, and this is just to acknowledge the people that we have worked with from Brandeis University to develop advanced algorithms, and I have to tell you that this was done under a grant from the Army.

[Slide.]

General Electric built the machine for us, and basically, the x-ray tube moves up in this housing.  This is just here so we don't whack the patient in the head as the x-ray tube moves, but it is a standard digital mammography system, digital detector right here, compression system. This just has to do with counterbalancing the movement of the x-ray tube.

Again, this is a prototype and not as elegant as a finished system would be, but to the patient, it is like having a mammogram.  Now, I will say just quickly as an aside, we are collecting data to see if we can reduce the compression needed for doing mammography.  I know most women don't care much about that, but we would like to reduce the amount of compression that is needed.

We are not sure we can get away with that.  Right now we are using the standard mammographic compression, but another advantage of tomosynthesis will be that we think we can just do one compression, that we won't need to do the two projections because, as I will show you, we will have 3- dimensional information.

So, the x-ray tube moves up here.  The exposure, as I say, takes 7 seconds.  I think we have had one case out of over 250 that we have done where we think there was some movement by the patient, but they seem to be okay, and again, 11 images in 7 seconds for 1.5 film-screen dose, not film-screen study, but film-screen image.

[Slide.]

Now, it is not going to be great here in this room because the lighting isn't terrific, and I will also give you the caveat that this looks much better on a workstation which is designed for this, but here is a conventional film-screen mammogram, and I am sure all the radiologists picked out this cancer.  Actually, it is a very subtle cancer.

I think we picked it up because there is some calcium that just happened to be nearby and drew the radiologist's eye.  This is a very subtle lesion.

[Slide.]

Here it is up close, a little better, but you can see you have got all the structure noise of the breast that is getting in the way.

[Slide.]

Here it is from the cranio-caudal projection, again pretty subtle.

[Slide.]

Even when you get up close, it is still pretty subtle.

[Slide.]

This is the tomosynthesis image, and, of course, it takes getting used to.  What I am going to do, we are going to be paging through.  We have reconstructed this at 1-millimeter thick slices.  We actually can go down to--we have done half-millimeter, we can go to thinner, it just requires much greater computer calculations.

That is right now the rate-limiting step.  It takes about two hours to process one of these images.  That cuts your throughput down a little bit, but we know how to do it actually within a minute or two, so that is really more the computational power, and computers get cheaper and cheaper by the second, so that is not an issue.

[Slide.]

Now, what I am going to do is just page through, and we are starting to move into the breast, and you will start seeing structures.  I am going to stop it here.  You can start seeing Cooper's ligaments coming into view, these curved lines.

[Slide.]

If you watch right here--it is hard here to stop and start with PowerPoint--but now you can see the cancer without the noise of the breast in the way.  I will show you this up close in a second.  Then, we will page through the cancer and out to the other side.

[Slide.]

Just to show you, this is a little easier to appreciate, these are the blood vessels just under the skin on the other side.  Notice that we just saw them at the end of the paging through, because they are not in the slices that preceded them.

[Slide.]

Let me just show you this up close.  Again, on a workstation, it is much easier to deal with this.  Here we go, just looking at the tumor.  Now you can see I think with much greater sharpness, the spiculations of the lesion, its irregular margins, and just the conspicuity of it.

[Slide.]

We will go to the next.  I tried to put them side to side, and adjust them the same.  This is the conventional mammogram on the left.  Here is the tomosynthesis on the right.  I think the ability to see this tumor is much greater than trying to pick it out of the background there.

[Slide.]

Here is another patient.  This was a great pickup by one of my associates.  There is some funny architecture right in here.  It is actually not this, it is up in here.  Let me see close up.  You can see there is some funny--again, the radiologists I think can appreciate some funny architecture there, everyone else is going what is he talking about.  Anyhow, this is a good pickup.

[Slide.]

Here, in just the projection, I am going to show you in this area, I don't even think I see it on the conventional mammogram, but anyhow, here is the tomosynthesis.  I think you can now see pretty easily the spicules coming out of this tumor, and then somewhat to our surprise, here is a second lesion.

To get into our study, you have to have an abnormality on the film-screen image to begin with, so that this is tying tomosynthesis one hand behind its back, but this we have had now three cases where we picked up a lesion that was not visible on the conventional mammograms.  We actually thought there might be a third one up here, but the surgeon wasn't willing to biopsy that, so we will just have to wait and see, hopefully, it is not.

Just to put them side by side, the conventional image here, and then we will run the tomosynthesis for you here.  You start seeing this condensation of tissue, I mean it just makes it much easier to see these things.

[Slide.]

There is the spiculate of the first one and then the second one right there, and just in case you don't believe me, we got both of them on the biopsy.  Actually, I think they are easier to see almost on the tomosynthesis than there on the specimen, which is a switch.

[Slide.]

This is another case of a patient who came with a palpable abnormality, that we just don't really see on the mammogram, we couldn't see it in the cranio-caudal view.

[Slide.]

Here, on the tomosynthesis, if you watch right where I have got it labeled, you will see the cancer--this was an invasive cancer come into view, and I am not sure what they were feeling.  We looked with ultrasound, and they were close together, but there is a second lesion, a fibroadenoma, which again shows up on tomosynthesis.

Now, we haven't looked at trying to differentiate benign from malignant at this point, but we think that we will have probably have a better shot.  We have done some preliminary reader studies that, in fact, suggest--and I can't give you all the data, because we are presenting it at the RSNA--but that suggests that we see the margins, as you might imagine, of lesions much better with tomosynthesis than conventional, so we think we should be able to differentiate benign from malignant more accurately.

[Slide.]

Just a couple more cases here and then I will wind it up.  This is the case of a patient who came in and had this asymmetry deep in her breast, and she didn't have it on the other side, so we were concerned.  It really wasn't that dense, but invasive lobular cancer can do funny things, so we were going to biopsy this if we didn't have any other information.

[Slide.]

The problem was that we couldn't see it on the cranio-caudal view.  We did extra views, and we really had a little bit of a quandary.  She agreed to have the tomosynthesis.  I put this in motion, you watch up here, you will see the lesion come into view.

Then, for the radiologists in the audience, you will notice--I don't know if you can see from the angle you are sitting up there, the panel--you will notice that there is a nice capsule around the abnormality, which we could not see on the conventional imaging.

We know exactly where this is now, because we know it was a 6-cm thick breast, and this was at millimeter 33, so not only do we know where it is 3-dimensionally, which we couldn't tell from the conventional mammography, but we also know that it is a mixed density lesion.  It has got a pseudocapsule around it with fat and dense tissue, and that to a radiologist means it's a benign hamartoma.

We were able to get this patient's old mammograms from California from eight years go, much to our surprise, and this was there eight years ago, so we are comfortable in just leaving this alone.  Just based on the tomosynthesis, we would have been comfortable leaving it alone, but knowing it has been there for eight years unchanged, that confirms our suspicion.

[Slide.]

I think this is the last case.  This is a patient who came in with a palpable abnormality.  These are the mammograms that you take off the pile and you put down below that you are not going to get to that day, so that your associate can read them the next day.

Very dense breast tissue and just very hard to tell what is going on.  Someone thought they felt something. There actually is a little architectural distortion here, and again you can sort of see it in here.  Everyone is going yeah, right.

Here, just a little bit digitized representation. Something up in here maybe, hard to tell on a cranio-caudal view.

[Slide.]

Here is the tomosynthesis.  As we page through, just showing the whole breast for a second here.  I think you can all see the strands coming out and the lesion right here.

[Slide.]

Just make it a little bigger and side by side.  It is very hard to see much on the conventional mammogram, page through.  Again, with the workstation, you can go back and forth, so it is not hard to see, but here, you can see the spiculations that really aren't even--I don't think you can see them there--much easier lesion to see.

We actually thought that there is another lesion here, and we are in the process of going through the pathology.  This turned out to be an 8-mm invasive cancer with DCIS, and I think what we are actually seeing is DCIS as ductal extension here.  We have got to confirm that with more detailed pathological review.  This is just the lesion on ultrasound.

[Slide.]

Just one more case to show you some of the other features with tomosynthesis.  Vary again, another one you put to the bottom of the pile, there is just all kinds of stuff going on here.  There is some funny architecture right here.

[Slide.]

Here is the tomosynthesis.  As we page through, you will notice that some of these are coming into very sharp detail with very sharp margins, and these turn out to be cysts.  There is another one here again as you page back and forth, very easy to see.

Then, we come into this spiculated architectural distortion.  Again, in this room, with this light, it may be hard to see, but much easier to see with all the structure noise moved out of the way or cut out of the way with tomosynthesis as we page through.

[Slide.]

I will skip that.

[Slide.]

We have actually done reader studies now with tomosynthesis and the lesions are much more conspicuous with tomosynthesis.  The borders of the lesions are more clearly defined.  We virtually can eliminate recall for superimposed structures because there aren't any.  When you are slicing through, you have eliminated everything that is in front or in back, so anything that is there is in a plane, and is not superimposed structure.

[Slide.]

The problem of where is it 3-dimensionally will go away because if you can see it in one view, you can figure out by what slice it is on where it is 3-dimensionally.  We think--we haven't proved this yet--but we think we will be able to better differentiate benign from malignant.  That, to me, would be nice, but it has got to be at least as good as a needle biopsy before I would rely on it, but that may be another feature of tomosynthesis.

The issues that we have to deal with, of course, are two hours for reconstruction per study is a little bit too long, but we have already done the math, if you will, and the computer design that will allow us to reconstruct these in just a few minutes, probably one to three minutes per image, which would be like the old days of processing a mammogram, and with faster computers, we can get that down even more.

[Slide.]

The difficulty now is that instead of one to look at, the radiologist has 60 more, or 120, or whatever, you know, your slice thickness is, and we think that there are ways of dealing with that.

It is actually not that bad.  You can go through these very quickly in a workstation, you know, you can go back and forth instantaneously or slow it down, whatever you want, or once you have--and this is great after lunch, we will have people barfing in the corridors--what you can do is take these slices and put them back together as a 3-dimensional projection.  This is what is called a maximum intensity projection.  There are different ways of doing this.

This is just a patient who had actually a previous biopsy.  You can see where her skin is thinned right here.  It is a benign biopsy.  This is just some postsurgical change from the biopsy.

But you can get an appreciation of how you could take these slices and put them together, so that the radiologist could very quickly look at this image and the computer programs are available today where you can just sit there and turn the breast as you want it.

Whether this is the way we will look at them, whether the slice is the way we look at them, I am not exactly sure, but having worked with this system now for several years, I am convinced that this is the way we are going to be doing mammographies.

If you think you had problems regulating conventional and the digital, wait until you get to tomosynthesis.  I can imagine the issues that we are going to face in terms of quality assurance, but there is little doubt in my mind that the sensitivity of tomosynthesis--I mean I am biased, you know that--but I think everyone who has looked at it has agreed that the sensitivity of tomosynthesis will be higher than conventional and digital mammography, and the specificity will be higher, as well.

We know we can eliminate 25 percent of the call-backs, so that right there is pretty desirable, and then the other points that I made.

Again, I appreciate your inviting me here, and I would be happy to take any questions.

DR. PISANO:  Dan, do you do all the images always in the oblique projection?

DR. KOPANS:  Yes.  So far we have only done them in the oblique projection.  Our thinking has been that one of the advantages of tomosynthesis would be that we could eliminate having to do two compressions, which in and of itself, I think women would appreciate.

So, we have really done most of our work that way. We have talked about it, maybe it would be even better doing it in two projections, and we will look at that in the future.

MS. HARVEY:  Dr. Ikeda.

DR. IKEDA:  How are you archiving these, and how big are the files?

DR. KOPANS:  That is a good question.  The files are whatever a digital mammogram is times 60.  I should have pointed that out, and thanks for the question.  These are done at the same spatial resolution as the General Electric detector 100 micron pixel size, so it is a very big file.

Right now we are archiving them on CD.  That is a good point, but archive gets cheaper and cheaper every year, too.

DR. IKEDA:  And you are displaying them on a regular GE Advantage workstation?

DR. KOPANS:  We are displaying them on a 2K by 2K monitor.  I have forgotten whose monitor it is.  It is not GE's.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  Dr. Kopans, we have also done tomosynthesis, and we see your excitement in that area.  I would like to ask you, how do you envision that we would be using tomosynthesis in a few years as the technology matures, because it is very difficult between diagnostic or there are certain groups of patients that you might say I want to go to tomosynthesis straight, bypassing the normal mammogram?

DR. KOPANS:  It is a good question, and of course, the question.  In my mind, finding early cancers is the only reason to image the breast.  I mean all the diagnostic imaging that we do, I think does have some benefit for patients, but the real benefit is in finding cancers early and saving lives.

As we look at tomosynthesis, we see it as the screening test.  It was interesting, I thought, well, it is only going to be in the dense breast, but even lesions in fatty breasts, small lesions in fatty breasts are much more conspicuous on tomosynthesis than they are on conventional mammography.

So, I don't have a feel yet as to whether you do the fatty breast with conventional digital and you do the tomosynthesis in the dense breast.  My prediction is--again, I am biased, but I think I am going to be right--is that it will become the screening mammogram.

Now, proving that is a monumental task.

MS. HARVEY:  Dr. Ikeda.

DR. IKEDA:  Have you been able to display side by side, left and right breasts, because oftentimes we look for symmetry.  Probably--I don't know if the computing power is there yet--but many people are thinking about, instead of unilateral MRI's, doing bilateral studies to look for symmetry, which can be a help sometimes.

DR. KOPANS:  Absolutely.  We have put them up, but what we end up doing is that everyone concentrates, you see so much detail that it is kind of like you almost forget about the other breast, but that is clearly a study.

Another thing that I didn't mention that we think is going to be very valuable is that computers now, and computer-aided detection, can look for morphologic features. They look for white spots, which are calcifications.  They look for certain linear projections to look for spiculation. They are not very good at looking at masses, but they can't look at last year's mammogram and see if there has been a change to this year's mammogram, you just can't do that with 2-dimensional imaging.

We think with 3-dimensional imaging, we will be able to teach computers to look for changes between last year's tomosynthesis and this year's tomosynthesis, because you have a 3D dataset that can be warped and registered, so you can see what has actually changed.

So, we think that having 3-dimensional datasets is going to open the door even further for computer-assisted detection and diagnosis.  I hesitate to talk about diagnosis because our needle biopsy techniques and even localization and surgery are so accurate and safe, that you really have to have a diagnostic system that was like 98 percent accurate to do away or even more than that, 99 percent accurate to do away with biopsies.

But in terms of finding cancers, having a 3-dimensional dataset and then adding, for example, contrast agents to the tomosynthesis, adding ultrasound to the tomosynthesis, which we think we can do in exactly the same position, so that everything is perfectly registered, a lot of opportunity for investigation.

That is why I say that, you know, digital is just in its infancy, and all the different things that digital is going to allow us to do are what are going to make it beneficial, not just that it is as good as a film-screen mammogram.

MS. HARVEY:  Dr. Ramos.

DR. RAMOS-HERNANDEZ:  Can you talk about cost, speculated cost?

DR. KOPANS:  It hasn't cost me a cent.

[Laughter.]

DR. KOPANS:  Cost, again, you really don't know until the companies really get involved, and we are trying to get the companies involved.  It will be more than a digital mammogram, but the real expense in digital mammography is the detector.

So, once you have the detector, what we are doing is actually, we asked for $25,000 from one of the companies to motorize the x-ray tube, we are going to do it ourselves, and they sent all kinds of people to talk to us.  They must have spent well over $25,000 visiting us, didn't give us the $25,000, and then we went out and got a grant, and the company got a million dollars to do it, so they were smarter than we were.

But it is just moving the tube, that doesn't cost much.  The computers are getting cheaper and cheaper and cheaper, so if a digital mammography system retails for, what, $400,000 or something, this might add $100,000 to it, but that will come down as the computers get less expensive and more systems are purchased.

We really haven't done an in-depth cost analysis because we are really in the very early study phase of this, but I don't think it is going to add that much expense, I mean per patient, that actually doesn't, that is a few dollars per patient.

MS. HARVEY:  Dr. Henderson.

DR. HENDERSON:  Jessica Henderson.

Just out of curiosity, the patient who had three suspicious places--

DR. KOPANS:  No, two, we don't know about the third one.

DR. HENDERSON:  Why did the surgeon only biopsy two?

DR. KOPANS:  Well, it gets to be tricky when you are doing research protocols, can you use the research to take care of the patient, and there are Institutional Review Board policies that start getting in the way of taking data from--I mean no one has ever done this before--so, we can't say to the surgeon, you know, we have got a track record, we know that is a cancer, and have the surgeon do a second biopsy.

So, it becomes an ethical problem.  There was enough debate in the group, that no, it isn't, yes, it is, no, it isn't, that we felt it was reasonable to follow her. She has got two cancers anyhow, she is being radiated and treated.

The issue of finding additional foci of cancer, I mean my excitement in finding those two cancers is that tomosynthesis found a cancer we didn't know about.  You know, the whole issue of do you really need to know all the cancer that is in a breast, this is going to sound a little strange, but that may not be a good thing.

For example, with magnetic resonance imaging, people are finding more cancers or more foci of cancers in a breast than they originally thought, so the patient, instead of having her breast conserved, is having a mastectomy.

Yet, at least in our practice, the recurrence rates for conservation therapy are incredibly low now.  Our radiation therapists just looked at their data, and it is 2 percent at eight years, which is very, very low.  So, maybe those cancers that we don't find now are being killed by the radiation, and finding them may be doing a disservice to the patients.

There is a lot of issues that come up in the issue of multifocality or multicentricity in terms of cancer, so it gets, you know, when you are doing a research project, it gets even more complicated as to how you deal with that information.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  Just as a follow-up to that, you probably did all the other things you would normally do with that area, right?  I mean just to clarify that.

DR. KOPANS:  Yes.

DR. PISANO:  You probably did extra views and ultrasound and all the things, and that is why the surgeon didn't want you to go after it.

DR. KOPANS:  Right.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  We often look at the cost of the procedure and the technology, but, Dr. Kopans, what do you think about the utilization or are there any potential costs to be saved if cancers are detected earlier, or of equal importance, if cancers or lesions are managed better, if you have better specificity, that way you might avoid procedures?

DR. KOPANS:  I think those are all very good points.  As I said, 25 percent of our recalls are for women who turn out to just have superimposed tissue, and we just have to get some extra mammographic views to kind of look around the trees to make sure that those are just superimposed, and not an actual abnormality.

So, eliminating 25 percent of recalls is a desirable thing from a cost-benefit point of view.  Then, better management of patients, those are sometimes hard to quantitate, but I think that my impression based on the work we have done so far is that our sensitivity will go up, so we will find smaller cancers, more small cancers, which I hope will translate into more lives saved.  We are already seeing the decreased death rate in the United States from screening.

I think that will help us with some of the cancers that we don't find now by mammography, and certainly don't find early enough, and then having the specificity improve will reduce some of the secondary costs of screening, as you point out.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  I have another question.  What is the time line or what is the status of this technology in terms of FDA approval?

DR. KOPANS:  Oh, FDA approval.  It has to be approved by the FDA?

[Laughter.]

DR. KOPANS:  That will be up to the FDA, and I don't want to speak for them.

DR. PISANO:  Well, it has been submitted.

DR. KOPANS:  My guess would be it's at least a  year for--I am being optimistic--a year for approval if we can get everything going very quickly, probably more like two years, and then approval for what is going to be the issue.  Obviously, it will be approval as a diagnostic device.  You can't do a screening test with it.

We will have to sit down with the FDA and figure out how we decide whether it's used for screening, because this really is different.  I mean I think most people who know me, know that I think making digital mammography have to go through a PMA process was a major mistake that the FDA has done, and it is going to make it very difficult to improve the conventional digital technology.

The reason I argued against it was that digital mammography is the same as film-screen mammography.  I think D-MIST is going to show that.  The other studies that have been available have shown that they are really the same.

This is different, so this is going to need a PMA and all the things that go along with that.  I mean I would like to see it out there in between three to five years, shorter if possible, I am not optimistic it can be shorter, but we will see.

MS. HARVEY:  Dr. Harrison.

DR. HARRISON:  This is fascinating.  You made a comment that you envision getting to a point where resolution of malignancy, if it gets up around 99 percent, we could proceed without tissue.

Do you really think we are ever going to get there, considering that many of the subsequent management decisions are based on histologic findings?

DR. KOPANS:   No, I think that is a good point.  I think that histology, although there are technologies that are now looking at this, spectral analysis using lasers, we fiddled around with this a number of years ago, where you put a needle in and try and get the spectral analysis.

All of our therapy, as you point out, is based on histologic analysis and margin analysis, and that is also going to get in the way of in vivo ablation, you know, what is the margin analysis, and so on.

That is why I am almost discouraged that we have to go through the diagnostic route to get technologies approved, when it is really screening that is going to be beneficial, but I understand the reasons, and they are good ones.

MS. HARVEY:  Thank you.

DR. KOPANS:  Thank you.

I believe we will have a break now, come back about 10 minutes after 3:00.  Thank you.

[Break.]

MS. HARVEY:  We are having a talk on the Inspection Demonstration Project.  It's an update by Charles Gunzburg of the Division of Mammography Quality and Radiation Program.

Welcome.

Inspection Demonstration Project

DR. GUNZBURG:  Thank you very much.

I am going to hopefully walk through this pretty quickly and let you know what our program is and I guess the who, what, when, and where of the program.

[Slide.]

When Congress first passed this Act, they included the requirement that there be annual inspections of all certified facilities.  Initially, the compliance rate, not the noncompliance rate, but the compliance rate was relatively low, but as soon as the facilities became aware of what we expected and became more familiar with the regulations, the compliance rate rose pretty dramatically.

Many facilities and some professional organizations were concerned, that were actually hopeful that we could inspect on a less frequent basis and actually save them some money and time.

We pointed back to the Act and said no, we can't do that, we have to do this annually.

[Slide.]

So, they went to Congress and they talked to them, and asked them to do something about it.  Congress listened and when they passed the Reauthorization Act in 1998, they kept the annual requirement, but they added a provision for an inspection demonstration program, and that would be a test program under which certain facilities could be inspected less frequently than annually.

[Slide.]

It couldn't be implemented before April 2001, so that was easy.  Facilities had to be substantially free of noncompliances, and the number of facilities had to be a statistically significant sample of the facilities available.

They didn't specify an inspection frequency, but they said that it had to be one that was capable of reasonably assuring compliance with the standards.

[Slide.]

So, FDA took that guidance and came up with some criteria of their own, and that was that the States and facilities would be selected according to a specific written criteria.  We would include both study and control groups, and conduct inspection of the study group every two years or at least two years on the first basis, and annual inspections of the control groups.

[Slide.]

The participation criteria for the States, first of all, the State had to be willing to do it, they had to agree to do it.  They couldn't have any laws or policies or requirements that meant they had to go to the facility more frequently than we were specifying in our project or our plan, and that was going to be two years again.

They had to agree to inspect at the frequency designated by FDA.

[Slide.]

They also had to be able to accept modifications to their contract, all these States that have contracts with us, about how many facilities they would inspect annually, and they had to be able to absorb the loss of income from not inspecting these facilities, and be willing to do that. They had to notify us if they found any problems that were important.

[Slide.]

That left us with 14 participating groups.  We have 11 states and 3 other testing jurisdictions.

[Slide.]

For a facility to be included, they had to have undergone at least two inspections under the final regulations, two annual inspections.  They had to have no citations in the last two inspections under the final, and they had to have never been considered, received, or being considered for any regulatory action by FDA.

[Slide.]

They also had to anticipate providing services throughout the length of the program, and they had to maintain their accreditation and certification, and lastly, they had to be kicked out of the program, the random selection process as being one of the facilities to be included.

[Slide.]

That left us with this.  It varies between one facility in some states to 24 in Ohio, and if you lump New York together, 35 facilities in New York State.

[Slide.]

Currently, we have notified everybody that we could think of that needed to know about this program, that it was in process.  We have 155 facilities, about 155, and approximately 155 facilities in the study and the control group.

Those numbers are approximate because we never know when somebody is going to drop out.  So, we don't know what we are going to finally end up with, but somewhere in that region we hope.

We actually began this process in May of this year.

[Slide.]

Where are we going with it?  We are going to continue the annual inspection of control groups.  The changes that this necessitates means we are going to have to make some changes in the procedure and the software.  We hope to have those changes in place by January, and we hope to have them tested shortly thereafter, and begin testing the study group facilities about the middle of the year.

Hopefully--this is not a real firm number--but we hope that it works.  It should finish about July of 2004 and begin data analysis at that point.

That is all I have.  Questions?

MS. HARVEY:  Any questions?

DR. GUNZBURG:  Good.  Thank you.

MS. HARVEY:  Thank you.

Dr. Chakrabarti and Ms. Butler will talk to us about full field digital mammography, accreditation and certification update.

Full Field Digital Mammography

Accreditation and Certification Update

[Slide.]

DR. CHAKRABARTI:  By now you have heard this several times, that GE's system was approved first by Office of Device Evaluation followed by Fischer and then Lorad, and you also know that there is no accrediting body, Dr. Finder explained that at the beginning of previous session, and we provide approval based on extending the existing screen-film certificate.

I will briefly discuss and summarize the approval process here.

[Slide.]

Until otherwise notified by FDA, an FFDM unit will be exempt from the MQSA accreditation requirement, and until FDA issues such notification, a facility must request FDA to extend its screen-film certification to cover its FFDM units.

Requests for FFDM certification extension need to supply all the information listed in the document MQSA Facility Certification Requirements.  This is on our web site, also the facility request has to provide application form, application package.  There is that information about the facility certification requirements.  I will give the gist of that here.

[Slide.]

In that requirement, the facility needs to furnish facility status information, FFDM Unit identification, digital image receptor identification, identification of printers for hardcopy output, monitor identification if softcopy display is available, phantom identification, personnel qualifications.

[Slide.]

Phantom image, personnel information, Report of Mammography Equipment Evaluation, manufacturer's quality control program.  I will take a brief pause and I will mention that we need, in FDA, we reviewed this with a phantom image, we need to have a phantom image.

Also, this third bullet, which says the Report of Mammography Equipment Evaluation, and that is 900.12(e)(10) of the final regulation, that is applicable to all modalities, but in case of digital, a new modality, the No. 4 bullet, which is very, very important, that means all QC programs, equipment evaluation must be performed according to the manufacturer's requirement.  That is the 900.12(e)(6).  That is in the final regulation.

So, when we review the Mammography Equipment Evaluation Report, we look at whether the facility has performed their tests according to the manufacturer's requirement of that FFDM system.  That is very important.

Also, we need the signature of the lead interpreting physician.  That signature tells us that all the information provided are true.  We don't need to have all the documentation about personnel qualifications.

We need the list of the personnel who will be working on the FFDM system, but whether that person is qualified to perform and have additional 8 hours of training, that is being sort of guaranteed by the signature when the inspector goes to the facility, inspector verifies those documentation or at the station of the personnel.

Now, once all this informations are furnished, we review the equipment evaluation report and phantom image, everything is satisfactory, we send a facility a letter of approval mentioning that your FFDM system is included with the conventional screen-film mammography certification, and that letter goes out in the name of our division director, and then the facility can perform using efficient imaging with the FFDM system.

[Slide.]

If a facility  receives a letter of acceptance, the approved FFDM unit will be added to the facility's certificate.

The facility must maintain its accreditation status for at least one screen-film unit in order to keep its certification alive and then can continue to utilize FFDM unit.

The facility is also subject to an annual on-site MQSA inspection of its FFDM unit at the same time its screen-film units are being inspected.

[Slide.]

The FFDM unit must be located within the same inspection jurisdiction as the certified screen-film unit. In most cases, this means that the FFDM unit must be located in the same state as the certified screen-film facility.

The lead interpreting physician must oversee the quality assurance program for both the screen-film and the off-site FFDM units.  That is very important.  We want to make sure that we make one person, who is the lead interpreting physician, responsible for overseeing the QA program in both screen-film and the FFDM system.

In general, we respond in three to five days from the time we get the application.  If the application is complete, I showed you the gist of the information that they have to furnish, plus if the mammography equipment evaluation report is complete and perform according to the manufacturer's requirement.

Indeed, there is the problem that we see, that for the first time facilities and first time medical physicists, there is physicists having a number of problems providing the proper MEE for FFDM system.

[Slide.]

I will go over a few tests where we see the problem.  Now, here is a GE system, GE Senographe 2000D. If you look at that, you see that GE requires that signal-to-noise ratio should be over 50 in an AOP mode, standard mode and SNR check.

Now, if the facility does not, the medical physicist does not perform the test and do not use the raw image, they use process image, they get more than 100 percent inflated value of SNR.  So, it does not tell us whether the SNR will pass the minimum of 50 requirement or not.

So, then, I call the physicist, they have to come back and redo the test, and this has happened a number of times.

Then, there is another one in this, that if you look at that, the GE requires the kV must not change, kV should be same if you are in the 2.5, 4, and 6 cm thickness, kV should be the target filter, and kV should not change.

There can be a range of mAs value, however, kV should be 27, 28, 32, depending on what thickness you use. We have seen in some cases, that that has changed, and we have to discuss with GE, and then that took a little more time for approval process.

[Slide.]

Another thing that I will mention in this regard is that the dose has to be performed in the AOP mode, three OP modes, and many first-time physicists haven't done that, have simply provided the dose value, but we are finding out that that is creating problem also, creating problem for us to give approval on time, which we believe we can do it within three to five days if the report is complete and tests are performed according to the manufacturer's specification.

[Slide.]

Now, these are the three tests which must be performed with raw images.  Very recently I am seeing even the signal-to-noise tests are different values than what is coming forth with the raw images, because some physicists have done it with process image.

So is MTF measurement and AOP mode and SNR check, all three tests must be performed in raw images, and it is very clear in the manual, and some physicists are missing it.

[Slide.]

I will mention here there is another reason that MTF values, some physicists have reported MTF value more than 100 percent.  One way you can do that, if you do not put your elliptical region of interest tightly within the bar pattern, if it goes beyond that, then, your value would be very incited [ph], and we can look at the number, and you can see that these are not done properly.

[Slide.]

Another thing that is very important and that physicists are missing, these are the list of things that GE wants and must be performed - room layout, room description, why this is important, because if all these things are moved around, then, particularly the dark level from the monitor will change drastically and the calibration would be disrupted.  So, you want to see that this information is provided properly as mentioned in the manual.

[Slide.]

Now, there is new change, there is revision of GE's manual is out, and I am seeing that many physicists are still using the older manual.  In the new manual, the physicists are not supposed to be doing any calibration, performing calibration, but they will check the calibration as performed by the service engineer, and they will perform the records of five reference luminance levels as given by this curve.

Many physicists are missing that, they are not giving those values or they are getting wrong numbers.

So, these are the things that are necessary for GE system to get approval within the proper time.

[Slide.]

This is the Fischer system.  Here, I see that many physicists are not performing the system resolution test properly.  They are using simply bar pattern and then counting the number.  They must be following what the manufacturer says in their manual.

[Slide.]

Also, I see that artifacts are not far from recording the window width that is specified in the manual, and this flat field tests are not performed at the region, at the corner and the center that the manufacturer requests.            Thirdly, the Fischer system has a contrast disc, unlike the GE system where GE prohibits use of a contrast disc with their phantom.

So, those are the things that there is a difference from one manufacturer to another manufacturer, and even for the same manufacturer, there are changes in the manual, the physicists and the facility must be aware of that and must be performing mammography equipment evaluations and the QC according to that, and that is the cause of the delay many times.

Sometimes facilities call us as soon as they send this thing, and says we are already schedule patients, and unless they have their report properly done, we cannot give approval.

Any question or should Penny speak first?

MS. HARVEY:  Any questions?

Thank you.

Sorry, Dr. Karellas has a question.

DR. KARELLAS:  Kish, according to FDA, are facilities required to have a printer?  I know the use of the printer and I realize that without a printer, it is going to be very difficult to operate, but does FDA require it?

DR. CHAKRABARTI:  Yes, we require for the foreseeable future, we require the original mammogram must be provided in the form of hard copy.  The facility has to have a printer available for a hard copy printout when the patients ask for the original image.

DR. KARELLAS:  What if the facility has an option of printing off site upon request, does the printer have to be on site,  or what if the printer is within the broader institution, another building or something like that?

DR. CHAKRABARTI:  That will work out fine, but when you apply, there was a prior mention of the list, the printer, you have to mention that that printer number is this, we have to make sure that the same printer as manufacturer, is comfortable with the manufacturer's system, and if it is available off-site, then you mention it is available there, and that will work out.

MS. HARVEY:  Thank you.

Ms. Butler.  Good afternoon.

MS. BUTLER:  Good afternoon.

I am Penny Butler from American College of Radiology, Senior Director for Breast Imaging Accreditation Programs.

I thought I would give everybody an update on where the ACR is with full field digital mammography accreditation.

[Slide.]

The last time we spoke, the full field digital accreditation module, it is not a new accreditation program, but it is a supplement or a module to the mammography accreditation program that has been in place since 1987.

It was complete and it was midway through our internal ACR leadership approval.  We had really hoped it was going to be out, be approved by the ACR leadership by September of 2001.  I think you all know perhaps what has slowed that down.

The module that we developed was manufacturer-specific.  GE was the first FDA-approved FFDM system, and the reason for this were multifaceted.  First, the exposure control mechanisms are different, meaning that our instructions to facility in order to make phantom and dosimetry measurements had to be unique to the manufacturer, and due to the FDA regulations, the required quality control is different.

[Slide.]

Since that time, in early October, ACR sent the full field digital mammography module document to the Executive Committee of our Board of Chancellors, and also at the same time, to the FDA for review.

In mid-October, these documents were approved by the Executive Committee.  In mid-November, FDA had instructed us to submit a formal application for approval of the full field digital mammography accreditation module, and this formal application needed to include a number of requirements that we hadn't addressed when we sent them the documents.  We weren't aware that we needed a formal application, and these were basically to address the elements that were in Part A of the regulations, similar to what we addressed in the accreditation body approval application that they approved on December 20th, 2000.

[Slide.]

At the beginning of July 2002, we submitted a complete formal full field digital mammography accreditation module application to the FDA.  At the end of July, after initial review of the application, FDA advised us that there was some of the information provided with an alternative standard request that we had submitted was insufficient.

This alternative standard request had to do with the exposure of the phantom and acquisition of dosimetry data during our mailed accreditation process.

In early August, I worked with some members of the Digital Subcommittee to collect some additional data to supplement this alternative standard request, and right now the material is under revision internally before we forward this to the FDA.

[Slide.]

So, that is the current status of the accreditation program.  I would like to review some of the  proposed accreditation process for full field digital mammography.

[Slide.]

In general, our process is not going to differ than what we do for film-screen mammography.  The paperwork that a facility submits to us is going to depend, just like film-screen, on how much time the facility has left on their MQSA certification and their accreditation.

In general, if they have less than 13 months left on their accreditation, all the units at the facility go through an early renewal process at the usual fee.

If they have greater than 13 months left on their accreditation, they will complete what we call a mid-cycle, we call it the New Unit Addendum Process at a reduced fee. The fee for accreditation for full field digital is not going to be any different from film-screen.

At that time, once the program is approved, the facilities will be able to have stand-alone digital systems and no screen-film will be required within the facility or associated with the facility as Kish has just described.

Keep in mind that right now, since the first application is for GE, right now we are talking about GE systems.

[Slide.]

The clinical image quality evaluation will not differ.  Our Digital Subcommittee and our Committee on Clinical Image Review have determined that they will be evaluating the same eight attributes in exactly the same way, and that is position and compression, exposure, contrast sharpness, noise, artifacts, and labeling.

All images will be submitted on hard copy at this time, and all of the ACR reviewers are qualified in digital mammography under the MQSA requirements.

[Slide.]

Phantom image quality evaluation is not going to differ.  Again, they have to be submitted on hard copy.  The scoring is going to be the same as with screen-film, that is, fibers, specks, masses, and the subtraction of artifacts.

We have made a minor revision to how we evaluate the phantom image quality to take into consideration some of the very special artifacts that you might see for digital, but they were relatively minor, so we could just supplement our standard evaluation form.

Also, as with clinical review, the ACR reviewers are qualified in digital under MQSA.

[Slide.]

So, if all these things are the same, why do we need a separate accreditation program or a separate module to accredit digital?  Well, most of this falls down to 900.12(e)(6), as Kish had described and let me just read this to you.  "For systems with image receptor modalities other than screen-film, the QA program shall be substantially the same as the quality assurance program recommended by the image receptor manufacturer except that the maximum allowable dose shall not exceed the maximum allowable dose for screen-film."

[Slide.]

So, we are working with the screen-film dose limit, but the QC program as specified by the manufacturers.  So, let's talk about the phantom exposure and dosimetry.

As you aware, ACR has a male dosimetry program.  We tried to get radiation dose estimates concurrently with the phantom image quality, and so these are done at the same time, so we can provide better information back to the facility of what possible causes for poor image quality may be, and certainly a dose that is too low is a strong reason for why image quality may be poor.

We do this currently through a mailed TLD.  The TLD dosimeter is in a little holder, it is several millimeters thick, and it is placed typically upon the phantom.  With the GE system, the exposure control mechanism that is typically used under an AOP is very different from what is used for a film screen.

The exposure control mechanism is different among the different manufacturers.  Some of the manufacturers are using just strictly manual techniques right now.  Consequently, our instructions to the facility have to be unit-specific, and we are very conscious about sending written instructions out to facilities because it is very easy for the technologists, who are usually the ones doing all the work, if the instructions are too technical or too physics like, it could be not clear enough.  We can get some very strange numbers back.

The GE exposure is impacted by the thickest or the densest part of the breast, and if you use the routine phantom, the plastic rim around the wax block that is in that phantom, and on top of that, the TLD holder that we use, it can result is a significantly higher exposure than one would anticipate under film-screen conditions, for example, for the 4.2 cm breast.

So, we have revised our instructions for GE to have the facility first expose an acrylic block, and that acrylic block is equivalent to the 4.2 cm tissue, which is what the center part of the ACR phantom is equivalent to under AEC to determine the appropriate technique.

Then, we asked the facility to exposure the accreditation phantom, and a dosimeter with the manual technique which is closest to the technique that came up under AOP mode.  This is one of the items that we are working with FDA on to revise, to make sure that it is appropriate under the regulations.

[Slide.]

In addition, we have tests listed in our application materials that the facility must submit either information on or a checklist showing that they do these tests, and these are specific to GE.

[Slide.]

Likewise the medical physicist test, they must submit equipment evaluation and if it is an annual survey, the annual survey report showing that they have performed all of these tests and all the tests appropriately meet the regulations.

[Slide.]

Once we receive FDA approval for the first manufacturer's module, the GE module, we are going to complete development of the modules for the other FDA-approved units.

This is my last slide, but I would like to add that some of the comments that were brought up here earlier regarding harmonizing the QC test among the different manufacturers wherever possible is something that would enable us to operationally make evaluation of applications submitted by facilities under the accreditation program a whole lot easier, and we fully support that effort.

Any questions?

MS. HARVEY:  Do we have an idea how long it will be before there is going to be an approval, like within the next six months, a lifetime?  This is the question I probably get most frequently these days.

DR. FINDER:  I would like to give you an answer. It is a process that is ongoing.  Obviously, both sides here are trying to accomplish this as quickly as possible.  We all understand the implications of having an accreditation body or not having an accreditation body.

I cannot give you a specific date or a time, but I can tell you that everybody is working as hard as they can to get this done as quickly as humanly possible.

MS. BUTLER:  Obviously, FDA, the next step, well, FDA is concurrently reviewing our full application that we sent in, and we are working to provide the supplemental material to them as quickly as possible for the alternative standard.

MS. HARVEY:  Thank you.

Any other questions?  Dr. Gray.

DR. GRAY:  Joel Gray with Lorad Corporation.

I have two questions, one for Penny and I believe one for Charlie or somebody.

You indicated that your image quality test is going to remain and the same, and the question is will the requirements for fibers, specks, and masses, 4, 3, and 3, remain the same, and the question for Charlie or someone, does this mean that each accrediting body is now going to have to go through and complete this process that ACR is, so you are going to have to go through this process of approval four more times?

MS. BUTLER:  I will take the easy question.  For GE, the standard was 4, 3, and 3.  That is what is in their QC manual.  That is what we were moving forward with.

DR. FINDER:  With respect to the accreditation bodies, the accreditation bodies are free to apply for FFDM. If they do, they will have to go through a process similar to what we are requiring of everybody else.  It is the same process.  So, the answer is yes.

MS. HARVEY:  Ms. Martin.

MS. MARTIN:  Penny and I both spent the last couple of days going to a class on the physics of digital mammography and how to do all these wonderful tests.  One comment that I got consistently from most of the attendees, and I would just pass this along, I am not sure where it will go, is that from what we could see, most of the units could do a 5-4-4 score, and we were wondering why that was not set.  If the digital is capable of being better, why are we setting the score so low for the phantom image, because if you can't get a 5-4-4 out of it, you really don't have your digital set up right.

I mean why was it set so low?

MS. BUTLER:  Well, we are going with the GE Quality Control Manual as far as meeting the specifications. That was outlined in that.

DR. PISANO:  I know Penny knows about this, and maybe some of you also do.  There is a phantom that has been developed for D-MIST, for digital mammography, which they are calling MISTY, which is really a much more challenging phantom for digital mammography, and we are going to have a lot of data on its performance across the trial through the D-MIST, that same presentation that I referred to earlier.

I think Dan Kopans commented that really, this technology makes the ACR phantom somewhat archaic, it's not challenging enough for digital, so perhaps over time there can be a evolution to another standard or any other phantom. I like what Penny has done or what the ACR has done, it is just adopt what GE did before this phantom became available, and I think that is a reasonable first step myself.

MS. BUTLER:  To expand on that a little bit more, what the subcommittee has been talking about is the current phantom the appropriate one, and I think there is a prevalent thought that there probably could be a better phantom out there, but being involved with developing new phantoms in the past and having them adopted is not something that happens overnight.  You think this process is long, try to develop a phantom.

To get going, we are staying with the ACR phantom the way it is right now and then hopefully look at this in the long term.

MS. HARVEY:  Dr. Karellas.

DR. KARELLAS:  The issue about the various phantoms was discussed, as the ACR is well aware, with Dr. Yaffe.  Dr. Yaffe is very familiar with that phantom.  I believe he developed it.  So, I just wanted to inform you that this decision was not totally arbitrary, and it has been decided, according to my understanding, and as Penny pointed out, that at this time, a decision was made to stay with the existing phantom.

I have no idea as to when we will be going to a new phantom.  Chances are at some point, something will change.

The other issue, though, that I think is of some importance, is to consider the minimum score required for these phantoms, that we cannot make an arbitrary decision, and frankly, I cannot tell you that if we increase the score one notch, something is going to change all of a sudden.

I find often on a phantom review, that some of my objections may not be as much with scoring, Pugh scoring of all the features in the phantom, but with the overall impression of the image, like excessive noise.  Although you see everything, you just don't like the noise and the image artifacts, that they are excessive and bothersome, and I would like to see something that addresses all these issues, giving a little more latitude for the reviewers to be a little more critical.

There is latitude right now, but perhaps to the point that we can actually reject something a little more easily in the future.

MS. HARVEY:  Another question?

CPT THOMAS:  My name is Jerry Thomas.  I am at the Uniformed Services University.  Kind of a comment and a question at the same time.

It has been clearly pointed out that there are substantial differences between quality assurance programs for each of the three approved digital systems.  Our current training requirements are eight hours for a new modality.

Do we have three different new modalities?  I think maybe eight hours, in my experience, and I ran this program, it was this past weekend that Melissa talked about, I think maybe eight hours is not enough training.

I would suggest you may want to consider what the impact of these new modalities are going to be on both the training requirements for the technologist, as well as the medical physicist.  I think probably the radiologist training could meet the eight-hour requirement without additional training, but probably not the other two.

I would like to hear other thoughts, as well.

MS. HARVEY:  Ms. Martin.

MS. MARTIN:  I guess my first response would be again coming from some of the other physicists, too, or their training, I think we have to look at it.  I guess I don't have a problem with the eight hours of the initial training.  The first time any of us have to go through one of these machines, you are going to have to go through it with the manufacturer's representative.

Certainly, eight hours of general training would not qualify you to walk in cold and do a different manufacturer's unit with absolutely no assistance, but eight hours of basic training in digital mammography imaging would qualify you to go with the engineer.

That is just an opinion, but I would agree with Captain Thomas, you are not going to be ready to do it without anyone around.

DR. PISANO:  I just have a comment about the technologists.  Does anybody who is a technologist want to comment about the technologist training?  Did you want to comment?  I really wanted a comment from a technologist if that is possible.

MS. ELLINGSON:  I am not working in the field myself, but we do have a lot of questions coming in to ASRT. They seem to think we know all the answers, but it is not very specific as to what eight hours of training is, is it applications training, is it a CE course where you heard a lecture on digital.

The questions that I am getting leads me to believe that it is not very clear what is intended for them to count as the eight hours of initial training.

MS. HARVEY:  Dr. Pisano.

DR. PISANO:  I just want to comment that in my experience, as I mentioned earlier, I have three different machines, and I find that the eight hours is more than adequate for the technologists to learn how to use the equipment and the tests that they are required to do.

I have probably been through this process with about, I would guesstimate 20 technologists at this point, because we have these three units, and we have a turnover in our place, so we have done it quite a bit.  I haven't found too big of a need for additional training ever, in fact, out of all those techs.

DR. FINDER:  Going back to the question about what types of training are involved, again, we have to go back to the history behind this.  At the time the regulations were written, these units didn't exist, so we tried to get the best opinions and expert advice that we could to try and settle on some type of initial training that was required, and we came up with the eight hours for the various personnel categories.

In our guidance, we have enumerated some of the things that can be used to meet this eight hours of training, and again, we were fairly flexible and general in the statement, so yes, hands-on experience can count in terms of training programs, CME courses, CEU courses would all, if they added up to the eight hours, would meet the requirement.

DR. PISANO:  I think the technologists are highly motivated also to learn how to use the equipment, and if they need more time, they are not shy about saying to the equipment manufacturer representative, who is present, that they need more time learning.

I haven't found it to be a real problem.  I can also say from the radiologist perspective, many radiologists have said to me--I actually run one of these programs for CME credits--and while it is a very nice way for our program to make some money, I have had many people say that they don't think eight hours is appropriate, that probably four would do, so maybe that could be shortened at some point for radiologists.  It is really not that different.

The main difference is reading on softcopy, so that is just another viewpoint on that.

MS. HARVEY:  Mr. Crocker.

MR. CROCKER:  This is Ken Crocker from Fischer Imaging again.

I just wanted to kind of reiterate that I think we need some urgency in developing some uniformity.  You know, if you look at some of the times lines here, GE had their PMA approval in December of 2000, and we are here today yet, and for one manufacturer and with one accrediting body, we don't yet have the process moved over to the accrediting bodies.

The other thing I would like to point out is that it appears that the approach that is being taken is very linear or sequential rather than in parallel.  I think as Penny mentioned, they are working very hard to work with one accrediting body, with one manufacturer, to get one approval.

If we continue that approach into the future, it is going to really tie a lot of people up and a lot of users up not having achieved their transfer over to what the regulation really intended.

That is just a comment.

MS. HARVEY:  Thank you.

Another question?

MS. MARTIN:  As a consulting physicist, I can only support Penny's comment and Ken Crocker's comment.  If there is one set of forms, one set of measurements that we are all expected to make, that is to the benefit of all the physicists.

MR. VASTAGH:  My name is Steven Vastagh.  I am with the NEMA, National Electrical Manufacturers Association.  I am pleased to recognize that it is wonderful that we have, not one, but three or four different solutions for digital mammography, so there are two sides for each issue, but I am pleased to tell you that NEMA and the manufacturers will begin to make an effort to harmonize QC tests.  I am real pleased to hear that the accreditation bodies are supportive of that and hope that this will contribute to speeding up the process.

MS. HARVEY:  Thank you.

Any further questions?  No?  Thank you.

Let's move on now to Dr. Burkhart, who will tell us a little bit about States as Certifiers.

States as Certification Agencies Update

DR. BURKHART:  I am going to give a brief update. The way I should start is to point out that these activities all originate from Subsection Q of the original Mammography Quality Standards Act of 1992.

Subsection Q permits FDA to authorize State agencies to carry out some of the functions under our oversight.  Perhaps most visible of these functions is the actual issuance of the certificates to the facilities, the certificates that they need to be able to do mammography, and because of the visibility of this particular function, that is why we refer to the whole effort as State's certifiers or we commonly use the acronym SAC to refer to these activities.

But it shouldn't be forgotten that this is not the only function that the States can be authorized to carry out.  Among other functions is administrative control of the inspection activities within their borders.  As I think probably everybody knows, the great bulk of the inspections are performed by State personnel under contract to FDA, under general FDA oversight and administrative control, but a SAC State can have the function of that general administrative control.

With this comes any associated follow-up actions to the inspections, any follow-up on the Level 1 or Level 2 citations can become the responsibility of the SAC State.

To go a step further, if compliance actions are necessary, these also can be a function which the State can be authorized to carry out although I should mention in connection with that, FDA still has the right, the authority to carry out compliance functions within a SAC State, as well as the State itself.

On the other hand, I should point out that perhaps a major function that can't be delegated to a SAC State is the function of developing the standards for the accreditation bodies or for the facilities.

This is specifically prohibited by law being delegated to a SAC State, and we define this as including not only the regulations, but also the guidance which interprets the regulations.  This remains an FDA function again the SAC States.

Before we could open up the possibility nationwide of States becoming SAC States, we needed to have implementing regulation, and to help us develop these regulations, about three years ago now, a SAC demonstration project was established, which the idea was that a limited number of States, for a limited period of time, would be given SAC functions to carry out.

They would be authorized to carry out the functions that I mentioned.  So, for about three years now, the States of Iowa and Illinois have been recognized as SAC States, and they have been carrying out the functions that I mentioned, and from this experience, we have gained a great deal of information which has been useful to us in developing regulations.

It has also been useful to use in our thinking about the long-term oversight activities, but now we are ready to move on to another plane because on February 6 of 2002, SAC regulations were published as final, and they became effective on May 7th.

So, now we have a third subpart to the MQSA regulations.  Subpart A is accreditation bodies.  Subpart B is the facilities.  Now we have Subpart C for the SAC States.

So, as new States enter the program, they will be looked at, their applications will be looked at under these new regulations, and the maintenance of activities also will be the oversight will be directed by the new regulations.

Probably one question which may come to your mind is are there other States that are interested in becoming in SAC States, and several States have mentioned some interest to us.  This interest is buried in inquiries in some cases, and in other cases, the States have gone further.

Probably the most time-consuming part of becoming a SAC State is in the development of regulations because it is required by the law that a SAC State have regulations in the mammography area equivalent to the MQSA regulations for the facilities.

On the State level, as on the national level, it takes time to develop regulations, so a State that is interested in becoming a SAC State, seriously interested, that is a logical first process to get started to begin developing their regulations.

It is also prudent if they are going to go this way, it is prudent for them to talk to us about their plans to begin with rather than go through the process, if the regulation is final, and then discover that they are not satisfactory and have to go through it again.

So, this is a prudent first step to discuss the regulation plans with us.  There have been States that have discussed regulations that they are working on with us, have discussed with us the regulations they are working on for this purpose.

But at the present time, we have no active applications in-house under review to produce SAC States, so at the present time, the only SAC States which exist are the two which we active under the demonstration project, the States of Iowa and Illinois are our current SAC States.

So, this brings us up to date to where we stand today.  The big news again since the last time this committee met was the publication of the regulations as final.  That has been the major change.

If there are any questions, I would be happy to try to answer them.

MS. HARVEY:  Any questions?  Dr. Pisano.

DR. PISANO:  In the guidance document, what about Arkansas, California, and Texas, that are listed in the guidance document?  I am a little confused maybe.

DR. BURKHART:  Arkansas, California, Texas, and Iowa, I think what you are referring to as accreditation bodies.  This is different from becoming a State's Certifier.

For a facility to become certified, as you know, it has to be accredited, and we can approve as accreditation bodies, we can approve States or private, nonprofit bodies, and we have the four State AB's plus ACR, of course, as accreditation body.

But this is the next step, issuing the certificates once a facility is accredited and making sure that they are inspected properly and, as I said, carrying out any compliance actions and follow-up which is necessary.

DR. PISANO:  So, Illinois is a certifying State, but not an accrediting State.

DR. BURKHART:  Right, Iowa is both.  California, Texas, and Arkansas are just accreditation bodies at the present time.

MS. HARVEY:  Thank you.

I think we have come to the last part of our meeting.  Dr. Finder.

DR. FINDER:  Just before we go and review the summary minutes, some issues again were brought up just the last few minutes, again about accreditation for FFDM, and I just wanted to clarify a few things.

One is FDA only can deal with what we get in-house.  The accreditation bodies obviously have to make their own decisions whether they are going to go ahead and accredit under FFDM or not.  That is their decision.  We can't force  anybody to do anything.

We certainly are willing to look at any comments that an accreditation body wants to submit to us if they want to apply for FFDM.  The same is true for alternative standards and some other aspects of the FFDM program.  Manufacturers are certainly free to submit materials to us if they believe that they are appropriate for us.

Just to go back to one of the earlier statements in terms of approved alternative standards, one manufacturer did come in to us for an alternative standard regarding, not the frequency, but the amount of time that a unit could be still used depending on the QC test that was failed.

Other manufacturers are certainly free to apply for the same thing.  That is their decision, and if they don't want to, facilities, if they want to, can also apply for an alternative standard.

We are certainly open to comments and suggestions and efforts by manufacturers and other entities with this process, we are certainly open to that, so the more, the merrier.

Review of Summary Minutes of August 2001

DR. FINDER:  Next, in terms of the review of the summary minutes, if anybody has any comments.

MS. HARVEY:  Are there any corrections or additions that any members of the committee found when they reread the summary minutes of our last meeting?

[No response.]

MS. HARVEY:  Very good.  Excellent.

Dr. Finder, do you want to talk to us a little bit about future meetings?

DR. FINDER:  Yes, but before I talk about future meetings, I do want to make mention of one fact.  Dr. Amy Lee has served on our committee, and her term is expiring in January of next year, so chances are we will not be having another meeting before her term expires, so we just want to thank her for all her efforts and hope that this has been an enjoyable experience.

We know that we have gained a lot from her insights into this area, and we thank her for her participation here.

MS. HARVEY:  Thank you, Dr. Lee.

Future Meetings

DR. FINDER:  As for future meetings, the plan probably is going to be to have a meeting sometime in the spring.  This is going to be somewhat based on what happens. As many of you are aware, the Mammography Quality Standards Act is in the process of reauthorization.  It terminates in October of 2002.  Hopefully, we will have some action by Congress to reauthorize the program for another five years.

When they reauthorize it, it is not uncommon--I shouldn't say uncommon--they reauthorized once and did put in specific items that required immediate attention from that reauthorization.

Depending on what is included in the reauthorization this time around, we may have to take some immediate actions to generate some new regulations depending on what they say.  So, the plan is to have a meeting sometime in the spring, and the topics may be dictated by what happens in the reauthorization process.

I would expect that if they do reauthorize and put in a few new items, we might be talking about a two-day meeting rather than a one-day meeting, so just to get everybody informed.

As for the exact timing, I will try and do the same thing that I did for this meeting, which is send out requests for days that are available from everybody and try and generate a suitable time that is applicable to everyone.

I will mention the fact that this was the first time that we tried to send out all the materials electronically.  It was quite an experience.  I got a lot of e-mails that were bounced back at me and a lot of comments about it, but I think we are going to work through that process and hopefully, this time around it will be smoother.

For those people who got their e-mails, but no attachments, I think it may be that your systems have recognized my name and are stripping off the attachments immediately.  I am on your spam list I guess.

MS. HARVEY:  Any other further comments, questions?

I wish you all a safe trip home and we will meet again.  The meeting is closed.

[Whereupon, at 4:10 p.m., the meeting was adjourned.]

- - -