NASA  NEPP

BOK

NDE of Electronic Parts and Packaging

David T. Mih

# JPL

### Background

The purpose of this BOK study is two fold.  First is to obtain and report available Nondestructive Evaluation (NDE) technologies for inspection of electronic parts and packaging; second is to survey the NDE equipment throughout all NASA centers that are applicable for this purpose.

### Introduction

For decades, visual inspection and cross-sectional Destructive Physical Analysis (DPA) were the mainstream of failure analyses in the electronic parts and packaging field.  However, they have become increasingly difficult as the packaging density improves and multi-layer assembly prevents a direct line of sight for optical viewing.

In addition, as the size of electronic packages decrease, deformation and stress by mechanical sectioning make reliable failure analysis difficult.  Moreover, due to the nature of NASA programs, projects usually do not have the luxury of performing large DPA samplings since the procedures are destructive to the hardware, and hence are not desirable, often because of part cost and limited availability.

The following two technologies will be covered by another BOK and hence are excluded in this study.

SAM		Scanning Acoustic Microscope

SLAM		Scanning Laser Acoustic Microscope

##### TECHNOLOGY EVALUATION REPORT

The use of flip chips and ball-grid arrays (BGAs), and dense PWBs have created a demand for inspection methods that do not require circuit access through rows of pins. Automated optical inspection (AOI) and metrology techniques such as laser mapping offer some help, but these techniques require that components and PWB features be visible.

The study has found that the mostly widely used NDE technique is the X-ray imaging method, including its several derivatives.  The other relatively popular technique is the infrared thermography.

### X-RAY

X-rays are attenuated exponentially by mass or atomic number of matter.  The detectability depends on the spatial resolution of the radiographic system and the image contrast obtained by the radiographic detector.  The spatial resolution is a function of the detector, which governs the signal to noise ratio (SNR); and the geometric resolution, which is related to the image magnification that is governed by source to object distance and object to detector distance.  The geometric unsharpness is governed by the x-ray source focal spot size, while image contrast is affected by several factors, such as attenuation coefficient of the material and anomaly defect, thickness of the anomaly in the direction of the x-ray beam and intensity of scattered radiation.

## Film X-ray

This method has been around for decades and is still the most commonly used method in the medical and dental fields.  The detector in this case is film and graininess determines the resolution of the images.  The film is then developed and a negative image is viewed.  This method still yields the best SNR and contrast.

However, the quality of the x-ray image and operator’s skill play a critical role with this method since the inspection involves manually adjusting contrast and brightness for different regions via adjustments of voltage and current.  This manual method plus the time for film development are not practical for most electronic inspections.

## Real Time Radiography

With the development of x-ray detectors and microfocus x-ray sources, a real time high-resolution inspection can be achieved.  The set-up is shown in Figure 1.  The small x-ray beam produced at the anode point source (4 um) compared with conventional beam dimensions allows magnification of up to 1000x without loss of image sharpness.  A detector such as an image intensifier coupled to a video camera provides an instantaneous image on a video monitor.

Figure 1.  Real Time High Resolution X-ray Imaging System

### Detector Technology

There are two common types of detectors used in today’s real time x-ray imaging systems, analog and digital.  The image intensifier is the most popular analog detector because it’s relatively inexpensive, easy to maintain and provides good contrast since it’s essentially a vacuum tube.  The intensifier tube is convex and covered by a thin layer of cesium iodide forming the cathode of the detector, this layer converts the incident x-ray into visible light.  Other analog detectors include photo-multipliers and glass-scintillaters coupled with vidicon cameras.

#### The Flat Panel Detector (FPD), Figure 2, has gained popularity recently, this type of detector uses amorphous Si or Se films as radiation sensitive layers. The active imaging areas (11” x 17”) are much larger than the image intensifier (6” Dia.), and the thicknesses of the devices are fairly thin to be comparable to the conventional x-ray film cassettes.  The spatial resolutions of the devices are similar to conventional image intensifiers, but the noise levels are low. Therefore, high image qualities can be achieved with the new FPD systems. They have been developed so far mostly for medical purposes, but recently have been applied to NDE applications.

Figure 2.  Amorphous Silicon Flat Panel Detector

### Source Technology

In order to view defects in small objects such as solder balls or bond wires with diameters as small as 25 microns, magnification and resolution are significant factors. High magnification is almost a prerequisite for the inspections of microelectronic components, but big image with low resolution offers little or no value.  The image resolution depends mostly on the dimensions of the x-ray source, or "focal spot".  Focal spot size of an x-ray source determines the sharpness of the image around the perimeters, as shown in Figure 3.  For the majority of x-ray systems currently in use in the electronics industry, focal spot size ranges up to 8 microns, sometimes described as “microfocus”. A typical microfocus x-ray tube is illustrated in Figure 4.  In the past two years, a few x-ray vendors have introduced x-ray source of less than 1 micron and described as “Nanofocus”.

###### Figure 3.  Magnification vs. Edge Distortion and Source Size

Figure 4.  Microfocus X-ray Source (Courtesy of FeinFocus)

### 3-D X-ray Systems

In two-dimensional (2-D) x-ray inspection, a source projects downward normal to the component’s surface, therefore it transmits through both sides and sees all the layers at the same time. Using such a simple irradiation technique, it is difficult to obtain information about the depth of the imaged structures.  Three dimensional (3-D) techniques such as tomography, laminography and tomosynthesis examine horizontal slices of the electronic parts or packaging at desired z-axis heights, and can therefore analyze several horizontal layers of a component.

Computed Tomography

X-ray computed tomography (CT) (Figure 5) has been widely used in the medical field as well as for NDE in many industrial applications.  The combination of X-ray transmission technique with tomographical reconstruction allows three-dimensional information about the internal microstructure. In this case any internal area can be reconstructed as a set of flat cross sections, which can be used to analyze the two- and three-dimensional morphological parameters.  For X-ray methods, the contrast in the images is a mixed combination of density and compositional information.

Figure 5.  Computed Tomography

### Laminography

Because of high absorption and limited access, CT is not always possible.  For example in the case of flat components such as multilayer printed circuit boards, a reconstruction algorithm may not work due to beam hardening effects.

Laminographic methods are able to overcome some of these difficulties.  They can be used to visualize the internal details of thin multilayer structures such as ball grid arrays and printed circuit boards.  Microlaminography allows layer by layer information to be measured in 'slices' as small as 10 microns. Questionable solder joints, disbonds and voids can be detected.  These methods yield images of object slices and allow the determination of the position of the object structures.  Linear translation of the object through the fan beam of an x-ray source is utilized.  Both the x-ray source and the detector remain stationary, or alternatively the object can remain stationary and the x-ray source and the detector can be moved synchronously but without relative movement as shown in Figure 6.  During movement, the object is irradiated by the x-ray at different angles due to the fan beam with a wide angle.  Thus, the elements of the detector successively obtain information of a given volume element of the object under consecutively changing angles and these digital projections contain the complete structure information of all object slices. The projection values then will be sorted and added to form a digital image.

###### Figure 6.  Schematic of a Laminographic System

Computed Tomosynthesis

Another variation of laminography is based on relative motion of the x-ray source, the detector and the object. The x-ray source and the detector are either moved synchronously in circles or are simply translated in opposite directions.  Due to the correlated motion, the location of the projected images of points within the object moves also.  Only points from a particular slice, sometimes called the focal slice, will be projected at the same location onto the detector and the image will be sharp.  Objects above and below the focal slice will be projected to different locations, and the resultant image will be fuzzy and will be superimposed as background intensity to the focal slice. This principle of superimposing projections is called tomosynthesis.

### Infrared Thermographic Inspection

Infrared Thermography or Radiometry, is a nondestructive and non-contact method for observing and measuring thermal energy from materials in all phases; solid, liquid and gas.  To observe gases and thin films which appear transparent to short-wave (3-5µm) or medium wave (8-12µm) IR camera’s, special filters can be used to selectively narrow the incoming radiation at a specified waveband of interest. For instance; carbon dioxide, an atmospheric gas, releases IR radiation in the 9-10.5µm wavelength, so a filter that attenuates the electromagnetic radiation below and above that waveband will allow the gas emissions to effectively be seen by the IR camera.  However; Silicon, Silicon Nitride and Gallium Arsenide, semiconductor substrate materials, appear transparent in the waveband of ~1µm to 15µm, a waveband were most infrared detectors are designed, so heat is virtually non detectable from these materials using a conventional IR camera.  One solution is to use a near IR (0.8 to 1.2µm) camera to measure radiation from these types of materials.  An infrared (IR) thermographic inspection system (Figure 7) can be used to obtain hot spots in the printed wiring board where local heating could be caused by electrical shorts.

Figure 7.  Thermographic Imaging System

Case Studies and Application Illustrations

X Ray

In the electronics industry, inspection systems using x-rays have proven particularly adept at examining hidden solder connections in PWBs, and provide a practical way to inspect the connections between ball-grid arrays (BGA) and a PWB, as well as other defects.

Open connections.  Open connections or cracks appear as light areas where a good connection would show solid dark images.  (Figure 8)

Figure 8.  Open Connection

Solder bridges and voids.  Solder bridges between contacts or short circuits usually result from excess solder or misplaced solder.  Bridges also can occur after rework of a BGA board, because of the difficulty of re-screening solder paste on the reworked area.  Bridges show up in an x-ray image as dark connections, while voids appear as light spots within the dark ball area in the x-ray image (Figure 9).  Entrapment of voids can result from improper re-flow temperature that prevents flux from escaping or from matter picked up by the solder paste from an improperly cleaned PCB.

Figure  9.  Bridges and Voids

Misregistration.  Misregistration occurs when PWB pads, solder balls, solder paste, or the PCB solder mask are not properly aligned, or “registered”.  An x-ray image of misregistered components (Figure 10) will show slight offset from PWB contacts.

Figure  10.  Misregistration or Alignment Offset

Summary

X-ray inspection is not a single technique but a family of techniques, each of which will perform more or less well in particular situations. Achieving success among the x-ray alternatives depends on a PCB's topology, technology, application, and other factors.

X-ray imaging falls into two broad categories: 2-D and 3-D. Two-dimensional (2-D), or transmission, x-ray inspection involves a source shining down perpendicular to the board's surface, thereby "seeing" both sides at the same time. Three-dimensional (3-D) techniques such as tomography, laminography and tomosynthesis examine horizontal slices of the board at desired z-axis heights, and can therefore analyze several horizontal layers of a PCB in a single inspection pass.

CT allows nondestructive imaging of object slices, but with the restriction that the objects have to be irradiated from the full angular region.

Laminography can inspect in adjustable layer-by-layer steps from 10um to 2.5 mm. A maximum of 2.56 mm /256 layers can be scanned in one cycle, with maximum specimen thickness of 10 mm that can be visualized by several scans.

2-D x-ray offers shorter inspection times than 3-D because 3-D needs to construct multiple images—one per "slice" through a PCB.  Therefore 2-D is recommended for real-time analysis in time critical situations.

In the case of thermal inspection for microelectronics, software is the key issue.  As most of the thermal applications rely on the detection of anomalies in the expected temperature pattern, there might be effects of artifacts (e.g. reflections from the environment, air turbulences, inhomogeneous infrared emission coefficient) that give rise to false alarms or, even worse, hide real and critical anomalies.  Unlike structural thermal inspection , most of the contrast in an unprocessed infrared image of a circuit is not due to a temperature differential; rather the differences in the emissivity create the most contrast.  A good software program with the proper algorithms must be used to measure and compensate for the emissivity of the material being tested in order to generate an accurate, calibrated thermal image.  Otherwise misleading material thermal signal will result.  Custom made focusing objective lenses are also required since semiconductor materials and related circuits emit low contrast thermal signals in addition to their small sizes.

Conclusion

As with all test and inspection issues, no ideal solution exists.  Neither 2-D nor 3-D x-ray techniques can address all inspection needs.  Solder opens on ball-grid arrays caused by lack of coplanarity on the board pad are visible only by 3D systems via a board slice, because on 2-D systems the open is hidden by the shadow of the solder ball itself.  3-D x-ray systems can quantitatively measure the volume of the solder joint, determining heel length, fillet height, and other characteristics.  2-D x-ray offers shorter inspection times than 3-D because 3-D methods take multiple slices or views and then reconstruct the images.

NASA NDE EQUIPMENT LIST

(For Microelectronic Inspection Only)

| Equipment Type | Center | Contact Information |
| --- | --- | --- |
| Micro-focus X-ray Imager 1 | JPL | David Mih (818) 354-9218 david.t.mih@jpl.nasa.gov |
| Infrared Microscope 2 | JPL | Ron Ruiz (818) 354-2184 Ronald.P.Ruiz@jpl.nasa.gov |
| Digital Radiography/Laminography | GRC | Richard Rauser (216) 433-5018 Don Roth (216) 433-6016 |
| Inner View X-ray System | SSC | William St. Cyr ((228) 688-1134 William.W.St_Cyr@nasa.gov |

- Practically every NASA center has a film x-ray system.  LaRC has a CT system, WSTF has a real time x-ray imager but with a regular sized source.  However, none of them are suitable for electronic inspection either due to a large x-ray source or large tomographic slices, since they are designed for large structural inspections.  All of the above lack the resolution required for microelectronic inspection.

2	MSFC, KSC and LaRC possess thermal imagers; but again, they are designed for large composite structure inspection rather than for microelectronics application.

Manufacturers of X-ray Equipment

Agilent Technologies  http://we.home.agilent.com/USeng/nav/-536887030.0/pc.html

Dage    http://www.dageinc.com/

Glenbrook Technologies   http://www.glenbrooktech.com

Faxitron  Wheeling  http://www.faxitron.com

Feinfocus  http://www.feinfocus.com/

Lixi, Downers Grove, IL   http://www.lixi.com/

Nicolet Imaging Systems    http://www.nicimg.com/

Operations Technology, Incorporated   http://www.optek.net/

Phoenix X-ray  http://www.phoenix-xray.com/

Photon Dynamics, Electronics Division, Aliso Viejo, CA  92656Formerly CR Technology

SkyScan/MicroPhotonics   http://www.skyscan.be/

Teradyne, Inc.  http://www.teradyne.com/axi/xstation.html

Viscom   http://www.viscom.de/

VJ Electronix   http://www.vjelectronix.com/vjebeta/index.asp

X-Tek Systems  http://www.xtek.co.uk/