November 21, 1997

Contract # AQ-1486A

***** Paragraph Numbered Format ***** ******* Copyright 1988, ARCOM ******* ARCHITECTURAL COMPUTER SERVICES, INC.

SPECIAL SPECIFICATION

SECTION 11610-S

LABORATORY FUME HOODS

Page

PART 1 - GENERAL

1.01	Summary	3

1.02	Design Requirements	3

1.03	Submittals	4

1.04	Quality Assurance	5

1.05	Delivery, Storage, and Handling	5

1.06	Project Conditions	5

1.07	Sequencing and Scheduling	6

1.08	Warranty	6

PART 2 - PRODUCTS

2.01	Manufacturers	6

2.02	Materials 	7

2.03	Benchtop Chemical Fume Hoods	8

2.04	Isotope Hoods	14

2.06	Walk-in Hoods	24

2.07	Canopy Hoods	24

2.08	Fume Hood Metal Finish	24

2.09	Source Quality Control Testing of Epoxy Resin Work Surface	25

2.10	Source Quality Control Testing of Poly Resin Liner	26

2.11	Source Quality Control Testing of Fume Hoods	29

2.12	Source Quality Control Testing of Fume Hoods	29

2.13	Source Testing of Fume Hood Sound Levels	29

PART 3 - EXECUTION

3.01	Installation	30

3.02	Field Quality Control Testing of Fume Hoods	30

3.03	Field Quality Control Testing of Fume Hoods	31

3.04	Adjusting	31

3.05	Cleaning	32

3.06	Protection of Finished Work	32

SPECIAL SPECIFICATION

SECTION 11610-S

LABORATORY FUME HOODS

THIS SECTION USES THE TERM "ARCHITECT."  CHANGE THIS TERM AS NECESSARY TO MATCH THE ACTUAL TERM USED TO IDENTIFY DESIGN PROFESSIONAL AS DEFINED IN THE GENERAL AND SUPPLEMENTARY CONDITIONS.

PART 1 - GENERAL

1.01	SUMMARY

A.	Section Includes:

1.	Laboratory fume hoods

B.	Related Sections:

Section 01700-S Commissioning.

Section 11601:  Laboratory Casework and Furnishings

3.	Division 15:  Furnishing and installation of plumbing utilities and final connections to fume hoods.

4.	Division 15:  Furnishing and installation of exhaust duct work and equipment, and final connection of fume hoods.

5.	Division 15:  Furnishing and installation of exhaust controls and final connections to hoods.

6.	Division 16:  Furnishing and installation of electrical utilities and final connections to fume hoods.

1.02	DESIGN REQUIREMENTS

A.	Fume hoods shall function as ventilated, enclosed work-spaces, designed to capture, confine and exhaust fumes, vapors and particulate matter produced or generated within the enclosure.

B.	Design fume hoods for consistent and safe air flow through the hood face. Negative variations of face velocity shall not exceed 10 percent of the average face velocity at any designated measuring point as defined in this section.

C.	Average illumination of work area:  Minimum 80 footcandles. Work area shall be defined as the area inside the superstructure from side to side and from face of baffle to the inside face of the sash, and from the working surface to a height of 28 inches.

D.	Fume hood shall be designed to minimize static pressure loss with adequate slot area and bell shaped exhaust collar configuration. Average static pressure loss readings taken 3 diameters above the hood outlet from 4 points, 90 degrees apart, shall not exceed the following maximums:

Face Velocity		Measured S.P.L. (W.G.)

75 FPM			0.18 inches

100 FPM			0.30 inches

125 FPM			0.45 inches

E.	Fume hoods shall be constructed in such a manner as to cause no perceptible increase in sound pressure level over the mechanical exhaust system.

F.	Maintenance Manuals:  Provide one written instruction manual for each hood and two additional for each type of hood outlining operating/safety instructions and proper maintenance procedures.

1.03	SUBMITTALS and SUBSTITUTIONS:

A.	Product Data:  Submit manufacturer's data for each type of hood specified. Include component dimensions, configurations, construction details, joint details, and attachments. Indicate location, size, and service requirement for each utility connection.

B.	Shop Drawings:

1.	Provide 3/4 inch = 1'-0" scale elevations of individual and battery of hoods showing cross sections, rough-in and anchor placements, tolerances, and clearances. Indicate relation to other laboratory equipment, surrounding walls, windows, doors, and other building components.

2.	Provide 1/4 inch = 1'-0" rough-in plan drawings for coordination with trades.

3.	Provide 1 set of sepia reproducible prints and 1 set of blueline prints.

C.	Top Samples:  Submit product sample of each type of benchtop.

D.	Finish Samples:  Submit 3 inch by 5 inch samples of color of finish for fume hoods, work surfaces and for other prefinished equipment and accessories for selection by the Owner's Representative.

E.	Test Reports:  Submit test reports verifying conformance to specified performance tests.

F.	Life Cycle Tests:  Provide independent test data for sash hardware components such as cables and pulleys.

G.	Maintenance Manuals:  Provide written instruction manuals outlining operating/safety instructions and proper maintenance procedures.

H.	Submit the above in accordance with section 1300-S.

I.	Substitutions on listed manufacturers in this specification shall be performed per the requirements of section 1600-S.

1.04	QUALITY ASSURANCE

A.	Single Source Responsibility:  Fume hoods, casework, work surfaces, laboratory furnishings, and accessories shall be furnished by a single laboratory furniture company.

B.	Manufacturer's Qualifications:  Modern plant with proper tools, dies, fixtures, and skilled production staff to produce high quality laboratory casework and equipment, and shall meet the following minimum requirements:

1.	Five years or more experience in manufacture of laboratory casework and equipment of type specified.

2.	Ten installations of equal or larger size and requirements.

C.	Installer's Qualifications:  Factory trained and certified by the manufacturer.

1.05	DELIVERY, STORAGE, AND HANDLING

A.	Delivery:  Schedule delivery of fume hoods and equipment so that spaces are sufficiently complete that material can be installed immediately following delivery.

B.	Handling:  Protect finished surfaces from soiling or damage during handling and installation. Keep covered with polyethylene film or other protective coating. Protect all work surfaces throughout construction period with 1/4 inch corrugated cardboard completely covering the top and securely taped to edges.

1.06	PROJECT CONDITIONS

A.	Do not deliver or install equipment until the following conditions have been met:

1.	Building areas requiring the installation of laboratory casework shall be dry and unexposed to adverse weather conditions which may damage finished materials.

2.	Interior building temperatures shall not register below 65 degrees Fahrenheit in areas of casework installation to permit the proper curing of epoxy sealants and adhesives.

1.07	SEQUENCING AND SCHEDULING

A.	All overhead mechanical, electrical and plumbing rough-in work shall be complete prior to laboratory casework deliveries.

B.	All mechanical, electrical and plumbing rough-in work required along walls and service islands, where lab furnishings are to be installed, is to be complete prior to delivery of materials.

C.	Walls and partitions must be in place and finished with at least the primer coat of paint. If finish painting is to take place after lab furnishing, installation, protect the casework and furnishings by covering and masking prior to commencement.

D.	All necessary wood or metal blocking must be installed within partitions prior to delivery of casework and furnishings.

E.	Overhead soffits and ceiling grid must be in place prior to casework installation.

F.	Overhead lighting must be installed and connected prior to casework installation.

G.	All flooring required to be placed under lab casework and furnishings must be installed prior to material delivery.

H.	Concrete floors must be level within 1/8 inch of level per 10 foot run, nonaccumulative, when tested with a straight edge in any one direction.

I.	Wet operations to be performed must be complete prior to material deliveries.

1.08	WARRANTY

A.	Provide a 1-year warranty against defects in materials and workmanship.

PART 2  - PRODUCTS

2.01	MANUFACTURERS

A.	Subject to compliance with requirements, provide products by one of the following:

1.	Fisher Hamilton Scientific Inc.

1316 18th Street

Two Rivers, Wisconsin  54241

2.	Kewaunee Scientific Corporation

Laboratory Division

PO Box 5400

Statesville, North Carolina  28677

3.	Ly-Line Airguard

P.O. Box 38

2251 Cole Street

Enumclaw, Washington  98022

Air Master Corp.

6480 Norton Center Dr.

Muskegon, MI 49441

5.	Substitutions: only under provisions of section 1600-S.

B.	Casework, furnishings and equipment specified in the following sections shall be furnished and installed by the same supplier:

1.	Section 11601:  Laboratory Casework and Furnishings

2.	Section 11605:  C-Frame Laboratory Furniture System

2.02	MATERIALS

A.	Sheet Steel:

1.	Metallic Furniture Stock:  ASTM A-366 mild steel, cold-rolled, pickled, double annealed, and free from rust, scales, scratches, buckles, ragged edges, and other defects.

2.	Minimum Gages:

a.	All gages indicated shall be US Standard.

b.	18 gage thick exterior casing, ceiling closure panels, and bottom sash rail.

B.	Stainless Steel:

1.	Type 304:  ASTM A240; 16 gage; exposed surfaces ground and polished to a Number 4 finish.

2.	Type 316:  ASTM A240; 16 gage; exposed surfaces ground and polished to a Number 4 finish.

3.	Welding:  All stainless steel welding material shall be of type similar to sheet material. Welds shall be made without discoloration; ground, polished, and passivated to blend harmoniously with a Number 4 satin finish.

C.	Epoxy Resin Sheets:

1.	Molded from modified epoxy resin that has been especially compounded and cured to provide optimum physical and chemical resistance required of a heavy duty laboratory working surface. Uniform mixture throughout, not dependent on a surface coating.

2.	Physical Properties:

a.	Compressive Strength (ASTM D695):  30,600 psi

b.	Flexural Strength (ASTM D790):  12,800 psi

c.	Tensile Strength (ASTM D638):  10,100 psi

d.	Heat Distortion (ASTM D648):  330 degrees Fahrenheit

e.	Water Absorption (ASTM D570):  0.018 percent

f.	Fire Resistance (ASTM D635):  Self Extinguishing

D.	Glass:  7/32 inch clear laminated safety glass.

2.03	BENCHTOP CHEMICAL FUME HOODS

A.	Superstructure:

1.	Wall:  Rigid, self supporting assembly of double wall construction, maximum 5 inches thick. Double wall shall consist of a sheet steel outer shell and a corrosion resistant inner liner and shall house and conceal steel framing members, attaching brackets and remote operating service fixture mechanisms and services. Panels must be attached to a full frame construction, minimum 14 gage thick galvanized members. Panels and brackets attached to eliminate screw heads and metallic bracketry from hood interior. Hoods shall be completely factory assembled to form a rigid, self-supporting structure.

2.	Exterior bypass grille:  Low resistance type, 18 gage steel, upward directional louvers.

3.	Interior Bypass:

a.	Provide an interior bypass plate on all variable volume or constant volume hoods with combination sashes and on all variable volume hoods with a vertical rising sash.

b.	The bypass area shall remain open on all constant volume hoods with a vertical rising sash.

4.	Bypass plate:  16 gage thick with 1/4 inch diameter holes; 1/2 inch on center in a staggered pattern; 23 percent open area. Provide with a urethane powder coat finish to match color of hood liner.

5.	Ceiling closure panel:  Minimum 18 gage thick, finish to match hood. Provide with hinged panel for access to hood lights, piping and exhaust duct connections.

6.	Fastenings:

a.	Exterior structural members attachments:  Sheet metal screws, zinc plated.

b.	Interior fastening devices concealed. Exposed screws not acceptable.

c.	Exterior panel member fastening devices concealed. Exposed screws not acceptable.

B.	Face Opening:

1.	Air foil or streamlined shape with all right angle corners radiused or angled:

2.	Air foil:  Bottom horizontal foil shall provide nominal 1 inch bypass when sash is in the closed position. Bottom foil shall be removable without use of special tools. Foil shall provide access areas for electrical cords. Bottom foil to be constructed of Type 304 stainless steel, with an acid and abrasion resistant paint coating applied by an electrodeposition process. Color to be selected by the Owner's Representative.

3.	Flush sill air foil:  Lower front assembly shall consist of a trough with hinged louvered horizontal panels forming a flush sill.

a.	Trough to be 16 gage thick, Type 304 stainless steel with an acid and abrasion resistant paint coating applied by an electrodeposition process. Provide 1 inch square 1/16 inch thick rail secured to fume hood end panels for reinforcement. Finish same as trough.

b.	Slotted angle horizontal panels, 16 gage thick stainless steel, Type 304, with linear polyurethane finish. Spot weld stand-offs to reinforcement rail and bolt at each end of horizontal panel to stand-offs.

c.	Panel slot patterns and size as follows:

1)	48 inch wide fume hood:  (2) 19 inch long panels

2)	72 inch wide fume hood:  (2) 31 inch long panels.

3)	96 inch wide fume hood:  (3) 28-1/2 inch long panels.

d.	Color selection of finishes by Owner's Representative.

C.	Sash:

1.	Full view type with clear, unobstructed, side to side view of fume hood interior and service fixture connections.

2.	Vertical rising sash:  2 inches maximum height bottom sash rail, 18 gage thick steel with a polyurethane powder coat finish. Provide integral formed flush pull the full width of bottom rail. Set safety glass into rails in deep form, extruded polyvinyl chloride glazing channels. Counter balance system:  Single weight, pulley, cable, counter balance system shall prevent sash tilting and permit one finger operation at any point along full width pull. Maximum 7 pounds pull required to raise or lower sash throughout its full length of travel. Design system to hold sash at any position without creep and prevent sash drop in the event of cable failure. Sash and counterbalance mechanism to be life cycle tested to withstand a minimum of 100,000 cycles without signs of fatigue. Submit independent test data. Open and close sash against rubber bumper stops.

3.	Horizontal sliding sash:  Bottom H-shaped rail to be 18 gage thick stainless steel with stainless steel roller wheels set on 18 gage thick stainless steel track. Top track to be extruded polyvinyl chloride set in 18 gage thick stainless steel U-channel assembly. All stainless steel components finished with polyurethane finish. Color selection by the Owner's Representative.

4.	Sash glass:  7/32 inch thick laminated safety glass.

5.	Sash cables:  Stainless steel, 1/8 inch diameter.

6.	Sash guides:  Corrosion resistant polyvinyl chloride.

7.	Pulley assembly for sash cable:  2 inches diameter, zinc dichromate finish, ball bearing type, with cable retaining device.

8.	Sash pull:  Steel with chemical resistant powder coating. Color selection by the Owner's Representative.

D.	Liners:

1.	Polyester liner:  Reinforced glass polyester panel; smooth finish and white color in final appearance. Flexural strength:  7,450 psi. Flame spread:  15 or less per UL 723 and ASTM E84-80. Manufactured by Western Slate Company, "Westliner LC 210". Provide gasketed access panels in the liner for access to valves and fittings. Panels shall be of the same material as the liner and be provided with a PVC gasket to eliminate air leakage and retain liquids inside the hood.

E.	Baffles:

1.	Provide baffles to control air vectors into and through the fume hood. Fabricate of the same material as the liner. All baffle supports/brackets to be nonmetallic.

2.	Baffle designs which permit close-off of all slots are not acceptable.

3.	Must comply with OSHA Lab Standard Guidelines.

4.	Provide acid resistant label indicating proper baffle operation. Locate label on sidewall of hood interior.

5.	Provide for baffle adjustment to allow repositioning prior to hood operation permitting setting for (1) high thermal loading, (2) heavier than air gases or fumes generated near work surface, and (3) normal or average operation.

6.	Provide with remote baffle adjustment. Toggle style, one handed, single point control accomplished while hood is in use, without disturbing apparatus, from outside right hand corner post of fume hood with sash in either the open or closed position, and permitting setting for (1) high thermal loading, (2) heavier than air gases or fumes generated near work surface, and (3) normal or average operation.

F.	Work Surfaces:

1.	Molded epoxy resin:  For hoods with Poly resin liners. Color to be selected by the Owner's Representative from the manufacturers standard black, gray or white colors. Work surface as manufactured by The Durcon Company, Inc.; Prime Industries; or Laboratory Tops, Inc. 1-1/4 inches thick surface, dished 1/4 inch to contain spills.

G.	Sinks

1.	Cup Sinks:  Epoxy resin

a.	Manufacturer:  The Durcon Company; Prime Industries; or Laboratory Tops, Inc.

b.	Size:  4-3/8 inches by 13-3/4 inches by 5-7/16 inches deep with a minimum wall thickness of 3/8 inch.

c.	Screen:  Type 316 stainless steel, 0.028 wire diameter, 8 by 8 mesh per square inch.

d.	Provide swivel sink strainer adapter, Town & Country Plastics Model PP-18; R&G Sloane Part Number 7218; or Scientific Plastics Company, Inc., Part Number 31595-158.

2.	Cup Sinks:  Stainless steel

a.	Size:  4-3/8 inches by 13-3/4 inches by 5-7/16 inches deep.

b.	Constructed of Type 304 stainless steel, minimum thickness 18 gage thick, Number 4 mil. polish finish. All joints butt welded and ground smooth. Pitch bottoms to drain outlet.

c.	Screen:  Type 316 stainless steel, 0.028 wire diameter, 8 by 8 mesh per square inch.

d.	Provide swivel sink strainer adapter, Town & Country Plastics Model PP-18; R&G Sloane Part Number 7218; Scientific Plastics Company, Inc., Part Number 31595-158.

3.	Lab Sinks:  Epoxy resin

a.	Manufacturers:  The Durcon Company; Prime Industries; or Laboratory Tops, Inc.

b.	Minimum wall thickness of 1/2 inch with all interior corners coved to 1-1/2 inch radius and bottoms pitched to outlet opening. Provide with 1-1/2 inch outlet and beehive overflow. Overflow to be 2 inches shorter than depth of sink.

c.	Provide swivel sink strainer adapter, Town & Country Plastics Model PP-18; R&G Sloane Part Number 7218; or Scientific Plastics Company, Inc., Part Number 31595-158.

H.	Lighting:

1.	Fixture:  2 T-8 lamps, electronic ballast, UL listed fluorescent light fixture with sound rated ballast installed on exterior of fume hood roof. Provide safety glass panel cemented and sealed to the hood roof.

2.	Interior of fixture:  White, high reflecting plastic enamel.

3.	Size:  Largest possible up to 48 inches for hoods with superstructures up to 6 feet. Provide (2) 36 inch fixtures for hoods with 8 foot superstructures.

4.	Include lamps with fixtures.

5.	Illumination:  80 footcandle minimum.

6.	Lamps and ballast to be the same type as specified in Division 16 for overhead lighting in laboratory spaces.

I.	Electrical Power:

1.	Three wire grounding type receptacles, as indicated on the Drawings, rated at 120 VAC at 20 amperes.

2.	Receptacles to be gray in color with stainless steel flush plates.

3.	Factory install receptacles, lighting, electrical fittings and wiring in accordance with applicable codes and Division 16 specifications.

4.	Provide a dedicated receptacle on top of hood for safety monitor and alarms.

J.	Service Fittings and Piping:

1.	Hoods shall be factory prepiped to a point of connection 2 inches above the hood roof. Piping materials and installation for each type of service shall be as specified under Division 15. Pressure test all prepiped lines in the factory.

2.	Service fittings within fume hoods shall have color coded, acid and solvent resistant plastic coating, applied over fine sandblasted surface, properly cleaned. Surfaces to be coated shall be sprayed and baked three times with minimum coating thickness of 6 mil.

3.	Control Valves:  Control valves shall be mounted on the front panel of the fume hood, with all components subject to wear accessible from the exterior of the hood. The valve stem shall be for fume hoods with airfoil 45 degrees to the side wall of the fume hood. The centerline of the valve inlet and outlet shall be parallel and 1-1/8 inches apart. Valves shall have a threaded collar to hold the valve in place.

4.	Color coding of hood interior service fittings and remote control valves shall be as specified for Laboratory Service Fittings in the Laboratory Casework and Furnishings Section of this Specification.

K.	Exhaust Outlet:

1.	Rectangular with ends radiused, shaped and flanged, 18 gage thick, steel. Finish with epoxy powder coating. Color to match liner.

2.	Provide exhaust duct transition from the hood rectangular exhaust outlet to the building exhaust system round duct connection. Fabricate transition of the same material as the exhaust outlet. Epoxy powder coat inside and outside.

L.	Safety Monitor/Alarm System:

1.	Each fume hood shall be equipped with a velocity control and safety alarm unit which is to be Sandia Furnished Equipment and installed and connected by contractor. Provide factory cut-outs in the front of the hood for field mounting of the unit. Coordinate location and size of cut-out for installation of Landis & Staefa keypad unit 537-720A.  Connection of control and safety alarm unit to sensors and control valves shall be under Division 13.

M.	Safety Monitor/Alarm Systems:

1.	Provide fume hood monitor with audible and visual alarm signals to activate if face velocity drops below 80 fpm or rises above 180 fpm. Alarm setpoints to be field adjustable.

2.	Audible alarm signal tone shall be a minimum of 85 decibels at 10 centimeters. Visual alarm signal shall be a flashing or continuous glowing red light. Alarm design shall permit silencing of tone signal, but red visual signal light shall glow until unsafe condition is corrected. Provide with circuitry and a pushbutton to test alarm circuits.

3.   The Fume Hood Control System shall be able to interface with the EMCS system called out in specification 13000.

4.	Provide with input/output accessory to send alarm signal to a remote location and for silencing alarm remotely.

5.	Electrical:  9 volt DC, maximum of 1 amp. Provide 120 Vac. Shall be provided by division 16.

N.	Instruction Plate:

1.	Permanent 4 inch by 6 inch plate attached to fume hood exterior with condensed information covering recommended locations for apparatus and accessories, baffle settings, use of sash, and recommended safe operating procedures.

2.	Provide a 4 inch by 6 inch corrosion resistant metal or plastic frame attached to front face of fume hood exterior with plastic or glass glazing to contain a removable data card to identify hood and show hood characteristics and field test data.

2.04	ISOTOPE HOODS

A.	Superstructure:

1.	Wall:  Rigid, self supporting assembly of double wall construction, maximum 5 inches thick. Double wall shall consist of a sheet steel outer shell and a stainless steel inner liner and shall house and conceal steel framing members, attaching brackets and remote operating service fixture mechanisms and services. Hoods shall be completely factory assembled to form a rigid, self-supporting structure.

2.	Exterior bypass grille:  Low resistance type, 18 gage steel, upward directional louvers.

3.	Interior Bypass:

a.	Provide an interior bypass plate on all variable volume or constant volume hoods with combination sashes and on all variable volume hoods with a vertical rising sash.

b.	The bypass area shall remain open on all constant volume hoods with a vertical rising sash.

4.	Bypass plate:  16 gage thick with 1/4 inch diameter holes; 1/2 inch on center in a staggered pattern; 23 percent open area. Provide with a urethane powder coat finish to match color of hood liner.

5.	Ceiling closure panel:  Minimum 18 gage thick, finish to match hood. Provide with hinged panel for access to hood lights, piping and exhaust duct connections.

6.	Fastenings:

a.	Exterior structural members attachments:  Sheet metal screws, zinc plated.

b.	Interior fastening devices concealed. Exposed screws not acceptable.

c.	Exterior panel member fastening devices concealed. Exposed screws not acceptable.

B.	Face Opening:

1.	Air foil or streamlined shape with all right angle corners radiused or angled.

2.	Air foil:  Bottom horizontal foil shall provide nominal 1 inch bypass when sash is in the closed position. Bottom foil shall be removable without use of special tools. Foil shall provide access areas for electrical cords. Bottom foil to be Type 304 stainless steel.

C.	Sash:

1.	Full view type with clear, unobstructed, side to side view of fume hood interior and service fixture connections.

2.	Vertical rising sash:  2 inches maximum height bottom sash rail, 18 gage thick, Type 304 stainless steel. Provide integral formed flush pull the full width of bottom rail. Set safety glass into rails in deep form, extruded polyvinyl chloride glazing channels. Counter balance system:  Single weight, pulley, cable, counter balance system shall prevent sash tilting and permit one finger operation at any point along full width pull. Maximum 7 pounds pull required to raise or lower sash throughout its full length of travel. Design system to hold sash at any position without creep and prevent sash drop in the event of cable failure. Sash and counterbalance mechanism to be life cycle tested to withstand a minimum of 100,000 cycles without signs of fatigue. Submit independent test data. Open and close sash against rubber bumper stops.

3.	Horizontal sash:  Bottom H-shaped rail to be 18 gage thick, Type 304 stainless steel with stainless steel roller wheels set on 18 gage thick stainless steel track. Top track to be extruded polyvinyl chloride set in 18 gage stainless steel U-channel assembly.

4.	Sash glass:  7/32 inch thick laminated safety glass.

5.	Sash cables:  Stainless steel, 1/8 inch diameter.

6.	Sash guides:  Corrosion resistant polyvinyl chloride.

7.	Pulley assembly for sash cable:  2 inches diameter, zinc dichromate finish, ball bearing type, with cable retaining device.

8.	Sash pull:  Type 304 stainless steel.

D.	Liner:

1.	Type 304 stainless steel liner:  16 gage thick, welded seamless construction. All vertical and horizontal corners and seamless joints between work surface and sides and back coved to a 1/2 inch radius, continuous welded, ground and blended to a fine grained finish. Seams between working surface and sides and back are not acceptable. Provide gasketed stainless steel access panels in the liner for access to valves and fittings.

E.	Baffles:

1.	Provide baffles to control air vectors into and through the fume hood. Fabricate of the same material as the liner. All baffle supports/brackets to be nonmetallic.

2.	Baffle designs which permit close-off of all slots are not acceptable.

3.	Must comply with OSHA Lab Standard Guidelines.

4.	Provide acid resistant label indicating proper baffle operation. Locate label on sidewall of hood interior.

5.	Provide for baffle adjustment to allow repositioning prior to hood operation permitting setting for (1) high thermal loading, (2) heavier than air gases or fumes generated near work surface, and (3) normal or average operation.

F.	Work Surfaces:

1.	Type 304 stainless steel:  16 gage, welded seamless construction with a raised 1/4 inch anti-spill edge. Stainless steel, reinforced with 16 gage steel channels, with an overall thickness of 1 inch.

2.	Reinforce work surface to handle loads of 200 pounds per square foot.

G.	Sinks

1.	Cup sinks:  Stainless steel

a.	Size:  4-3/8 inches by 13-3/4 inches by 5-7/16 inches deep.

b.	Constructed of Type 304 stainless steel, minimum thickness 18 gage , Number 4 mil. polish finish. All joints butt welded and ground smooth. Pitch bottoms to drain outlet.

c.	Screen:  Type 316 stainless steel, 0.028 wire diameter, 8 by 8 mesh per square inch.

d.	Provide swivel sink strainer adapter, Town & Country Plastics Model PP-18; R&G Sloane Part Number 7218; or Scientific Plastics Company, Inc., Part Number 31595-158.

H.	Lighting:

1.	Fixture:  2 lamp, rapid start, UL listed fluorescent light fixture with sound rated ballast installed on exterior of fume hood roof. Provide safety glass panel cemented and sealed to the hood roof.

2.	Interior of fixture:  White, high reflecting plastic enamel.

3.	Size:  Largest possible up to 48 inches for hoods with superstructures up to 6 feet. Provide (2) 36 inch fixtures for hoods with 8 foot superstructures.

4.	Include lamps with fixtures.

5.	Illumination:  80 footcandle minimum.

6.	Lamps and ballast to be the same type as specified in Division 16 for overhead lighting in laboratory spaces.

I.	Electrical Power:

1.	Three wire grounding type receptacles rated at 120 VAC at 20 amperes.

2.	Receptacles to be gray in color with stainless steel flush plates.

3.	Factory install receptacles, lighting, electrical fittings and wiring in accordance with applicable codes and Division 16 specifications.

4.	Provide a dedicated receptacle on top of hood for safety monitor and alarm.

J.	Service Fittings and Piping:

1.	Hoods shall be factory prepiped to a point of connection 2 inches above the hood roof. Piping materials and installation for each type of service shall be as specified under Division 15. Pressure test all prepiped lines in the factory.

2.	Service fittings within fume hoods shall have color coded, acid and solvent resistant plastic coating, applied over fine sandblasted surface, properly cleaned. Surfaces to be coated shall be sprayed and baked three times with minimum coating thickness of 6 mil.

3.	Remote Control Valves:  Remote control valves shall be mounted on the front panel of the fume hood, with all components subject to wear accessible from the exterior of the hood. The valve stem shall be for fume hoods with airfoil 45 degrees to the side wall of the fume hood. The centerline of the valve inlet and outlet shall be parallel and 1-1/8 inches apart. Valves shall have a threaded collar to hold the valve in place.

4.	Color coding of hood interior service fittings and remote control valves shall be as specified for Laboratory Service Fittings in the Laboratory Casework and Furnishings Section of this Specification.

K.	Exhaust Outlet:

1.	For hoods with stainless steel liners, 18 gage thick, Type 304 stainless steel collars welded in place.

2.	Provide exhaust duct transition from the hood rectangular exhaust outlet to the building exhaust system round duct connection. Fabricate transition of the same material as the exhaust outlet.

L.	Safety Monitor/Alarm System:

1.	Each isotope hood shall be equipped with a safety alarm unit which is to be provided under Division 15. Provide factory cut-outs in the front of the hood for field mounting of the unit. Coordinate location and size of cut-out with Division 15. Connection of safety alarm unit to sensors and fans shall be under Division 15.

M.	Safety Monitor/Alarm Systems:

1.	Provide fume hood monitor with audible and visual alarm signals to activate if face velocity drops below 80 fpm or rises above 180 fpm. Alarm setpoints to be field adjustable.

2.	Audible alarm signal tone shall be a minimum of 85 decibels at 10 centimeters. Visual alarm signal shall be a flashing or continuous glowing red light. Alarm design shall permit silencing of tone signal, but red visual signal light shall glow until unsafe condition is corrected. Provide with circuitry and a pushbutton to test alarm circuits.

3.	Monitor shall utilize thermal sensors, located in the monitor, to determine hood face velocity.

4.	Provide with input/output accessory to send alarm signal to a remote location and for silencing alarm remotely.

5.	Electrical:  9 volt DC, maximum of 1 amp. Provide 120 VAC to 9 VDC transformer, UL listed.

N.	Instruction Plate:

1.	Permanent 4 inch by 6 inch plate attached to fume hood exterior with condensed information covering recommended locations for apparatus and accessories, baffle settings, use of sash, and recommended safe operating procedures.

2.	Provide a 4 inch by 6 inch corrosion resistant metal or plastic frame attached to front face of fume hood exterior with plastic or glass glazing to contain a removable data card to identify hood and show hood characteristics and field test data.

O.	Isotope Filter

1.	Stainless steel filter assembly with rough and HEPA filter. Housing shall be stainless steel.

2.	Capacity:   1300 cfm at 1.6 inches static pressure with clean filters.

3.	Size:  Overall dimensions shall be 24-1/2 inches wide, 24 inches high, 25-1/4 inches deep, with a 12 inch diameter inlet and outlet.

4.	Provide manometer mounted on housing to indicate filter pressure drop and report on filter condition.

5.	Each HEPA filter shall be tested and rated at 99.9 percent efficiency, minimum for particles 0.3 to 0.5 microns in diameter. Housing shall be stainless steel.

2.06	WALK-IN HOODS

A.	Walk-in fume hoods shall be as detailed on the Drawings and as described under Article 2.03, Benchtop Chemical Fume Hoods, unless modified under this Article.

B.	Floor:  Provide a removable stainless steel floor, Type 316, mounted on marine plywood.

C.	Support channels:  Type 316 stainless steel as detailed.

2.07	CANOPY HOODS

A.	Fabricate of 18 gage thick, mild steel as detailed. Paint entire hood. Color selected by Owner's Representative.

B.	Fabricate, as detailed, of 18 gage thick, Type 304 stainless steel with a Number 4 finish.

C.	All exposed corners are to be welded and ground smooth.

D.	Support from steel rods attached to structure above and from brackets attached to adjoining wall. Refer to detail on Drawings.

2.08	FUME HOOD METAL FINISH

A.	Preparation:  Spray clean metal with a heated cleaner/phosphate solution, pretreat with iron phosphate spray, water rinse, and neutral final seal. Immediately dry in heated ovens, gradually cool, prior to application of finish.

B.	Application:  Electrostatically apply epoxy or urethane powder coat of selected color and bake in controlled high temperature oven to assure a smooth, hard satin finish. Surfaces shall have a chemical resistant, high grade laboratory furniture quality finish of the following thicknesses:

1.	Exterior and interior surfaces exposed to view:  1.5 mil average and 1.2 mil minimum.

2.	Surfaces not exposed to view:  1.0 mil average.

2.09	SOURCE QUALITY CONTROL TESTING OF EPOXY RESIN WORK SURFACE

A.	Test procedure:  Apply 5 drops of each reagent to surface and cover with 25 mm watch glass, convex side down; test volatiles using 1 ounce bottle stuffed with saturated cotton. After 24 hour exposure flush surface, clean, rinse and wipe dry.

B.	Evaluation ratings:  Change in surface finish and function shall be described by the following ratings:

1.	No Effect:  No detectable change in surface material.

2.	Excellent:  Slight detectable change in color or gloss, but no change to the function or life of the work surface material.

3.	Good:  Clearly discernible change in color or gloss, but no significant impairment of work surface function or life.

4.	Fair:  Objectionable change in appearance due to surface discoloration or etch, possibly resulting in deterioration of function over an extended period.

5.	Failure:  Pitting, cratering or erosion of work surface material; obvious and significant deterioration.

C.	Test Results:  Submit a report of the test results. The results shall be equal to or better than the following:

Reagent					Rating

1.	Hydrochloric Acid, 37 percent	Excellent

2.	Sulfuric Acid, 33 percent		No Effect

3.	Sulfuric Acid, 77 percent		No Effect

4.	Sulfuric Acid, 96 percent		Failure

5.	Formic Acid, 90 percent		Excellent

6.	Nitric Acid, 20 percent		Excellent

7.	Nitric Acid, 30 percent		Excellent

8.	Nitric Acid, 70 percent		Good

9.	Hydrofluoric Acid, 48 percent	Fair

10.	Phosphoric Acid, 85 percent	No Effect

11.	Chromic Acid, 60 percent		Failure

12.	Acetic Acid, 98 percent		Excellent

13.	3 & 8 Equal Parts			Excellent

14.	Ammonium Hydroxide, 28 percent	No Effect

15.	Sodium Hydroxide, 10 percent	No Effect

16.	Sodium Hydroxide, 20 percent	No Effect

17.	Sodium Hydroxide, 40 percent	No Effect

18.	Sodium Hydroxide Flake		No Effect

19.	Sodium Sulfide			Excellent

20.	Zinc Chloride			No Effect

21.	Tincture of Iodine			Excellent

22.	Silver Nitrate			No Effect

23.	Methyl Alcohol			No Effect

24.	Ethyl Alcohol			No Effect

25.	Butyl Alcohol			No Effect

26.	Benzene				Excellent

27.	Xylene					No Effect

28.	Toluene				Excellent

29.	Gasoline				No Effect

30.	Dichloroacetic Acid			Good

31.	Dimethylformamide			Excellent

32.	Ethyl Acetate			No Effect

33.	Amyl Acetate			Excellent

34.	Acetone				Excellent

35.	Chloroform				Excellent

36.	Carbon Tetrachloride		No Effect

37.	Phenol					Excellent

38.	Cresol					Excellent

39.	Formaldehyde			No Effect

40.	Trichlorethylene			Excellent

41.	Ethyl Ether				Excellent

42.	Furfural				Good

43.	Methylene Chloride			Excellent

44.	Monochlorobenzene			Good

45.	Dioxane				Excellent

46.	Methylethyl Ketone			Excellent

47.	Acid Dichromate			Fair

48.	Hydrogen Peroxide			No Effect

49.	Naphthalene				Excellent

2.10	SOURCE QUALITY CONTROL TESTING OF POLY RESIN LINER

A.	Test Procedure:

1.	Test Number 1 - Spills and Splashes:

a.	Suspend in a vertical plane a 42 inch (horizontal) by 12 inch (vertical) panel divided into 3/4 inch wide vertical columns.

b.	Apply 5 drops of each reagent listed with an eye dropper.

c.	Apply liquid reagents at top of panel and allow to flow down full panel height. (CAUTION!  Flush away any reagent drops.)

2.	Test Number 2 - Fumes and Gases:

a.	Divide 24 inch by 12 inch panel into 2 inch squares.

b.	Place 25 milliliters of reagent into 100 milliliter beakers and position panel over beaker tops in the proper sequence. Note:  Beaker pouring lip permits atmospheric oxygen to enter and participate in the reaction of the reagent fumes.

3.	After 24 hours remove panel, flush with water, clean with naphtha and detergent, rinse, wipe dry and evaluate.

B.	Evaluation ratings:  Change in surface finish and function shall be described by the following ratings:

1.	No Effect:  No detectable change in surface material.

2.	Excellent:  Slight detectable change in color or gloss, but no change to the function or life of the work surface material.

3.	Good:  Clearly discernible change in color or gloss, but no significant impairment of work surface function or life.

4.	Fair:  Objectionable change in appearance due to surface discoloration or etch, possibly resulting in deterioration of function over an extended period.

5.	Failure:  Pitting, cratering or erosion of work surface material; obvious and significant deterioration.

C.	Test Results:  Submit a report of the test results. The results shall be equal to or better than the following:

Reagent List				Test 1			Test 2

Concentrations by Weight		Rating Spills 		Fumes

1.	Sodium Hydroxide Flake		No Effect		No Effect

2.	Sodium Hydroxide, 40 percent	Excellent		No Effect

3.	Sodium Hydroxide, 20 percent	Excellent		No Effect

4.	Sodium Hydroxide, 10 percent	Excellent		No Effect

5.	Ammonium Hydroxide, 28 percent	No Effect		No Effect

6.	Methylene Chloride			No Effect		No Effect

7.	Chloroform				No Effect		No Effect

8.	Carbon Tetrachloride		No Effect		No Effect

9.	Trichlorethylene			No Effect		No Effect

10.	Monochlorobenzene			No Effect		No Effect

11.	Tincture of Iodine			No Effect		Good

12.	Methyl Alcohol			No Effect		No Effect

13.	Ethyl Alcohol			No Effect		No Effect

14.	Butyl Alcohol			No Effect		No Effect

15.	Phenol, 85 percent			No Effect		Excellent

16.	Cresol					No Effect		No Effect

17.	Sodium Sulfide, Saturated		Good			Excellent

18.	Furfural				Good			No Effect

19.	Dioxane				No Effect		No Effect

20.	Zinc Chloride			No Effect		No Effect

21.	Benzene				No Effect		No Effect

22.	Toluene				No Effect		No Effect

23.	Xylene					No Effect		No Effect

24.	Gasoline				No Effect		No Effect

25.	Naphthalene				No Effect		No Effect

26.	Methylethyl Ketone			No Effect		No Effect

27.	Acetone				No Effect		No Effect

28.	Ethyl Acetate			No Effect		Excellent

29.	Amyl Acetate			No Effect		No Effect

30.	Ethyl Ether				No Effect		No Effect

31.	Silver Nitrate, 10 percent		Excellent		No Effect

32.	Dimethylformamide			No Effect		No Effect

33.	Formaldehyde, 37 percent		No Effect		No Effect

34.	Formic Acid, 88 percent		No Effect		No Effect

35.	Acetic Acid, Glacial			No Effect		No Effect

36.	Dichloroacetic Acid			No Effect		No Effect

37.	Chromic Acid, Saturated		Fair			No Effect

38.	Phosphoric Acid, 85 percent	Excellent		No Effect

39.	Sulfuric Acid, 33 percent		Excellent		No Effect

40.	Sulfuric Acid, 77 percent		Excellent		No Effect

41.	Sulfuric Acid, 93 percent		Excellent		No Effect

42.	Hydrogen Peroxide, 30 percent	No Effect		No Effect

43.	Acid Dichromate			Excellent		Excellent

44.	Nitric Acid, 20 percent		No Effect		No Effect

45.	Nitric Acid, 30 percent		No Effect		No Effect

46.	40 & 47 Equal Parts			No Effect		Excellent

47.	Nitric Acid, 70 percent		No Effect		No Effect

48.	Hydrofluoric Acid, 37 percent	No Effect		No Effect

49.	Hydrochloric Acid, 48 percent	Excellent		Excellent

2.11	SOURCE QUALITY CONTROL TESTING OF FUME HOODS

A.	Test one fume hood of each type and size specified in accordance with the method prescribed in American National Standard ANSI/ASHRAE 110-1995. The minimum overall performance rating of each test shall be 4.0 AM 0.05 with 4.0 equal to liters per minute of tracer gas release, AM identifying an "as manufactured" test, and 0.05 indicating the level of tracer gas control in parts per million.

B.	Hood test shall take place in the manufacturer's test facility with testing personnel, samples, apparatus, instruments, and test materials supplied by the manufacturer.

2.12	SOURCE QUALITY CONTROL TESTING OF FUME HOODS

A.	Submit a test report, for each type and size of hood, for the standard product previously tested, if the product is identical to equipment being provided for this project.

B.	Evaluation of standard product shall have been conducted in the manufacturer's test facility in accordance with the method prescribed in ANSI/ASHRAE 110-1985.

C.	Hoods shall achieve a rating of 4.0 AM 0.05 PPM or better.

2.13	SOURCE TESTING OF FUME HOOD SOUND LEVELS

A.	Provide the following historic certified sound test data for each size and type of fume hood:

1.	Background sound pressure level readings for the test facility with exhaust system operating but without connection to the fume hood.

2.	A second set of readings recorded with the fume hood connected and operating and the sensor located 36 inches in front of the sash assembly.

3.	Certified report:  Octave band sound pressure level, Db re 20 micro Pa in the 31.5 to 8,000 Hertz frequency range, for the fume hood operating.

PART 3 -  EXECUTION

3.01	INSTALLATION

A.	Installation:

1.	Install fume hoods and equipment in accordance with manufacturer's instructions.

2.	Install equipment plumb, square, and straight with no distortion and securely anchored as required.

3.	Secure work surfaces to casework and equipment components with material and procedures recommended by the manufacturer.

B.	Accessory installation:  Install accessories and fittings in accordance with manufacturer's recommendations.

3.02	FIELD QUALITY CONTROL TESTING OF FUME HOODS

A.	Testing Requirements:

1.	Perform tests in field to verify proper operation of the fume hoods before they are put in use, using only qualified personnel.

2.	Perform tests after installation is complete, the building ventilation system has been balanced, all connections have been made, and written verification  has been submitted that the above conditions have been met.

3.	Verify that the building makeup air system is in operation, the doors and windows are in normal operating position, and that all other hoods and exhaust devices are operating at designed conditions.

4.	Correct any unsafe conditions disclosed by these tests before request of test procedures.

B.	Testing Equipment:

1.	Properly calibrated hot wire thermal anemometer.

2.	Supply of 30 second smoke bombs.

C.	Face Velocity Testing:

1.	Conduct tests in accordance with procedures as outlined in SEFA-1-1992.

2.	Check room conditions in front of fume hood using a thermal anemometer and a smoke source to verify that the velocity of cross drafts does not exceed 20 percent of the specified average fume hood face velocity. Eliminate any cross drafts that exceed these values before proceeding.

3.	Perform the following test to verify conformance of actual fume hood face velocities to those specified. With the sash in the open position, as scheduled on the Drawings, determine the face velocity by averaging the velocity of 6 readings taken at the fume hood face; at the centers of a grid made up of 3 sections of equal area across the top half of the fume hood face and 3 sections of equal area across the bottom half of the fume hood face. If not in accordance with scheduled face velocities, refer to manufacturer's Troubleshooting Guide for aid in determining cause of variation in airflow.

4.	Check sash operation by moving sash through its full travel. Verify that sash operation is smooth and easy, and that vertical rising sash shall hold at any height without creeping up or down.

D.	Smoke Testing:

1.	With sash in the open position check air flow into the fume hood using a smoke source. Verify that air flow is into the fume hood over the entire face area by a complete traverse of the fume hood 6 inches inside the face. Reverse flow is evidence of unsafe conditions. Take necessary corrective actions and retest.

2.	Move a lighted smoke bomb throughout the fume hood work area directing smoke across the work surface and against the side walls and baffle. Verify that smoke is contained within the fume hood and rapidly exhausted.

3.03	FIELD QUALITY CONTROL TESTING OF FUME HOODS

A.	Refer to Division 15 specification sections for air balancing and testing of fume hoods.

3.04	ADJUSTING

A.	Repair or remove and replace defective work, as directed by the Owner's Representative upon completion of installation.

B.	Adjust sash, fixtures, accessories and other moving or operating parts to function smoothly.

3.05	CLEANING

A.	Clean equipment, touch up as required.

3.06	PROTECTION OF FINISHED WORK

A.	Provide all necessary protective measures to prevent damage to equipment from exposure to other construction activity.

B.	Advise Contractor of procedures and precautions for protection of material and installed fume hoods from damage by work of other trades.

END OF SECTION