DETECTION AND MONITORING EQUIPMENT

Multi-RAE Plus Gas Detector  [10 to 30 seconds]

- Oxygen (O2) Level = 21% total volume

- % Lower Explosive Limit (LEL) = 0 %

- Toxic Sensor Readings = 2 parts per million (ppm)

M-8  - Paper turns BLUE-BLACK or DARK GREEN.  [30 seconds]

- Will not detect chemical agents in gas or vapor form.

- Will give a false positive response when exposed to cleaning solutions, pesticides, aromatics, hydraulic oil and some decontaminants.

AIMS 4501 – Detects up to 31(programmed) chemical gases. Will not test for organophosphates.

Colorimetric Tubes - Draeger’s Chemical Warfare Agent Detection Tubes (Civil Defense Kit)

Phosphoric Acid Esters 0.05/a  (0.05 ppm or 0.5mg/m3).  From Set B, the YELLOW changes to RED.  [Approximately 5 minutes]

Pesticide Tickets - Detects organophosphates, thiophosphates, and carbamate pesticides.  Works on a

fail-safe system.  If the test fails, the ticket remains white and will keep the responders on the safe

side, indicating that there may be contamination until proven otherwise.  [10-15 minutes]

- Blue change = No pesticide detected.

- White = Possible pesticide present.

- Range of detection 0.1-10 ppm.

Photoionization Detector  (PID)  - Use 10.6eV or 11.7eV ultraviolet lamp to ionize chemicals with an

ionization potential (IP) below 10.6 to measure their concentration in parts per million

(ppm) with a CF (correction factor) of about 0.5 for VX in gas/vapor form.  [Instant to 20 seconds]

M-256A1 - Use the Star test spot.  Used primarily for monitoring because the reaction time of its contents

is too slow to allow for effective warning capability.  Sensitivity:  0.02 mg/m3   [Approx. 15 minutes]

# M-18A2 - Enzyme ticket turns blue.  Used to confirm the results of the M-256.  [2-3 minutes]

Blue color on both ends (color may not be evenly distributed) = NO AGENT

Square end colorless to orange.  Round end blue = nerve agent present.

ICAM APD  (Improved Chemical Agent Monitor – Advanced Portable Detector)

Sensitivity:  V – 0.04 mg/m3 < 30 seconds.  [High concentration of this agent < 10 seconds]

Operation temperature:  -30oC to +52o C.  Alarms:  Easy to read visual display and two

audio alarms:  warning and test.

TMX412 Multi-Gas Monitor [45 seconds]

- Oxygen (O2) Level = 21% total volume

- % Lower Explosive Limit (LEL) = 0 %

- Toxic Sensor Readings = 0 parts per million (ppm)

M-8A1 – Alarm system that can automatically detect nerve agents (VX) in vapor concentrations only.

When a nerve agent is detected at a remote location, a warning light and horn sound.  [2 minutes]

APD2000  (Advance Portable Detector) – Detects chemical warfare agents and provides actual identity and relative concentration of agent under most environmental conditions. APD2000 has selectable operating modes (monitor or detect). Easy to read visual and audible alarms. Sensitivity response time for VX is 30 sec at 4 parts per billion (ppb), for high concentration of agent response time is 10 seconds

READINGS FOR PARATHION

STOCKTON CWFSE

Multi-RAE Plus Gas Detector  	Reading:  2 ppm

M-8 Paper			Reading:  Paper turns BLUE-BLACK or DARK GREEN.  [30 seconds]

AIMS 4501 			Will not test for organophosphates.

Colorimetric Tubes – Draeger Tubes (Civil Defense Kit)

Phosphoric Acid Esters 0.05/a  (0.05 ppm or 0.5mg/m3).  From Set B, the YELLOW changes to RED.  [Approximately 5 minutes]

Pesticide Tickets	 	Reading:  [10-15 minutes]

Blue change = No pesticide detected.

White = Possible pesticide present.

Photoionization Detector  (PID)  	Reading:  0.5 ppm for gas/vapor form.  [Instant to 20 seconds]

M-256A1 			Reading:  Positive for nerve agents

# M-18A2 				Reading:  Enzyme ticket turns blue

Blue color on both ends - NO AGENT

Square end colorless to orange

TMX412 Multi-Gas Monitor 	Reading:  0 ppm

M-8A1 				Sounds when vapor is detected

APD2000  			Reading:  4 parts per billion (ppb)