Chilled Mirror Cleaning & Recalibration

It is best to do the recalibration and cleaning in an indoor environment.  Stable temperatures and dew points are needed.  If the unit is outdoors or in an environment where temperature and dew point is fluctuating during the recalibration the settling time of the unit will take much longer.

Items needed:  Fluke meter able to read 0 – 5 V DC, distilled water, lacquer thinner, isopropyl alcohol, cotton swabs.

- Remove CMH unit from the tower and take inside the GW or back to the duplex.

- Leave power disconnected from unit at this time.

- Turn Auto balance dial to 000 so that 0 is read in the “window” and 0 is set to the mark below the window.

- Remove mirror assembly from the sensor housing assembly (white mushroom).

  - Remove retaining screw from side of housing assembly

  - Grab sensor (from the bottom) by the handle and pull assembly out, you may want to disconnect the sensor data cable.  The unit connects through two banana connectors and slides in and out of assembly only one way.

- Once unit is removed you can begin the cleaning of the mirror.  Begin by taking a clean cotton swab and distilled water, thoroughly wet mirror surface using a moistened swab.  Once mirror is wet, wash with a gentle circular motion.

- Wipe the surface with a clean dry swab until dry and all loosened material is removed.

- Repeat steps 4 and 5 until the cleaning has no further effect, i.e. dry swab is clean after drying sequence.

- Repeat steps 4 & 5 using lacquer thinner.

- Continue lacquer thinner cleaning until it has no further effect.

- Repeat steps 4 & 5 using Isopropyl Alcohol.

- Continue alcohol cleaning until it has no further effect.

- NEVER use soap or any other detergent to wash mirror.

- Re-install mirror assembly into the sensor housing assembly.  Ensure the banana connector line up before seating mirror assembly.  Tighten retaining screw and ensure the grounding cable is connected also between the assembly cap and retaining screw.

- Reconnect the sensor data cable to the main housing unit.

- Apply AC power to the unit.

- Turn the Mode Selector Switch to HEAT, and leave on for 1 minute.  This will ensure a dry mirror after following the cleaning steps above.  Turn mode Selector Switch to OPR position after 1 minute time period is up.

- Remove AC power.

- Connect a digital voltmeter set to read 0 – 5 V DC to the transmit logic board.  Connect the positive lead to the left side of R3 (junction of R8, R3, and C2) on the transmit logic board and the negative lead to ground.

- Reapply AC power.

- THE FOLLOWING STEPS MUST BE DONE WITH A COMPLETELY DRY MIRROR.  TO INSURE A DRY MIRROR CHECK THAT THE DEW POINT IS ABOVE THE AMBIENT AIR TEMPERATURE BY TOGGLING THE MONITOR SELECTOR SWITCH BETWEEN Ta and Td AND PUSHING THE MONITOR DISPLAY SWITCH TO READ THE VALUES.  NEED TO KEEP Td ABOVE Td.  THE FOLLOWING ADJUSTMENTS MUST NOT BE MADE WHILE HEATING THE MIRROR.  YOU MUST MONITOR THE Td, AND HEAT AS NECESARRY.  MAKE MEASUREMENTS AND ADJUSTMENTS IN BETWEEN HEATING CYCLES.  SELECTOR SWITCH SHOULD BE IN OPR POSITION WHEN MAKING ADJUSTMENTS

- Turn the Sd gain adjustment (R21) clockwise until the Sd LED (CR5) just goes on.  This is the threshold level.  The voltmeter should read approximately 4.9 Volts.

- Adjust the Si gain adjustment (R22) Clockwise until the Si LED (CR9) goes on then let system stabilize then turn the Si gain adjustment CounterClockwise until the Si LED (CR9) just goes off.