MSFC Problem Reporting and Corrective Action (PRACA) System

WHOLE RECORD REPORT( + ADDENDUM)

| MSFC Record # A10000 | In-Flight Anomaly Number -- | Contractor Report Number E-098 | JSC# -- | KSC# -- |
| --- | --- | --- | --- | --- |
| Problem Title ELECTRICAL TEST EQUIPMENT FAILURE |
| EICN# -- | ELEMENT ET | Contractor MMMSS | FSCM# -- | FCRIT 1 |
| HCRIT -- | Sys_Lvl Y | Misc Codes A (3) B C D E F G H I J K L M N O |
| HARDWARE EIM | NOMENCLATURE ET COMPLETE | PART# 80901010000 | SER/LOT# LWT-32 | MANUFACTURER MMC |
| HARDWARE LRU | NOMENCLATURE N/A | PART# N/A | SER/LOT# N/A | MANUFACTURER N/A |
| HARDWARE NCA | NOMENCLATURE ET COMPLETE | PART# 90901010000 | SER/LOT# LWT-32 | MANUFACTURER MMC |
| Test/Operation A - ATP | Prevailing Condtion F - FUNCTIONAL | F / U F | Fail Mode EV - NOT-TO-SPEC | Cause ETT - EI-TEST-EQUP |
| System ELECTRICAL | Defect -- | Material Y - OPER-N | Work Contact J. ADAMS | Fail Date 03/06/1986 |
| Received at MSFC 04/09/1986 | Date Isolated -- | FMEA Reference N/A | IFA: Mission Phase -- | Mission Elapsed Time -- |
| Location MAF | Symptom EV - NOT-TO-SPEC | Time Cycle -- |
| Effectivity Text LWT-27, -28, -30 AND -32 |
| Vehicle Effectivity Codes |
| Vehicle 1 -- | Vehicle 2 -- | Vehicle 3 -- | Vehicle 4 -- | Vehicle 5 -- |
| Mission Effectivity Codes |
| Mssn 1 -- | Mssn 2 -- | Mssn 3 -- | Mssn 4 -- | Mssn 5 -- |
| Estimated Completion Dates |
| MSFC Approved Defer Until Date -- | Contractor Req Defer Until Date -- | LVL 3 Close -- | Remark / Action -- |
| Investigation / Resolution Summary |
| Last MSFC Update 02/13/1995 | CN RSLV SBMT 07/16/1986 | Defer Date -- | Add Date -- | R/C Codes 3 - F/TE -- -- |
| Assignee |
| Design W. COBB | Chief Engineer -- | S & MA R. JACKSON | Project M. PESSIN | Project MGR -- |
| Approval |
| Design W. COBB | Chief Engineer -- | S & MA R. JACKSON | Project M. PESSIN | Project MGR -- |
| PAC Assignee G. MILLER | PAC Review Complete GM | MSFC Closure Date 08/08/1986 | Status C - CLOSED | F/A Completion -- |
| Problem Type -- | SEV -- | Program Name -- | REVL -- | OPRINC -- |
| FUNC MOD -- | Software Effectivity -- -- -- -- -- | Software Fail CD -- | SUBTYPE -- | Software Closure CD -- |
| RES PERSON L2 -- | Approval Signature L3 -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Related Document Type -- | Related Document ID -- |
| Related Document Title -- |
| Contractor Status Summary |
| Reliability/Quality Assurance Concerns, Recommendations: |
| Problem Description A FAILURE OF THE ELECTRICAL TEST EQUIPMENT USED FOR ET FINAL ACCEPTANCE TESTING IN MAF BUILDING 420, CELL 2, WAS DISCOVERED DURING RECALIBRATION ON MARCH 6, 1986. THE MEGOHMMETER FUNCTION OF THE APS CIRCUIT ANALYZER APPLIED LOWER VOLTAGES THAN THOSE REQUIRED FOR ISOLATION RESISTANCE TESTING. THE ANALYZER IS CALIBRATED AT SIX MONTH INTERVALS. THOSE TESTS, WHICH WERE PERFORMED SINCE THE LAST CALIBRATION, WHICH OCCURRED ON SEPTEMBER 6, 1985, ARE SUSPECT OF BEING INADEQUATE TO MEET THE TM04 TEST REQUIREMENTS |
| Contractor Investigation/Resolution 4/10/86 - CONSTRAINTS: LWT'S 27, 28 AND 30 EFFECTIVITIES: LWT'S 27, 28, 30 AND 32 4/10/86 FAILURE INVESTIGATION - A. THE MAF BUILDING 420, CELL 2, ELECTRICAL TEST EQUIPMENT CAME DUE FOR THE SIXTH MONTH INTERVAL CALIBRATION CHECKS ON MARCH 6, 1986. DURING CALIBRATION, THE APS CIRCUIT ANALYZER WAS FOUND TO BE DEFECTIVE WHEN USING THE MEGOHMMETER FUNCTION. AN INTERNAL HIGH VOLTAGE POWER SUPPLY HAD PARTIALLY FAILED AS A RESULT, WHEN A 500 VOLT MEGOHMMETER TEST WAS SELECTED, ONLY 140 VOLTS WAS APPLIED. THE RESISTANCE MEASUREMENT CIRCUIT WAS ACCURATE, BUT WITH THE LOWER VOLTAGE APPLIED, THE CIRCUIT UNDER TEST WOULD NOT BE PROPERLY TESTED. THE CIRCUIT ANALYZER IS USED TO PERFORM CONTRACTUALLY REQUIRED TM04 TESTING OF EXTERNAL TANK CIRCUITRY. A REVIEW OF THE CIRCUIT ANALYZER SCHEMATIC DIAGRAMS FOUND THAT THE EQUIPMENT MANUFACTURER HAD MODIFIED THE RESISTANCE MEASUREMENT CIRCUITRY DURING AN UPGRADE OF THE SYSTEM. THE MODIFICATION HAD NOT BEEN REQUESTED BY MMC AND WAS NOT DETECTED UPON RETURN OF THE ANALYZER TO MAF FOLLOWING THE UPGRADE. THE CIRCUIT ANALYZER HAS A PARTIAL RECALIBRATION ("QUICK LOOK") PERFORMED ON A WEEKLY BASIS. TEST OPERATIONS, DEPARTMENT 3613, HAD INTENDED THAT WEEKLY TESTING BE AN ADEQUATE TEST OF ALL CIRCUIT ANALYZER FUNCTIONS, INCLUDING MEGOHMMETER TESTS. THE MODIFICATON TO THE CIRCUIT ANALYZER MADE RESISTANCE MEASUREMENTS MORE STABLE AND ACCURATE BUT MADE OBSOLETE THAT PORTION OF THE WEEKLY TEST WHICH CHECKED THE MEGOHMMETER FUNCTION. B. A REVIEW OF THE PRODUCTION FLOW AND TEST EQUIPMENT CALIBRATION IN MAF BUILDING 420, CELL 2, REVEALED THE FOLLOWING: - THE LAST SIX MONTH CALIBRATION TEST PRIOR TO MARCH 6, 1986 HAD BEEN PERFORMED ON FIRST SHIFT SEPTEMBER 6, 1985. - LWT-27 "OPERATIONAL INSTRUMENTATION, ALL SYSTEMS TEST," TP-6D201-AT WAS STARTED ON SEPTEMBER 14, 1985. - LWT'S 28 AND 30 - ALL ELECTRICAL TESTS WERE PERFORMED IN CELL 2. NOTE: LWT'S 29 AND 31 WERE TESTED IN CELL 1 AND ARE, THEREFORE, NOT AFFECTED. - LWT-32 - "OPERATIONAL INSTRUMENTATION, RESISTANCE AND ISOLATION RESISTANCE TEST" TP-6C204-AT WAS COMPLETED ON FEBRUARY 5, 1986. - THE CIRCUIT ANALYZER FAILED THE CALIBRATION TEST ON MARCH 6, 1986. CORRECTIVE ACTIONS - A. TEST OPERATIONS, DEPARTMENT 3613, HAS REVISED THE WEEKLY CALIBRATION TESTS FOR ALL CIRCUIT ANALYZERS AT MAF. THE TESTS NOW VERIFY THE MEGOHMMETER VOLTAGES. - TP-13A150-AT, PCN-14 HAS BEEN RELEASED FOR THE ANALYZERS IN CELLS 1 AND 2 IN BUILDING 420. - TP-13A50-HA, PCN-21 HAS BEEN RELEASED FOR THE ANALYZER IN HARNESS FABRICATION SHOP. - TP-13A55-FA, PCN-NEW HAS BEEN RELEASED FOR THE HUGHES CIRCUIT ANALYZER IN FINAL ASSEMBLY. B ELECTRICAL ENGINEERING, DEPARTMENT 3513, AND SYSTEMS TEST INTEGRATION ENGINEERING, DEPARTMENT 3545, WILL IDENTIFY THE RETEST REQUIREMENTS AND PARTICIPATE IN THE TEST PROCEDURE WRITING FOR LWT'S 27, 28 AND 30 CHANGE SUMMARY B0 1754 IDENTIFIES THE TEST REQUIREMENTS. CONTRACTS, DEPARTMENT 3300, HAS ISSUED RMO 86 MI-0312 TO AUTHORIZE THE WORK. NOTE: THIS ITEM WILL BE REVISED TO INCLUDE THE ESTIMATED COMPLETION OF THE RETEST OF THE ET'S WHEN THE SCHEDULE HAS BEEN SET. ITEM OPEN - ECD IS 5-2-86. 4/10/86 CLEARANCE EFFECTIVITIES - LWT-27 - CONSTRAINT TO FLIGHT. THE RESULTS OF MAF FINAL ACCEPTANCE TEST PROCEDURE TP-6D201-AT ARE SUSPECT. THE NECESSARY RETESTING WILL BE PERFORMED PRIOR TO VAFB RECEIVING INSPECTION TESTS. LWT'S 28 AND 30 - CONSTRAINT TO FLIGHT. THE RESULTS OF MAF FINAL ACCEPTANCE TEST PROCEDURES TP-6C204-AT AND TP-6D201-AT ARE SUSPECT. THE ET'S ARE TO BE RETESTED.. LWT-32 - NO CONSTRAINT. THE FAILURE OF THE TEST EQUIPMENT WAS DISCOVERED PRIOR TO THE COMPLETION OFTESTING AND DOCUMENTED ON MARS T-78769. THE DISPOSITION OF THE MARS WILL ACCOMPLISH ALL NECESSARY RETESTING PRIOR TO DELIVERY TO THE CUSTOMER. 4/17/86 PRB STATUS - THE AUTOMATED ELECTRICAL TEST EQUIPMENT IN CELL 2 OF MAF BLDG. 420 WAS FOUND TO BE DEFECTIVE DURING THE 6 MONTH CALIBRATION CHECK PERFORMED ON 3-6-86. THE MEGOHMMETER FUNCTION APPLIED A LOWER VOLTAGE THAN REQUIRED. ONLY LWT'S 27, 28, AND 30 WERE AFFECTED. MMC WILL SEND A TEAM WITH TEST SET TO VANDENBERG TO RETEST LWT-27. LWT'S 28 AND 30 WILL BE RETESTED AT MMC ECD NOT DETERMINED. 5/15/86 PRB STATUS - MSFC SUGGESTED THAT MMC INVESTIGATE THE POSSIBILITY OF ESTABLISHING A TREND CHART HISTORY TO ALERT PERSONNEL OF ANY CHANGE IN THE CALIBRATION TOLERANCE. MMC FEELS THE WEEKLY EQUIPMENT CHECK IS ADEQUATE. PROBLEM STILL OPEN. ECD IS 10-30-86. 6/23/86 - DURING THE INTERIM PERIOD OF NO LAUNCH ACTIVITY, ALL CONSTRAINTS ARE BEING TRACKED AGAINST STS-34 (61E); THEREFORE, THIS PROBLEM REPORT EFFECTIVITY IS BEING CHANGED FROM 61N TO 61E. 6/19/86 PRB STATUS - THE AUTOMATED ELECTRICAL TEST EQUIPMENT IN CELL 2 OF MAF BLDG. 420 WAS FOUND TO BE DEFECTIVE DURING THE 6 MONTH CALIBRATION CHECK PERFORMED ON 3-6-86. THE MEGOHMMETER FUNCTION APPLIED A LOWER VOLTAGE THAN REQUIRED. ONLY LWT'S 27, 28 AND 30 WERE AFFECTED. MMC SENT A TEAM WITH A TEST SET TO VANDENBURG TO TEST LWT-27. TEST HAS BEEN COMPLETED AND NO DEFECTS WERE FOUND. LWT-28 AND LWT-30 ARE TO BE RETESTED AT MAF PER FINAL ACCEPTANCE TESTS TP-6C204-AT AND TP-6D201-AT (REF. BO-1754). CAPS CLOSURE IS BEING PREPARED AT MMC. 7/24/86 RESOLUTION - DURING A SIX MONTH INTERVAL RECALIBRATION OF THE CIRCUIT ANALYZER IN MAF BUILDING 420, CELL 2, IT WAS FOUND THAT THE MEGOHMMETER VOLTAGE WAS LOW. THE WEEKLY INTERVAL CALIBRATION CHECK HAD FAILED TO DETECT THE LOW VOLTAGE. THE ANALYZER HAS BEEN REPAIRED. THE FOUR ET'S WHICH HAD BEEN TESTED USING THE CIRCUIT ANALYZER DURING THE PRECEDING SIX MONTHS ARE BEING RETESTED. THE WEEKLY CALIBRATION CHECKS FOR ALL CIRCUIT ANALYZERS HAVE BEEN IMPROVED SO THAT FUTURE SIMILAR FAILURES WILL BE DETECTED. ELECTRICAL ENGINEERING, DEPARTMENT 3513, AND SYSTEMS TEST INTEGRATION ENGINEERING, DEPARTMENT 3545, IDENTIFIED ALL RETEST REQUIREMENTS FOR LWT'S 27, 28 AND 30 IN CHANGE SUMMARY B01754 CONTRACTS, DEPARTMENT 3300, ISSUED RMO 86MI-0312 TO AUTHORIZE THE WORK THE FAILURE OF THE TEST EQUIPMENT WAS DISCOVERED PRIOR TO TEST COMPLETION ON LWT-32, AND WAS DOCUMENTED ON MARS T-78769. THE MARS DISPOSITION ACCOMPLISHED ALL NECESSARY RETESTING ON LWT-32. NOTE: LWT'S 27, 28, 30 AND 32 WERE THE ONLY TANKS TESTED WITH THE DEFECTIVE CIRCUIT ANALYZER. RECOMMEND THIS PROBLEM REPORT BE CLOSED |
| MSFC Response/Concurrence |

MSFC Problem Reporting and Corrective Action (PRACA) System

ASSESSMENT ADDENDUM REPORT

| MSFC Report# A10000 | IFA# -- | Contractor RPT# E-098 | JSC# -- | KSC# -- | EICN# -- |
| --- | --- | --- | --- | --- | --- |
| Asmnt Part# 80901010000 | Asmnt Part Name ET COMPLETE | Asmnt Serial/Lot# LWT-32 |
| HCRIT CD -- | FCRIT CD 1 | CAUSE CD ETT - EI-TEST-EQUP | FAIL MODE EV - NOT-TO-SPEC |
| Asmnt FMEA N/A | Asmnt FM N/A | FMEA CSE N/A | FMEA SCSE N/A |
| Asmnt FMEA -- | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Asmnt FMEA -- | Asmnt FM -- | FMEA CSE -- | FMEA SCSE -- |
| Correlated Part# -- | Correlated Part# -- | Correlated Part# -- |
| Associated LRU# -- | Associated LRU# -- | Associated LRU# -- |
| MAJOR DESIGN CHANGES |
| APRV DATE -- | DESCRIPTION OF CHANGES -- |
| ASSESSMENT TEXT |