## CSN URG Worksheet
Chemical Speciation Network |  |  |  |  | US Environmental Protection Agency
Performance Audit Worksheet |  |  |  |  | Office of Air and Radiation
URG 3000N
Location |  |  |  |  | Date
AQS Site ID |  |  |  |  | Latitude:
Longitude:
Audit Information
Auditor(s) |  |  |  |  | Affiliation
Operator |  |  |  |  | Affiliation
Phone No.
Sampler Model |  |  |  |  | Controller S/N
Pump S/N
Sampler S/N
Last Calibration Date |  |  |  |  | Collocated? | Yes
No
Reference Std Model |  |  |  |  | Reference Standard S/N
Calibration Date
Clock Test:
If Local Time is under daylight savings, convert Ref Std to Local Standard Time. Daylight Saving Time begins for most of the United States at 2:00 a.m. on the first Sunday of April. Time reverts to standard time at 2:00 a.m. on the last Sunday of October.
| Time (hh:mm) |  |  |  | Difference Minutes |  | 5 minutes or less?
| Ref Std |  | URG |  |  |  | Pass | Fail
Audit
Recalibrated
Date
Leak Test
| Initial Audit |  |  | After Correction |  | 225 mm Hg drop or higher fails
| A mm Hg |  |  | B mm Hg |  | Fail A | Fail B | Pass
Channel 1 |  |  | Channel 1
Flow Test |  |  |  | Calibration
For the reference standard, enter "UR" for under range and "OR" for over range flow readings.
|  | L/min |  |  |  |  | Less than 10%?
| Ref Std |  | URG |  | % Difference |  | Pass | Fail
Channel 1
Retest after Calibration
|  | L/min |  |  |  |  | Less than 10%?
| Ref Std |  | URG |  | % Difference |  | Pass | Fail
Channel 1
Reference Standard vs Design Flow
|  | L/min |  |  |  |  | Less than 10%?
| Ref Std |  | URG |  | % Difference |  | Pass | Fail
Channel 1
Retest after Calibration
|  | L/min |  |  |  |  | Less than 10%?
| Ref Std |  | URG |  | % Difference |  | Pass | Fail
Channel 1
Ambient Temperature Test
|  | Degrees C |  |  |  |  | Less than 2 degrees?
| Ref Std |  | URG |  | Difference |  | Pass | Fail
Retest After Recalibration
Pressure Test
|  | mm Hg |  |  |  |  | Less than 10 mm Hg?
| Ref Std |  | URG |  | Difference |  | Pass | Fail
Retest after recalibration
Please save file as <Location-CSN-Date>where "Location" is the common network site name.

## CSN TSA
Chemical Speciation Network
Monitoring Site Location:
AQS Site ID:
GPS Coordinates: |  | (Latitude)
(Longitude)
Assessor (Auditor) Name:
Affiliation:
Observer(s) Name and Affiliation:
Assessment Date:
Section 1. Organization and Responsibilities
1. Field Operations Manager
| Name: |  | Affiliation:
Phone:
Address:
Phone:
E-mail:
- Management of monitoring network;
- Coordination of field operations;
- Logistical support of field operations;
- Training on routine sampler operation, data validation and quality control data.
2. Monitoring Site Operator(s)
| Name: |  | Affiliation:
Phone:
Address:
Phone:
E-mail:
| Name: |  | Affiliation:
Phone:
Address:
Phone:
E-mail:
- Operation of samplers;
- Calibration of samplers;
- Maintenance of samplers;
- Maintenance of monitoring site.
Section 1. Organization and Responsibilities (Cont'd)
Audit Questions (Block for the correct answer is highlighted yellow. If answer other
| than correct answer, enter response in Comments Section.) |  |  |  | (O = Other)
RESPONSE
[5 empty] | Y | N | O
1. | Does the site operate under an approved quality assurance project plan (QAPP)?
Date of QAPP approval?
2. | Is the approved QAPP being used by this organization?
3. | Is a copy of the approved QAPP available for review by field operators?
If no, briefly describe how and where QA and QC requirements and SOPs are
documented.
If yes, does the approved QAPP contain the field operations SOP(s)?
4. | Have all appropriate personnel reviewed the QAPP?
5. | Are there any deviations from the field SOP(s) at your site?
If yes, briefly describe why.
6. | Have the site operators attended Chemical Speciation field operations training ?
When?
7. | Are corrective actions in place when Measurement Quality Objectives (MQOs) are not met
(e.g. out-of-control calibration data)?
If no, briefly describe why.
Comments Section for Section 1 (Place question number and comment)
Section 2. Safety
Audit Questions (Block for the correct answer is highlighted yellow. If answer other than
| correct answer, enter response in Comments Section.) |  |  |  | (O = Other)
RESPONSE
[5 empty] | Y | N | O
8. | Is anyone authorized to halt the program in the event of a health or safety hazard or
safety hazard or inadequate quality?
Authorizer:
9. | Has the operator been trained in the particular hazards of the instrument/materials
with which they are operating? (E.g., electrical, slip/fall, platform/security)
10. | Are personnel outfitted with any required safety equipment? E.g., extreme weather
clothing, harnesses, head gear, repellants.
11. | Are personnel trained regarding safety procedures? (E.g., vermin, weather, electricity)
Comments Section for Section 2 (Place question number and comment)
Section 3. Sampler Siting
Use 40 CFR Appendix A and E for siting requirements
Audit Questions (Block for the correct answer is highlighted yellow. If answer other than
| correct answer, enter response in Comments Section.) |  |  |  | (O = Other)
RESPONSE
[5 empty] | Y | N | O
12. | Does the location for the samplers and collocated sampler(s) conform to the siting
requirements of 40 CFR 58, Appendices A and E?
13. | Has permission been given for violations of the siting criteria? If yes, please explain.
14. | Are there any noticeable problems at the site that would affect sample integrity?
15. | Are there any visible sources that might influence or impact the monitoring instrument?
16. | What are the nearby sources?
Comments Section for Section 3 (Place question number and comment)
Basic siting criteria from 40 CFR Appendix A and E
1. | The height of the inlet to the sampler should be between 2 and 15 meters above ground surface.
2. | For samplers located on roofs or other structures, the minimum separation distance between the inlet and any
structure should be greater than 2 meters.
3. | The sampler should be located away from obstacles so that the monitor is at a distance least twice the height of the
obstacle. For example, a tree is 10 meters tall and is east of the sampler. The sampler would need to be placed at
least 20 meters away from a tree.
4. | An unrestricted air flow of 270° must exist around the inlet.
5. | If the sampler is located on the side of a building, a 180° air flow clearance is required.
6. | Sampler inlet should be placed at least 20 meters from the drip line of any tree.
7. | Minimum distance to any roadway is 10 meters, but this value is determined by the average daily number of vehicles
(refer to 40 CFR Part 58 Appendix E for exact table).
8. | The inlet for a co-location sampler and audit sampler should agree vertically within 1 meter.
9. | The closest horizontal distance to place a co-location sampler to a Lo-Vol sampler or Hi-Vol sampler is 1 and 2 meters,
respectively. The maximum horizontal distance a co-location sampler can be from any sampler is 4 meters.
Section 4. Monitoring Siting
Audit Questions (Block for the correct answer is highlighted yellow. If answer other than
| correct answer, enter response in Comments Section.) |  |  |  | (O = Other)
RESPONSE
[5 empty] | Y | N | O
17. | Are site logbooks present and maintained?
18. | Are required data sheets filled in properly, clearly, and completely?
19. | Does the operator keep the module handling area neat and clean?
20. | Is there adequate room to perform the needed operations?
21. | Do the sampler(s) appear to be well maintained and free of dirt and debris, bird/animal/insect
nests, excessive rust and corrosion, etc.?
22. | Are the walkways to the station and equipment kept free of tall grass, weeds, and debris?
23. | Is the sampling platform (if any) clean and in good repair?
Comments Section for Section 4 (Place question number and comment)
Section 5. Sample Handling
Audit Questions (Block for the correct answer is highlighted yellow. If answer other than
| correct answer, enter response in Comments Section.) |  |  |  | (O = Other)
RESPONSE
[5 empty] | Y | N | O
24. | Are all samples handled to avoid contamination and/or loss of material?
25. | Does the operator know how to perform trip and field blank events?
Have operator show steps. Document any discrepancy from SOP.
26. | Observe the following handling steps for routine samples, verifying that the operator follows the sample handling
SOPs correctly:
a. | Receipt and temporary storage of samples at the operator’s facility
b. | Documents receipt of sample on chain of custody form
c. | Inspection of the sample prior to sampling
d. | Installation of sample in the sampler
e. | Retrieval from the sampler after sampling
f. | Completion of chain of custody and field data forms supplied by the reporting organization
g. | If samples are transferred to the local operator’s facility, follows temporary transport
procedures
27. | How do you communicate sample handling problems and to whom?
Comments Section for Section 5 (Place question number and comment)
Section 6. Shipping
Audit Questions (Block for the correct answer is highlighted yellow. If answer other than
| correct answer, enter response in Comments Section.) |  |  |  | (O = Other)
RESPONSE
[5 empty] | Y | N | O
28. | Is there adequate freezer space for blue ice on site or in the office?
29. | Does site operator have knowledge of filter holding/use/shipping times?
30. | Are there weekend storage procedures in place?
31. | Are the field and chain of custody forms being filled out properly?
32. | Are the coolers and samples being packed according to the SOPs? Have site operator
demonstrate procedure and document any discrepancies.
Comments Section for Section 6 (Place question number and comment)
Section 7. Samplers |  |  |  |  | RESPONSE
[5 empty] | Y | N | O
33. | Does this site use a MetOne SASS or MetOne |  |  | MetOne SASS?
| SUPER SASS sampler? |  |  | MetOne SUPERSASS?
34. | What type of filters are used on this sampler?
Teflon?
Nylon?
Quartz?
35. | Does this site have a URG 3000N sampler in use?
36. | What type of filters are used on this sampler?
Teflon?
Nylon?
Quartz?
37. | When was the URG 3000N put into operation?
Comments Section for Section 7 (Place question number and comment)
Assessment Date:
Checks/Maintenance |  | Frequency | Requirement | Last Date | Performed Correctly?
[5 empty] | Y |  | N
Clock Check |  | Every Run | Current date, time ± 5 minute
Leak Check |  | Monthly | < 100 mL/min
Flow Rate check |  | Monthly | ± 10% working standard
Filter Temperature Check |  | Monthly | Current temp ± 2oC via NIST traceable thermometer
Ambient Temperature Check |  | Monthly | Current temp ± 2oC via NIST traceable thermometer
Ambient Pressure Check |  | Monthly | Current pressure ±10 mm Hg NIST traceable barometer
Inspect/Clean Cyclones |  | Monthly | Per Service Manual
Denuder Swap Out |  | Quarterly | As sent by Service Lab
Clean Inside of Housing |  | Semiannual | Per Service Manual
Clean Air Screens |  | Semiannual | Clear Obstructions to Flow
Flow Rate Calibration Device |  | Annually | Verify against NIST standard or sent to factory
Temperature Calibration Device |  | Annually | Verify against NIST standard or sent to factory
Pressure Calibration Device |  | Annually | Verify against NIST standard or sent to factory
Clock Check |  | Quarterly | Current date, time ± 5 minute
Leak Check |  | Quarterly | < 80 mL/min
Flow Rate Audit |  | Quarterly | +/- 10% of Audit standard
Filter Temperature Check |  | Quarterly | Current temp ± 2o C via NIST audit thermometer
External Temperature Audit |  | Quarterly | Current temp ± 2o C via NIST audit thermometer
Ambient Pressure Audit |  | Quarterly | Current pressure +/- 10 mm Hg NIST traceable barometer
Clean Cyclones |  | Quarterly | Per Service Manual
Flow Rate Audit Device |  | Annually | Verify against NIST standard or sent to factory
Temperature Audit Device |  | Annually | Verify against NIST standard or sent to factory
Pressure Audit Device |  | Annually | Verify against NIST standard or sent to factory
Comments Section for Part 2 MQOs
Part 3 - Drawing
Briefly draw the monitoring location and illustrate all obstructions including distances to
the nearest roadways and/or obstructions.
After your sketch, please photograph the sampler from 8 cardinal directions, and then take
photographs looking from the sampler in the 8 directions.
N
E
S
W
MetOne Sampler
4.6 m
R&P Seq, Sampler
Black top road is 25 m away
X
Anenometer Tower (T, RH, Wind, Pressure)
15 m

## CSN SASS Worksheet
Chemical Speciation Network |  |  |  |  | US Environmental Protection Agency
Performance Audit Worksheet |  |  |  |  | Office of Air and Radiation
MetOne SASS
Location |  |  |  |  | Date
AQS Site ID |  |  |  |  | Latitude:
Longitude:
Audit Information
Auditor(s) |  |  |  |  | Affiliation
Operator |  |  |  |  | Affiliation
Phone No.
Sampler Model |  |  |  |  | Sampler S/N
Head S/N
Pump S/N
Last Calibration Date |  |  |  |  | Collocated? | Yes
No
Reference Std Model |  |  |  |  | Reference Standard S/N
Calibration Date
Clock Test:
If Local Time is under daylight savings, convert Ref Std to Local Standard Time. Daylight Saving Time begins for most of the United States at 2:00 a.m. on the first Sunday of April. Time reverts to standard time at 2:00 a.m. on the last Sunday of October.
| Time (hh:mm) |  |  |  | Difference Minutes |  | 5 minutes or less?
| Ref Std |  | SASS |  |  |  | Pass | Fail
Audit
Recalibrated
Date
Leak Test
| Initial Audit |  |  | After Correction |  | 0.10 L/min or greater fails
| A L/min |  |  | B L/min |  | Fail A | Fail B | Pass
Channel 1 |  |  | Channel 1
Channel 2 |  |  | Channel 2
Channel 3 |  |  | Channel 3
Channel 4 |  |  | Channel 4
Channel 5 |  |  | Channel 5
Channel 6 |  |  | Channel 6
Channel 7 |  |  | Channel 7
Channel 8 |  |  | Channel 8
Flow Test |  |  |  | Calibration
For the reference standard, enter "UR" for under range and "OR" for over range flow readings.
|  | L/min |  |  |  |  | Less than 10%?
| Ref Std |  | SASS |  | % Difference |  | Pass | Fail
Channel 1
Channel 2
Channel 3
Channel 4
Channel 5
Channel 6
Channel 7
Channel 8
Retest after Calibration
|  | L/min |  |  |  |  | Less than 10%?
| Ref Std |  | SASS |  | % Difference |  | Pass | Fail
Channel 1
Channel 2
Channel 3
Channel 4
Channel 5
Channel 6
Channel 7
Channel 8
Reference Standard vs Design Flow
|  | L/min |  |  |  |  | Less than 10%?
| Ref Std |  | SASS |  | % Difference |  | Pass | Fail
Channel 1 |  |  | 6.7
Channel 2 |  |  | 6.7
Channel 3 |  |  | 6.7
Channel 4 |  |  | 6.7
Channel 5 |  |  | 6.7
Channel 6 |  |  | 6.7
Channel 7 |  |  | 6.7
Channel 8 |  |  | 6.7
Retest after Calibration
|  | L/min |  |  |  |  | Less than 10%?
| Ref Std |  | SASS |  | % Difference |  | Pass | Fail
Channel 1 |  |  | 6.7
Channel 2 |  |  | 6.7
Channel 3 |  |  | 6.7
Channel 4 |  |  | 6.7
Channel 5 |  |  | 6.7
Channel 6 |  |  | 6.7
Channel 7 |  |  | 6.7
Channel 8 |  |  | 6.7
Ambient Temperature Test
|  | Degrees C |  |  |  |  | Less than 2 degrees?
| Ref Std |  | SASS |  | Difference |  | Pass | Fail
Retest After Recalibration
Filter Temperature Test
|  | Degrees C |  |  |  |  | Less than 2 degrees?
| Ref Std |  | SASS |  | Difference |  | Pass | Fail
Channel 1
Channel 2
Channel 3
Channel 4
Channel 5
Channel 6
Channel 7
Channel 8
Retest After Recalibration
Channel 1
Channel 2
Channel 3
Channel 4
Channel 5
Channel 6
Channel 7
Channel 8
Pressure Test
|  | mm Hg |  |  |  |  | Less than 10 mm?
| Ref Std |  | SASS |  | Difference |  | Pass | Fail
Retest after recalibration
Please save file as <Location-CSN-Date>where "Location" is the common network site name.