##### United States Response

Contact: Wayne Stiefel (stiefel@nist.gov)

Comments for 1st Working Draft “General Requirements for Software Controlled Measuring Instruments”

| Document Clause | Page Number | Comments | Proposed Change |
| --- | --- | --- | --- |
| 2 | 5 | The D-SW document scope provides no guidance on the matter of “universal” (user programmable) vs. “built for purpose” (upgradeable, but not user programmable) equipment. The WELMEC 7.2 Software Guide document immediately breaks out these two categories, noting the differences between the two. In section 5.2.2, D-SW notes some restrictions on the use of an “off the shelf general purpose computer”. Why does the D-SW document not address this matter more fully? | Define categories and appropriate restrictions in much more detail, similar to WELMEC Guide. |
| Sections 3, 5 and 6 |  | This restates our previous comment on the Pre-draft that the Secretariat noted for future discussion. The term “legally relevant” is found in many places in the document and refers to both hardware and software. Example from 5.1.1: “Legally relevant software shall be clearly identified.” To improve clarity we prefer “under metrological control” | Replace “legally relevant” with “under metrological control”. |
| 3.3.3 | 13 | Editorial - Add the words “an” and “the”. | To set a special protection to serve as an indicator for the case of … |
| 5.1.1 | 15 | We concur with the desirability of a single software identification number for a piece of equipment. We can, however, envision a machine containing multiple measurement channels, each containing metrological software, each having a different purpose and identification. For example, a PLC equipment rack might have a number of empty slots into which a measurement channel card could be inserted, perhaps for measuring flow, mass, temperature, etc. Each card contains metrological software. The equipment rack also has its own software. The customer can configure the equipment rack contents. Can/should a single identification cover all possible configurations? What about the case of “universal” measurement option cards that may be used in multiple types of equipment chassis? Perhaps special cases such as this are treated separately? | Topics for discussion and may require further clarification. |
| 5.1.1 | 15 | Editorial - Add suffix “ant”. | Purpose: Each measuring instrument in use has to conform to the approved type. The software identification enables surveillance personnel and persons affected by the measurement to determine whether the instrument under consideration is conformant. |
| 5.2.1.1 | 18 | Editorial - Remove repeated word “it”. | Requirement (b): It it shall be shown… |
| 5.2.6 | 23 | Clause needs to be added regarding the update of universal operating systems (i.e., Linux, Windows, etc.) that the software may reside on. Some updates are small patches while other updates fix potential security issues. | 5.2.6 Requirement – add the following: …Only the software used for measurement needs to be verified after updating. If the software resides on a universal operating system and the universal operating system is updated, verification of the software is not required. Verification of the software is only required when the software is updated. |
| 5.2.6.1 | 23 | This restates our comment on the Pre-daft. Traced electronic updates appear to require an electronic audit trail entry for the instance of the update; however, non-traced updates in the Figure appear not to require an audit trail. Potentially, non-traced updates might require an audit trail that is manifested as an update to the approval certificate or in essence sort of a paper audit trail. The Secretariat added a sentence “A person responsible for verification must be at place.” We do not understand what this means and it does not address our concern to document the update. | A suggestion might be to indicate that a manual (i.e. paper) re-approval or paper audit trail is necessary for these updates. |
| 6.1 | 26 | In some cases… What cases and what will be done with the source code? Source code is proprietary property of the manufacturer. In lieu of source code, a high level flow diagram provides needed information in conjunction with “black box” testing. The Secretariat in responding to the comment noted that: Added an explanation. Where is this explanation added? We see your point as you also refuse/deny the validation methods that are based on source code analysis. We assume that if high level conformity (5.2.5 (d)) or high level fraud protection is required, the high level validation methods are suitable. We disagree. Fraud protection is not achieved through code inspection but through various sealing methods and interface protections in conjunction with field enforcement. The responsible TCs should decide about the level of validation, we only want to provide the tools for realising this level. We would like to keep the paragraph and all the corresponding requirements in section 6.4. We agree with providing options for TCs however, the application of procedure B should be restricted to the rare instances involving risks to life. Example (relevant today in some countries): Evidential measurements require high trust in the software of the device which can only be attested before the court on the basis of source code analysis. We disagree that evidential measurements can only be attested on the basis of source code analysis. This is insufficient what is required is black box testing of the instrument and software to prove it meets requirements in the laboratory. and further that the field processes including calibration and operational considerations have been achieved. | Delete requirement to provide source code. |
| 6.1 and throughout the document |  | Restating our comment to the pre-draft: We have difficulty understanding the circumstances in which type approval would require Validation procedure B. Data flow analysis (6.4.3.4), Code inspection and Walk through (6.4.3.5) and Software module testing (6.4.3.6) are all beyond what should be required for type approval, and almost certainly beyond the capability of most type approval laboratories. We believe that only the software developer has the full capability to perform these types of evaluations. In lieu of source code, a high level flow diagram provides needed information in conjunction with “black box” testing. The Secretariat asked that we refer to the U.S. FDA’s “General Principles of Software Validation; Final Guidance for Industry and FDA Staff” which considers software validation. We are familiar with the FDA’s Guidance and its intended audience of medical instrument manufacturers. The FDA’s General Principals of Software Validation: Final Guidance for Industry and FDA Staf” provides guidance on what FDA expects medical device manufacturers to do to ensure compliance with the Quality System regulation for software validation. It is the device manufacturer who must apply the guidance in implementing their Quality System not the FDA compliance official The FDA regulation also states: “When source code and design specifications cannot be obtained, ‘black box testing’ must be performed to confirm that the software meets the user’s needs and its intended uses.” Again it is the device manufacturer who performs the ‘black box testing’. | Delete requirements to provide source code and Procedure B throughout the document. |
| 6.1 to 6.4 | 26 to 32 | We have received comments from both users and manufacturers of weighing and measuring instruments that provide additional rationale for not providing source code. A summary follows: Data Flow Analysis, Code Inspection, Walk-Through and Software Module Testing, are of limited value because they do not test for the occurrence of unexpected interactions between seemingly unrelated modules operating together in real time. This type of analysis requires the highest level or specialist technical expertise, and requires access to all the necessary debugging tools and support tools that the examiner is unlikely to possess. Requiring that source code ‘shall’ be made available to the examiner is impractical in some cases. Modern system designs utilize dozens if not hundreds of files and libraries. Quality management tools in use today are also a barrier to distributing copies of source code. These QC tools are designed to restrict access to, and impose version control of software modules. Software development tools such as build scripts, pre-parsers, compilers, editors, assemblers and custom code for final file generation compilers are usually licensed to run on specific machines or network servers and limit the number of licensed users. Many of these tools are licensed per seat and are locked to development system via some licensing mechanism. Potential requirement of source code by type approval examiner may pose problem for non-open-source software. Might be difficult to understand the software as a whole; usually, software is written in parts and the parts interact and thus it is difficult for one person to understand source code for the entire software package. For software examination, performance-based testing of software should be used. Performance-based testing of software enables simulation of operating scenarios. Code inspection, walk-through, and data flow analysis do not enable the examiner to test potential operating conditions. Distributing electronic copies of proprietary source code, measuring system algorithms, cryptographically functions, or critical timing constraints, puts at risk the intellectual property rights of the manufacturer and introduces the potential for legal liability to the testing authority. | Emphasize methods AD, VFTM and VFTSw. Eliminate or de-emphasize methods DFA, CIWT and SMT |