Pressure Relief and Plumbing Features

Cryostat Return-End

1.8K Magnet Test-Facility at LBL

This paper lists the basic functions and system features required of the return-end plumbing in a horizontal cryostat used for 1.8K testing of 1.3 meter superconducting magnets at LBL’s Magnet Test Facility.  Four functions are required.  Their associated features and purposes are listed below:

| Feature | Purpose |
| --- | --- |

Cryostat cool-down gas-return:

| Gas exit: top , return-end | Removal of the warmest gas from the cryostat. |
| --- | --- |
| Warm seals: vacuum-tight | Clean “pump and purging” of the cryostat. |
| Radiation shields: Al, flapper-valve style | Radiation & thermal-acoustic block; R-V conductance. |
| Radiation shields: external removal & repair | Easy repair of design oversight or quench damage. |
| Thermal interception: metered gas-outflow | Controllable cold-gas heat interception. |
| He return valve: cold operation | “Cross-over” (T < 30K) control. |

Cryostat pressure-relief:

| Plumbing: large diameter | Low gas-flow impedance. |
| --- | --- |
| He-II R-V: spring-loaded (10 psi) | Auto-close, auto-pressure-relief. |
| He-II R-V: low super-fluid & heat-leak | Minimize heat leak to cold volume. |
| He-II R-V: manual, lever-controlled | Quick, easy over-ride of the auto-close feature |
| psi R-V: automatic | Safety: baseline pressure relief. |
| psi R-V: automatic, flat-plate | Safety: back-up, high conductance pressure relief. |
| psi R-V: double-captured O-ring | Reliable resealing. |
| Manual R-V: easy access, 2 inch ball valve | Quick insurance against frozen R-V seals. |
| All valve seals: Teflon | Cold seal-able, no cracking, no sticking. |
| All valve bodies: copper or brass | Facilitate torch resealing of leaking seals. |
| All valves: downward venting | Safety: no water condensation failures, easy reach. |
| All valves: vent deflectors | Safety: freezing gas away from people or equipment. |

Liquid helium removal:

| Insertion valve: warm, ball valve | Quick insertion capability. |
| --- | --- |
| Insertion valve seals: Teflon | Cold operation capability. |
| Insertion seal: beveled Teflon | Reliable, cold- sealing capability. |
| Insertion seal location: 2 inches above valve | Safety, permits sealing before opening the valve. |
| Super-fluid seal: Teflon plug on end of tube | Easy, inexpensive, removable, replaceable seal. |
| Radiation shield: Nema-G spaced Al rods | Easy, inexpensive radiation blockage. |
| Heat intercept: gauged gas-outflow | Controllable heat interception. |

Diagnostic:

| Return pressure: Bourdon, 30” - 15 psi | Facilitate valve operation decisions. |
| --- | --- |
| Cryostat pressure: Bourdon, 30” - 15 psi | Facilitate valve operation decisions. |
| Cryostat transient pressure: electronic | Permit remote recording. |
| R-V heat intercept flowmeter: 20 CFH | Facilitate controlled heat-leak. |
| LHe, F/O heat intercept flowmeter: 5 CFH | Facilitate controlled heat-leak. |
| All gauge mounts: warm stems, upward | Avoid excess temperature and/or water condensation. |
| All gauge locations: away from R-V’s | Avoid torch damage to glass, plastic, or wiring. |
| All gauge orientations: toward control room | Facilitate visibility to operators. |
| Fiber-optic seals: 2, short,potted-pipe type | Easy sealing/installation/removal of fragile cables. |

Valve Status

Standard Operations

Operation

He-II

R-V

Manual

R-V

He Gas

Return

R-V

Flow

LHe

Flow

LHe

| Access |
| --- |
| Cool-down (300 - 30K) | Open | Closed | Open | 0 | 2 CFH | Rod Plug5 |
| Cool-down (30K - 4.3K | Open | Closed |

Closed1

| sequentially | <10 CFH | <2 CFH | Rod Plug |
| --- | --- | --- | --- |
| 4.3 Liquid Accumulation | Closed | Closed | Closed | 10 CFH | 2 CFH | Rod Plug |
| Quench (P < 20 psi) | Closed | Closed | Closed | Large | N/A | Rod Plug |
| Quench (P > 20 psi) | Open3 | Open4 | Open2 | 0 | N/A | Rod Plug |
| Super-Fluid (T < 2.17K) | Closed | Closed | Closed | 10 CFH | 2 CFH | Rod Plug |
| Liquid Helium Recovery | Closed | Closed | Closed | Closed | Closed | LHe Stinger |

Cross-over (30K to 4.3K):  It is preferable to close a return valve that is warm (as far as practical from the cryostat), in order to reduce the damage to the sealing surfaces and give better control over the valve’s conductance.

Quench Recovery:  He gas-return is opened at the start of the lead-end’s 20 psi R-V blow-off.

Quench Recovery:  He-II R-V is opened to keep the lead-end’s 20 psi R-V from getting too cold.

Quench Recovery:  Manual R-V is opened to keep all “warm” R-V’s from getting too cold.

Rod plug: A rod, having the same diameter as LHe removal stinger, is composed of a series of Nema-spaced Al radiation blocks.  It has a beveled Teflon tip for the bottom super-fluid seal.

Super-Fluid Seals

Reliability and Low Thermal Conductance Issues

| Assertion | Conclusions |
| --- | --- |
| All seals will leak some superfluid. | Expect helium accumulation above the seal. |
| Heat-flow to the “1.8K” system is controlled by the temperature and the flow-rate at which normal helium is driven through the leak. |

Minimize by reducing:

the inlet temperature to the leak,

the mass flow-rate through the leak.

The normal helium flow-rate is governed by the pressure difference across the leak, and the leak’s conductance.

Minimize by reducing:

the pressure difference across the leak,

the conductance through the leak.

A leak’s conductance is minimized by minimizing the leak’s cross-sectional area, maximizing the leak path-length.

Minimize by providing:

mating or auto-matching surfaces.

| a long leak-path. |
| --- |
| Differential contractions can distort surfaces. | Use the same thick, isotropic material for all pieces. |
| It is unlikely to duplicate the exact mechanical conditions after each valve actuation. | Expect the leak-rate to vary with each actuation of a valve. |
| The leak’s inlet temperature is controlled by the system’s ability to stratify helium at the leak’s inlet. |

Arrange the geometry to facilitate stratification:

Intercept any heat-leak at the 4.3K layer.

Feed any leak from the 2.2K layer.

Return-End Plumbing

Schematic Diagram

# Text boxes

10 psi

Auto R-V

Super-fluid

Isolation R-V

P

P

He-Return

Pressure

He-Return

Valve

Flowmeter

P

Optical-Fiber Feed-Thru

Flowmeters

Flapper-Valve

Radiation Shields

Cryostat

Super-fluid Seal

Liquid Helium

Access Line

Liquid Helium

Access Valve

Beveled Teflon

Sliding Safety-Seal

To

Helium Return

Cryostat

Pressure

20 psi back-up R-V

Manual

Cold R-V

Control

Manual R-V