Magnet Cryostat Tasks

The following magnet cryostat tasks have been identified as relevant to the successful, safe, and/or efficient testing of LBQ and D20 magnets at 1.8K, and are listed in rough order of immediate concern.  Among various options that have been examined, the option which currently seems the wisest is listed, along with the reason for the task.  Each task has been assigned a rough priority (P) for implementation. The priority groups have the following meaning:

Essential for safe, reliable, successful operation of LBQ at 1.8K.

Presently deferrable, but essential for safe, reliable operation of D20 at 1.8K.

Sanity/time/cost saving, and quality enhancements.

Disaster insurance.

Non-essential simplifications.

The initials of the primary coordinator (Co) for each task is listed (PB = Paul Bish, JC = Julian Cortella, DS = Dick Schafer, AL =  Al Lietzke, HO = Hans Van Oort).

| Item | Task | Reason | P | Co |
| --- | --- | --- | --- | --- |
| Power-lead insulation: | Surround leads with loose Kapton sheet on ground side. | Safer electrical operation. | 1 | PB |
| Cold R-V sealing surface: | Reverse plate to provide a 3/4” thick SS sealing surface. | Copper gasket causes 30W conduction load to 1.8K. | 1 | PB |
| Fiber-optic tubing feed-thru pipe, lower seal: | Beveled Teflon seal onto 3/4” OD tube (or plug). | The 3/4” Teflon seal is already finished and waiting. | 1 | PB |
| 19-wire V-tap wire copper tubing: | Provide 2” long silicone-rubber sealed Teflon or Nema plug. | >10W conduction to 1.8K. | 1 | PB |
| Power-lead feed-thru from 4.3K: | Provide 3” of silicone-rubber sealed, clamped Nema tube (>1/8” wall) attached to the present Teflon seal. | >4.8W/lead conduction load to 1.8K (RRR = 40). | 1 | PB |
| R-V riser thermal-acoustic & radiation baffles: | Install 4-5 “Schafer” butterfly baffles with irregular spacing. | 2W radiation & 2W thermal acoustic load to 1.8K. | 1 | DS |
| Holes in LN shields: | Plug with Al tape or foil. | 3-4W radiation load to 1.8K blankets. | 1 | PB |
| Cryostat diagnostics: | Install diagnostics. | Diagnostic capability. | 1 | PB |
| Liquid helium removal warm valve: | Provide ball-valve with upper tapered Teflon seal at least 2” above valve. | Safer, easier insertion & removal of stinger or plug. | 1 | PB |
| South-end cold R-V actuator rod threads: | Chase, and lubricate threads on cold R-V actuator rod. | Reduce threat of seizure. | 1 | PB |
| All R-V’s: | Provide H2O condensation protection for all R-V’s having frequent usage. | Safer, more reliable R-V operation and resealing. | 1 | PB |
| Glue-joint, lower bellows seal: | Provide plastic guide-pins, and backing screws. | Safer assembly & disassembly of blind glue-joints. | 1 | PB |
| North-end cold R-V spring: | Provide 10 psi capability. |

Safer automatic operation.

| Easier manual operation. | 1 | JC |
| --- | --- | --- |
| South-end cold R-V spring: | Provide 10 psi capability. |

Safer automatic operation.

| Easier manual operation. | 2 | JC |
| --- | --- | --- |
| Fiber-optic feed-thru sealing pipe: | Replace 1-long epoxy-filled pipe with 2 short pipes. | Easier and safer magnet insertion and removal. | 2 | HO |
| Fiber-optic tube super-fluid seals: | Provide super-fluid seals between well-spaced 1.8K-2.2K contours. | 0.086W/tube-in. conduction load on exposed 1/8” tubing. | 2 | HO |
| Fiber-optic tube warm seals: | Provide warm helium seals on all feed-thrus. | Reduce helium loss thru tubing. | 2 | HO |
| Cryostat diagnostics: | Warm-check data-acquisition capability | Catch errors before cool-down. | 2 | AL |
| Liquid helium displacer: | Provide safe, low enthalpy, low helium absorption displacer. | Reduce cool-down and LHe accumulation time. | 3 | JC |
| Manually operated, warm R-V: | Provide a 2” Teflon-sealed ball-valve that can be operated quickly. | Safer operation & more reliable sealing of automatic R-V’s. | 3 | PB |
| Cold R-V manual actuation rod: | Provide for external, threaded insertion. | Easier modification and/or repair. | 3 | DS |
| South-end cold R-V shim: | Split brass spacer tube. | Easier super-fluid sealing. | 3 | PB |
| Glue-joint (LL1 - LL2): | Provide guide-pins and backing screws. | Safer assembly & disassembly of glue joints. | 3 | PB |
| V-tap wires: | Route away from power leads. | Reduce inductive pick-up on V-tap wires. | 3 | AL |
| LN blanket, top ring: | Cut ID smaller to help old shield. | Reduce LL1 load; keep O-ring warm. | 3 | PB |
| Local pressure gauges: | Add gauges to read local pressure. | Better-informed valve operation. | 3 | PB |
| Super-fluid penetrations: | Provide small gas-flow ports | Super-seal & heat intercept diagnostic access. | 3 | PB |
| Local flow-meters: | Add leak-rate flow-meters. | Super-seal & heat intercept diagnostic. | 3 | PB |
| R-V plumbing: | Use threaded pipe, where practical. | Cheaper, easier mods and repairs. | 3 | PB |
| Vacuum gate-valve: | Provide RGA port on pump side of gate-valve. | Permit RGA observation of pump baseline. | 4 | PB |
| 19 V-Tap wire tube shield: | Wrap Al foil and shield copper tube separately, down to start of SS. | Reduce heat load to 1.8K. | 4 | PB |
| Vacuum gate-valve: | Provide back-up valve. | Cheap insurance against bad vacuum. | 4 | PB |
| Flat-plate R-V, O-ring groove: | Provide for a double-captured/glued silicone-rubber O-ring seal. | Decrease chances of O-ring ejection during R-V operation. | 4 | PB |
| V-Tap wires: | Provide with Teflon sheath. | Extra insulation insurance | 4 | PB |
| Liquid helium removal stinger insertion-stop: | Provide Aero-seal clamp stop on stinger. | Protect super-fluid sealing surface. | 4 | PB |
| Liquid helium removal superfluid seal: | Provide Teflon or SS super-seal on end of a warm-sealed stick. | Back-up super-seal | 4 | PB |
| Flat plate R-V plate: | Use copper plate, instead of SS. | Easier resealing of frozen O-ring. | 4 | PB |
| Fiber-Optic feed-thru: | Bevel lower end of 1/2” bevel-seal. | Safer, easier feed-thru insertion. | 4 | PB |
| Cold R-V sealing plate: | Use a flat SS R-V plate (no seal-rings). | 1.5W conduction thru sealing rings to 1.8K. | 5 | PB |
| All cold valve-stems: | Disconnect (or provide 80K intercept). | 200mW (100mW) conduction down stem to 1.8K. | 5 | PB |
| LHe valve-stem pipe: | Cut pipe above the LN blanket, not below. | 40mW radiation-load to 1.8K. | 5 | PB |
| S-I between LL1 & LL2: | Remove S-I between cold systems. | Not needed | 5 | PB |