STANDARD OPERATING PROCEDURE  - 2009

## WASHER/STERILIZER, OPERATION AND MAINTENANCE

PURPOSE: To establish a standard operating procedure for the use of washer/sterilizers.

POLICY: Washer/sterilizers are used in the cleaning/disinfecting process for heat and moisture tolerant instruments in the hospital decontamination area.

RESPONSIBILITY: All SPD personnel or staff engaged in decontamination type activities.

INFORMATION: Washer/sterilizers are effective devices for lowering bioburden on surgical instruments. A washer/sterilizer cycle is 250-270 degrees F. for 4-30 minutes, depending on the manufacturer.  They are designed to eliminate bioburden, therefore, it is recommended that instruments be passed directly from the washer/sterilizer to the preparation room.

PROCEDURE:

1.  Instruments are removed, inspected, and sorted according to the cleaning need upon return from the using area.

2.  Items that will tolerate heat and immersion will be cleaned using the following process:

Step 1 - Instruments are opened and placed in the ultrasonic cleaner.

Step 2 – Upon completion, instruments are placed in a double door washer/sterilizer to reduce bioburden. to the lowest possible level.

Step 3 – Instruments are removed on the Preparation side for inspection and assembly.

3.  Items requiring manual cleaning (those instruments that cannot be cleaned by immersion and agitation, items with lumens, pockets, rubber or plastic, delicate instruments without special containers, and items that cannot be immersed) must not be placed in the washer/sterilizer.

4.  As with other sterilization devices, charts or graphs are monitored to assure that cycle parameters are met.

5.  A schedule is followed for enzymatic cleaner replenishment and basic maintenance of the washer sterilizer.

Reference:  Manufacturer’s User Manual