Spallation Neutron Source Beam Loss Monitor System(

D. GASSNER*; R. WITKOVER**; P. CAMERON*; J. POWER***

BNL, Upton, NY 11973,**TechSource, Inc., Santa Fe, NM 87594,***LANL, Los Alamos, NM 87545

Abstract. The Spallation Neutron Source facility to be built at ORNL is designed to accumulate 2x1014 protons at 1.0 GeV and deliver them to the experimental target in one bunch at 60 Hz.  To achieve this goal and protect the machine from excessive radiation activation, an uncontrolled loss criteria of 1 part in 104  (1 W/m) has been specified.  Measured losses will be conditioned to provide machine tuning data, a beam abort trigger, and logging of loss history.  The design of the distributed loss monitor system utilizing argon-filled glass ionization chambers and scintillator-photomultipliers will be presented.

# BACKGROUND

The SNS project1 is a collaboration of 6 national laboratories to build an accelerator based Spallation Neutron Source at Oak Ridge National Laboratory (ORNL). The H-minus ion source, RFQ, LEBT and MEBT will be provided by Lawrence Berkeley Laboratory (LBL), the Linac will be designed and built by Los Alamos National Laboratory (LANL) and Thomas Jefferson National Laboratory (TJNL) and Brookhaven National Laboratory will design and build the HEBT, accumulator ring and Ring-to-Target Beam Transport (RTBT) line. The Spallation Target and conventional facilities will be the responsibility of ORNL and Argonne National Laboratory. With such a segmented work force it was clear that a consolidated group should handle all controls.  It was also realized that beam instrumentation should be similarly organized so that the same detectors and interface hardware should be used, where ever possible, to measure the same parameters throughout the facility. This should eliminate questions of differences in measurements due to different instrumentation design. This consolidation has been done for beam loss, beam position and beam current monitoring, with BNL having primary responsibility for the Loss Monitor System.

The SNS (Fig. 1) will consist of an H-minus ion source injecting a beam, chopped to pre-form the Ring bunch, into a 2.5 MeV RFQ. This beam will be accelerated to 87 MeV in a Drift Tube Linac (DTL), then to 186 MeV in a Cavity Coupled Linac (CCL). A Superconducting RF Linac (SRF) raises the beam energy to 1 GeV, after which it is transported to the Ring via the HEBT line.

Figure 1.  Drawing of SNS site layout

The Ring accumulates the beam in a single bunch over a 1 msec macro-pulse duration. It is then ejected into the RTBT and transported to the liquid mercury neutron target. The beam from the source through the Linac and Ring will have the same temporal character (1 msec macro-pulse width, chopped to 695 nsec at the Ring revolution frequency of 1.05 MHz).  Beam current will vary from a peak of 70 mA at the source, to 56 mA through the Linac, while in the Ring it will increase from the 56 mA first turn to more than 56 A at the end of the accumulation period, for a total of 2 x 1014 protons in the single bunch.

The design average current of 2 mA will produce 2 MW at the target at a 60 Hz rate. This very high average beam power makes it crucial that losses be kept extremely low to allow normal accelerator maintenance without high background radiation. This will be done by:

- Carefully designing the accelerator to prevent halo formation and having apertures sufficiently large to reduce scraping.

- Providing beam dumps in areas where known losses will occur, such as,

- HEBT where stripped H-minus beam from the Linac must be removed

- Ring injection where H-minus beam which misses the stripping foil and beam which is only partially stripped can be safely removed.

- Installing collimators to clean the beam tails, designed to fully contain the resulting radiation.

- Installing a Beam-in-Gap cleaner in the Ring. This fast rising resonant kicker deflects any beam in the gap into the collimators over approximately a 20 turn period. This will remove a possible source of instability in the Ring and eliminate beam which might otherwise be lost on the extraction equipment.

- Installing a beam loss monitor system to inhibit the beam when high loss is observed.

Beam losses can be divided into “Controlled” and Un-Controlled” losses. Losses in the dumps or collimators which produce no increase in background are considered Controlled Losses and are acceptable as long as they do not exceed the design loss rate. The Beam Loss Monitor (BLM) system will be designed to minimize the Un-Controlled losses by providing data for tuning the machine and by inhibiting the beam when excessive losses occur. It will be a crucial tool in reaching the goal of low losses.

# SYSTEM PHILOSOPHY AND DESIGN

The very high intensity of the SNS will make the BLM system the primary diagnostic tool for tuning and preventing excessive activation and equipment damage. As the primary diagnostic tool there must be sufficient detector coverage in all sections of the SNS. For many years the AGS accelerator complex has used extended ion chambers made from hollow large diameter coaxial cables to assure complete spatial coverage2. These "line" detectors also offer the advantage of 1/r spatial dependence rather than the 1/r2 of point detectors. However, their length prevents them from being placed close to the beam line, lowering their sensitivity, and mechanical considerations preclude high bias voltage, limiting their linear range. For these reasons smaller volume sealed glass ion chambers were used in RHIC3 and will be used in SNS. To obtain uniform coverage similar to that of the long coaxial cable ion chambers, many more detectors are required. In SNS detectors will be placed at essentially every quadrupole, since that is where the beta-max will occur, and at other key points. In addition to these fixed ion chamber locations, there will be other detectors which can be moved to points requiring more detailed study to reduce losses.

Because the ion chambers will not have the time response to see losses within the bunch, scintillator-photomultipliers will be installed at Ring injection and extraction, and other strategic points.

The system must have enough sensitivity to observe low level losses over many pulses at all significant locations, as well as high local losses of short duration. This is further complicated by the 3 decade range of intensity as the Ring accumulates beam. A high level loss can cause physical damage if allowed to continue for several pulses. Quenching of the super-conducting cavities in the Linac can also occur from a rapid local loss. To prevent excessive activation, the Linac, transfer lines and Ring are all limited to a 1 W/m loss over a 1 second duration. This is equivalent to a loss of 10-4 of the full intensity per pulse summed over the entire Ring.

These considerations lead to multiple requirements:

- In all regions there must be a real time limit imposed on local losses. This will be implemented in hardware using a programmable threshold detector and will shut off the beam at the Front End for the remainder of the pulse within 5-10 (sec.

- For each detector a 1-second moving sum will be computed and compared against the 1 W/m limit. This will be done in higher level code and used to warn the operators.

- A similar computation will be required for the 10-4 criterion in the Ring but in this case the sum must include all detectors in the array.

The amount of data generated by the BLM system with several hundred detectors and a pulse rate of 60 Hz is far more than an operator can accept. It is expected that comfort displays of a subset of the data will be always presented while other more detailed data will be available on request. For example, a "waterfall" style presentation4 of the 1-second sum losses in each section of SNS would at a glance give the loss levels as color variations with the x-axis the detector location and the y-axis the history of previous 1-second sums. Another comfort display might be a pseudo-time chart of the 1-second sum losses for the various Linac sections, HEBT, Ring and RTBT. Displays on request might include the output of a single detector at a function of time in the macro-pulse or within the mini-pulse. Another on-request display might be a waterfall display of all the detectors in a given section but with the y-axis now on a pulse-by-pulse basis. This could be useful in looking for line-frequency dependent loss mechanisms.

# EXPECTED LOSSES

The beam loss design criterion has been set at 1 W/m to allow hands-on accelerator maintenance. This has come about from operating experience at LAMPF and the PSR, and design studies for the Accelerator Production of Tritium (APT) project at LANL5.  Applying an average loss criterion to a specific local geometry raises questions. Some losses, e.g., background gas stripping of the H-minus beam, do occur uniformly. Shafer6 has estimated the gas scattering losses at 0.13 W/m for 10-7 Torr vacuum. In the Ring, where the vacuum is much better and the H-minus beam has been converted to protons, this is not a factor.  Most other losses will take place at specific locations. For example: off-energy beam will be lost at high dispersion points, while transverse losses will take place at collimators or quadrupoles where the beta-max occurs. Simulations and activation measurements at LANL7 confirm that losses are most significant downstream of quadrupoles. In the Ring, if the same 1 W/m loss is distributed among the number of quads in the Ring rather than uniformly, the loss would be 4.43 W/quad. This is a significant difference for the BLM system design.

The SNS CDR8 correlates the 1 W/m loss with a 115 Rem/hr dose rate at 28 cm from the beam pipe center. If the detector is mounted on the beam pipe (13 cm from the pipe center) then the dose rate is 533 Rem/hr which is equal to 0.15 Rad/sec or 2.5  mRad/pulse. However, this is based on the 1 W/m average loss. If we mount the detectors downstream of the quadrupoles then we can apply the geometry factor of 4.43 to get 11 mRad/pulse, or, 11 Rad/sec during the 1 msec pulse. While this provides an estimate of the losses for the Linac, HEBT, and RTBT, the Ring losses have a different character since the beam current is continually growing during accumulation. The 1 W/m loss tolerance corresponds to a loss of 10-4 of the total beam injected. But, in the Ring the same particles may make many turns and the loss rate must be much less. To achieve a total loss of 1 W/m the loss mid-way during the pulse must be the same as the average loss in the Linac and HEBT, but the first turn loss must be about 500 times lower, assuming the loss is linear with intensity. It is also clear that both instantaneous and integrated losses must be tested to provide proper beam inhibit protection.

# DETECTORS

## Ion Chambers

The SNS will use about 300 (Linac 94, HEBT 45, Ring 96, RTBT 57) sealed, argon-filled glass ion chambers as the primary detectors for monitoring beam losses.  Among the alternative detectors considered were long coax ion chambers, air filled ion chambers and solid state detectors. The long coax cable ion chamber used at AGS, and ISIS9. Although they do not suffer from the 1/r2 response dependency due to their length, they do have reduced dynamic range due to limited bias voltage potential they can achieve. Argon filled chambers are used in preference to the simpler air filled option10 to take advantage of their faster (electron collection) time response, and better saturation characteristics. Solid state LM’s are not used because they utilize PIN photodiodes which have radiation damage thresholds at about 108 rads. The maximum count rate of these detectors may limit the response during the 1 msec cycle.

The SNS detector consists of a sealed glass ion chamber housed in an aluminum cylinder with signal connectors on one end and HV connectors on the other.  The basic ion chamber11 was designed by R. Shafer at FNAL in 1982 and subsequently packaged by various FNAL engineers to the configuration used at the Tevatron.  This design was further developed by BNL for use in the RHIC.  These modifications include improving radiation hardness by replacing the PTFE insulation in the BNC (signal) and SHV (HV bias) connectors with Rexolite, breaking a ground loop by using isolated versions of these connectors, and improvements to the mechanical supports of the glass chamber within its aluminum case.

Ion Chamber Parameters:

- Net internal volume is 113 cc

- Filled to 725 Torr with pure Argon gas

- 4” long concentric cylinder electrodes made of nickel

  - Inner signal electrode diameter 0.25”

  - Outer HV electrode diameter 1.5”

- Calculated response is 70 nano-Coulombs per Rad or 19.1 pA/Rad/hr

- Measured response is 19.6 pA/Rad/hr with 95% within (1.5 pA/Rad/hr

### Detector Assembly

The ion chamber bottle is manufactured by Troy-Onic12 (~ $120 each, 1999). It is then acceptance tested at BNL and shipped to a local contractor for the remaining mechanical/electrical assembly (~$200-300 each). An RC network built into each ion chamber housing provides decoupling in the case of a short, noise filtering, and storage capability to handle large fast losses.  The guard ring around the glass chamber, covering the signal port end of the outer bias cylinder, reduces leakage currents from the HV to the signal lead as low as 0.1 pA. The double connector configuration on each end allows daisy chain capability of signal and HV.

FIGURE 2.  Ion Chamber Assembly

The detectors may be mounted directly (electrically isolated) to the beam pipe, magnet, or moved away from the radiation source to reduce signal output or integrate losses over a larger area.

FIGURE 3. Typical BLM installation in RHIC transfer line

Ion Chamber signals will be transmitted on a special low-tribo-electric RG59 equivalent cable (similar to Belden 9054).  The HV bias will be carried on red RG-59, which is daisy chained from detector to detector.  Two HV cables will be used in each group of approximately 32 detectors, picking up alternate BLM’s to provide degraded but continued coverage in the event of a high voltage short.

## Scintillator-Photomultiplier Detectors

Scintillator-photomultipliers (SPMs) offer a wide dynamic range and nano-second response to resolve losses within a single turn. They will be located in critical areas around the ring.  They will be installed around injection, extraction, collimators, downstream end of each straight section, and in the transfer lines.  Details of the fast losses might provide insight about clipping of the bunch ends or losses due to beam in the gap and losses due to instabilities if they occur.  They will not be used as the detectors in the global loss monitoring system for a number of reasons:

- They show significant unit-to-unit variation, thus requiring individual calibration over the full dynamic range, which may vary over time.

- The PMTs and the scintillators are subject to radiation damage.

- The cost per channel is many times greater than for an ion chamber based system.

Though the solid scintillators are more efficient, liquid cell scintillator has the advantage of increased radiation hardness (1 Mega-rad dose to reduce light transmission to 1/e in solids, 5-10 Mega-rad for liquids), this is due to the lack of polymer chain bonds to break in the liquids.  This type of detector, also know as a “paint can”, has been used at several laboratories including Fermilab13. However, in previous home built detectors the leakage of the liquid scintillator has raised safety concerns. The use of these detectors at Fermilab was discontinued due to the difficultly and cost of mixed waste disposal.

Typical detector response is on the order of 100 micro-Coulombs per Rad.  Each SPM signal will be buffered and acquired through a standard 60 Msa/s, 14-bit ADC. In addition, all units in a given location will be selectable through a wideband mux to be read into a fast digitizer of 1 Gsa/s to observe details of losses within the bunch duration.

FIGURE 3.  Drawing and photograph of BicronTM liquid cell scintillator-photomultiplier assembly.

Bicron ™ offers a commercially available liquid cell scintillator/PMT assembly with a “Bicrocell” container made of glass or aluminum14.  These cells have expansion reservoirs containing oxygen free nitrogen to maintain stability and light output.  The liquid cell is enclosed in an aluminum cell housing for mechanical strength, and is sealed with a gasket and to prevent leaks.  A liquid scintillator under consideration is BC-505 (chemical name 1,2,4 -Trimethylbenzene), it offers the highest light output of any liquid (80% anthracene), excellent light transmission (2.5 meters minimum 1/e attenuation) and a high flash point (48o C, 118oF).

# ELECTRONICS

Circuitry for the ion chamber BLM system has not yet been designed but will take advantage of the fast electron signal to provide rapid interrupt in the case of high beam losses. A potential problem will arise from the difference in collection times for the electrons and ions since ions due to early losses will contribute to the signal in the latter portion of the 1 milli-second pulse. If the signal is sensed as a current then the electron signal will be much larger than the ion signal but if integrated over a sufficiently long time, the two should result in equal charge collection. In the Ring and RTBT at least one gain change will be required.

While most of the other beam instrumentation in SNS are expected to interface to a standard 60 Msa/s, 14-bit digitizer, this would seem to be overkill for such low bandwidth signals. Since there will be a large number of BLM channels in each of the electronics locations it might be more reasonable to use an digitization technique more suited to multi-channel acquisition.

Programmable threshold detectors will be included to provide rapid sensing of excessive beam loss and allow turn off of the beam within 5-10 (sec.

The SPMs will be buffered and viewed by the standard 60 Msa/s digitizers for general acquisition, allowing 50-60 samples during the micro-bunch. They will also go to a wideband mux for acquisition on a 1 GSa/s scope for more detailed viewing.

# ACKNOWLEDGMENTS

The authors would like to R. Shafer for providing valuable insight and discussions regarding the topics in this work.

REFERENCES

. Alonso, J., "The Spallation Neutron Source Project", Proceedings of the 1999 PAC, NY, 1999, p574. For current status see http://www.sns.gov

2. Witkover. R. L. ,"Microprocessor Based Beam Loss Monitor System for the AGS", IEEE Trans. Nucl. Sci 26 3313 (1979)

3. Witkover, R. L., Michnoff, R. J. and Geller, J. M., "RHIC Beam Loss Monitor System Initial Operation", Proceedings of the 1999 PAC, NY, 1999, 2247

4. Bai, M., et al., “RHIC Beam Loss Monitor System Commissioning in the Year 2000 Run”,To be published these proceedings.

5. Hardekopf, R. A., "Beam Loss and Activation at LANSCE and SNS", LA-UR-99-6825, Los Alamos National Laboratory, Los Alamos, NM. Sept. 1999

6. Shafer, R. A., "Beam Loss From H- Stripping in the Residual Gas", SNS Technical Note LA-UR-99-2497, May 1999

7. Loc. Cit. Reference 5

8. NSNS Conceptual Design Review, June 1997,  Chapter 3, p97

9. M. A. Clarke-Gayther, et al. “Global Beam loss Monitoring Using Long Ionisation Chambers at ISIS” , 4th European Particle Accelerator Conference, London, England, 1994, p1634

10. Plum, M., and Brown, D., "Response of Air-Filled Ion Chambers to High Intensity Radiation Pulses", Proceeding of the 1993 PAC, p2181

11. Shafer, R. E., et al., “The Tevatron Beam Position and Beam Loss Monitoring Systems”, Proc. The 12th Int’l Conf. High Energy Accel., p609, 1983

12. Troy-Ionics Inc., 88 Dell Avenue, P.O. Box 494, Kenvil, NJ., 07847

13. M. Awschalom, H. Howe, R. Shafer, D. Theriot, “An Inexpensive Beam Loss Monitor for Use at NAL” NAL-TM-274 (10/30/70)

14. Bicron, Saint-Gobain Industrial Ceramics, Inc “Organic Scintillators, Related Materials and Detectors”, 12345 Kinsman Road, Newbury, OH.  (http://www.bicron.com)

# Footnotes

( Work performed under the auspices of the U. S. Dept. of Energy