## Sheet1
Medical - Electrical Equipment Inspection
These columns are entered by the equipment owner | [5 empty] | These columns are entered by the EEI
Total
Manufacturer/Model#/Equipment Name | Model No. /Manufacturer | Equipment Location | Equipment Origin | Equipment ID Number | Cognizant Group/Person | Equipment Reference Number | Inspection Information | Date | Inspector | 2 | Comments
Electromiography | Cadwell Modell XL | MRI Bldg. 560 | COTS | 9104E01303 | Imagining Group | BP 000246 | A | 2007-04-23 | J. Durnan | 1
Electromiography -TECA | TECA Corp model GR -20 | MRI Bldg. 561 | COTS | 66672 | Imagining Group | BP 000245 | A | 2007-04-23 | J. Durnan | 1
Note on A4: This is the equipment manufacturer, model # and name. If built "in-house", enter the name you assign to the equipment. Example1: Hewlett Packard 8752a Network Analyzer Example2: Experimental Sample Temperature Controller This is entered before inspection by the Equipment Owner
Note on D4: This is either COTS (commercial off the shelf) CABS (constructed according to BNL specification) BIHB (built in house by BNL) BBEC (brought by external collaborator) O (other) This is entered before inspection by the Equipment Owner
Note on E4: This is either the commercial serial number BNL inventory number, or group assigned number. If none exists, enter NA This is entered before inspection by the Equipment Owner
Note on F4: Person or group responsible for equipment This is entered before inspection by the Equipment Owner
Note on G4: This is the number on the scan tag. This is entered at time of inspection by the EEI
Note on H4: This is either A (approved Lt Grn) CA (Conditionally Approved Lt Yel) NRTL (NRTL Approved Lt Blu) DNA/NIS (Does not Apply/Not in System Lt Vio) R (Rejected Red) This is entered at time of inspection by the EEI
Note on I4: Date inspected This is entered at time of inspection by the EEI
Note on J4: This is entered at time of inspection by the EEI
Note on C4: This is the equipment location. This is entered before inspection by the Equipment Owner

## Biology-463
Biology Electrical Equipment Inspection
These columns are entered by the equipment owner |  |  |  |  | These columns are entered by the EEI
Total
Manufacturer/Model#/Equipment Name | Equipment Location | Equipment Origin | Equipment ID Number | Cognizant Group/Person | Equipment Reference Number | Inspection Information | Date | Inspector | 38 | Comments
Ultek Ion Pump Unit | Bldg 463 rm 195 | COTS |  | STEM | BP 702 | A | 2007-03-28 | J. Durnan | 1 | Unit in storage- original experimental equipment upgraded
STEM 3 -Electron Microscope Unit | Bldg 463 rm 195 | BIHB |  | STEM | BP 703 | A | 2007-03-28 | J. Durnan | 1 | Electron microscope that was installed by BNL
Varian Ion Pump | Bldg 463 rm 196 | COTS |  | STEM | BP 705 | A | 2007-03-28 | J. Durnan | 1
STEM Data acquisition rack | Bldg 463 rm 196 | BIHB |  | STEM | BP 704 | A | 2007-03-28 | J. Durnan | 1 | BNL built Data acquisition rack composed of COTS
Power Supply | Bldg 463 rm 196 | BIHB |  | STEM |  | R | 2007-03-28 | J. Durnan | 1 | Item is excessed/discarded -power cord cut
Temperature Indicator | Bldg 463 rm 196 | BIHB |  | STEM |  | R | 2007-03-28 | J. Durnan | 1 | Item is excessed/discarded -power cord cut
Power Supply | Bldg 463 rm 196 | BIHB |  | STEM |  | R | 2007-03-28 | J. Durnan | 1 | Item is excessed/discarded -power cord cut
Controller box (fast scan controller) | Bldg 463 Lab 195 | BIHB |  | STEM |  | R | 2007-03-28 | J. Durnan | 1 | Item is excessed/discarded -power cord cut
Flash controller | Bldg 463 Lab 196 | BIHB |  | STEM |  | R | 2007-03-28 | J. Durnan | 1 | Item is excessed/discarded -power cord cut
J. T.'s Solid state Unidirectional Pulse switch | Bldg 463 Lab B-239 | BIHB |  | J. S. Biophysics | BP 000209 | R | 2007-03-28 | J. Durnan | 1 | Require Solid ground to chasis and fuse
Varian Excitation Monochronometer | Bldg 463 Lab B-239 | COTS-modified |  | J. S. Biophysics | BP 000210 | CA | 2007-03-28 | J. Durnan | 1 | Modification -removed light source for use at NSLS; if returned to service require inspection
Pre-amp Power Supply | Bldg 463 Lab B-239 | BHIB |  | J. S. Biophysics | BP 000296 | A | 2007-03-28 | J. Durnan | 1
Controller box (fast scan controller) | Bldg 463 Lab B-239 | BHIB |  | J. S. Biophysics | BP 000297 | R | 2007-03-28 | J. Durnan | 1 | Cord cut -deciding whether to repair or discard
Sample Chamber -Spectrometer | Bldg 463 Lab B-239 | BHIB |  | J. S. Biophysics | BP 000298 | DNA | 2007-03-28 | J. Durnan | 1 | Not electrical
Varian Spectrometer | Bldg 463 Lab B-239 | BHIB |  | J. S. Biophysics | BP 000299 | A | 2007-03-28 | J. Durnan | 1 | Computer interface for manual spectrometer
IR light detector | Bldg 463 Lab B-239 | BHIB |  | J. S. Biophysics |  | DNA/NIS | 2007-03-28 | J. Durnan | 1 | Not electrical
voltage Regulator | Bldg 463 Lab B-132 | BHIB |  | Electron Microscopy | BP 000247 | R | 2007-04-26 | J. Durnan | 1 | Replace dry rotted line cord
Controller box | Bldg 463 Lab B-136 | BHIB |  | Southerland Group | BP 000248 | A | 2007-04-26 | J. Durnan | 1
Power Supply | Bldg 463 Lab B-136 | BHIB |  | Southerland Group | BP 000249 | A | 2007-04-26 | J. Durnan | 1
Distiller relay box | Bldg 463 Lab B-136 | BHIB |  | Southerland Group | BP 000250 | A | 2007-04-26 | J. Durnan | 1
Controller box timer | Bldg 463 Lab B-136 | BHIB |  | Southerland Group | BP 000251 | CA | 2007-04-26 | J. Durnan | 1 | Top panel requires bonding wire
Controller box timer | Bldg 463 Lab B-136 | BHIB |  | Southerland Group | BP 000252 | CA | 2007-04-26 | J. Durnan | 1 | Top panel requires bonding wire
Water distiller | Bldg 463 Lab B-136 | BHIB |  | Southerland Group | BP 000253 | A | 2007-04-26 | J. Durnan | 1
Control box -3 dial | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000254 | DNA | 2007-04-26 | J. Durnan | 1
Power Supply switched | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000255 | R | 2007-04-26 | J. Durnan | 1 | Line cord dry rotted
Power Supply -EG&G Photo diode | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000336 | A | 2007-04-26 | J. Durnan | 1
Power Supply -Polarizing Voltage Adj. | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000337 | R | 2007-04-26 | J. Durnan | 1 | Neutral fused
Power Supply -Amp Bal | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000338 | R | 2007-04-26 | J. Durnan | 1 | no cover -exposed terminals
Power Supply | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000339 | R | 2007-04-26 | J. Durnan | 1 | Chassis not grounded; no fuse for small wire
Power Supply - Intronics | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000340 | R | 2007-04-26 | J. Durnan | 1 | Case bond wire to small; no fuse
Power Supply for IO-57 | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000341 | R | 2007-04-26 | J. Durnan | 1 | Chassis not grounded; no fuse for small wire
Control box | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000342 | DNA | 2007-04-26 | J. Durnan | 1
Rotating Arm | Bldg 463 Lab 221A | BHIB |  | Dunn Group | BP 000343 | CA | 2007-04-26 | J. Durnan | 1 | Should have fuse for motor
Electric Shutter System | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000346 | DNA | 2007-04-26 | J. Durnan | 1
Power Supply -line receiver to Xenon | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000347 | R | 2007-04-26 | J. Durnan | 1 | Chassis not bonded
Electric Shutter System | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000348 | DNA | 2007-04-26 | J. Durnan | 1
Controller box | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000349 | DNA | 2007-04-26 | J. Durnan | 1
Line receiver | Bldg 463 Lab 220 | BHIB |  | Hind Group | BP 000350 | DNA | 2007-04-26 | J. Durnan | 1
Note on A4: This is the equipment manufacturer, model # and name. If built "in-house", enter the name you assign to the equipment. Example1: Hewlett Packard 8752a Network Analyzer Example2: Experimental Sample Temperature Controller This is entered before inspection by the Equipment Owner
Note on C4: This is either COTS (commercial off the shelf) CABS (constructed according to BNL specification) BIHB (built in house by BNL) BBEC (brought by external collaborator) O (other) This is entered before inspection by the Equipment Owner
Note on D4: This is either the commercial serial number BNL inventory number, or group assigned number. If none exists, enter NA This is entered before inspection by the Equipment Owner
Note on E4: Person or group responsible for equipment This is entered before inspection by the Equipment Owner
Note on F4: This is the number on the scan tag. This is entered at time of inspection by the EEI
Note on G4: This is either A (approved Lt Grn) CA (Conditionally Approved Lt Yel) NRTL (NRTL Approved Lt Blu) DNA/NIS (Does not Apply/Not in System Lt Vio) R (Rejected Red) This is entered at time of inspection by the EEI
Note on H4: Date inspected This is entered at time of inspection by the EEI
Note on I4: This is entered at time of inspection by the EEI
Note on B4: This is the equipment location. This is entered before inspection by the Equipment Owner

## CFN-703
Electrical Equipment Inspection Information
These columns are entered by the equipment owner |  |  |  |  | These columns are entered by the EEI
Manufacturer/Model#/Equipment Name | Equipment Location | Equipment Origin | Equipment ID Number | Cognizant Group/Person | Equipment Reference Number | Inspection Information | Date | 77 | Inspector | Comments
Vacuum Pump | Bldg 703--lab W-6 | COTS | SN-305350-2003 | Meigen Aronson | BP 000 000 646 | NRTL | 2007-01-25 | 1 | W Rasmussen
Pressure gauge | Bldg 703--lab W-6 | COTS | N/A | Meigen Aronson | BP 000 000 647 | A | 2007-01-25 | 1 | W Rasmussen
Arc Melting | Bldg 703--lab W-6 | COTS | SN-040004 | Meigen Aronson | BP 000 000 648 | A | 2007-01-25 | 1 | W Rasmussen
Elma Tech Welder | Bldg 703--lab W-6 | COTS | SN-4004000010 | Meigen Aronson | BP 000 000 649 | CA | 2007-01-25 | 1 | W Rasmussen | Not wired up yet
Glove Box | Bldg 703--lab W-6 | COTS | BNL-131197 | Meigen Aronson | BP 000 000 650 | NRTL | 2007-01-25 | 1 | W Rasmussen
Edwards vacuum pump | Bldg 703--lab W-6 | COTS | SN-23813 | Meigen Aronson | BP 000 000 651 | A | 2007-01-25 | 1 | W Rasmussen
Weller soldering iron | Bldg 703--lab W-6 | COTS | WES51 | Meigen Aronson | BP 000 000 652 | NRTL | 2007-01-25 | 1 | W Rasmussen
Ohaus Scale | Bldg 703--lab W-6 | COTS | SN-1125520979 | Meigen Aronson | BP 000 000 653 | NRTL | 2007-01-25 | 1 | W Rasmussen
Lamp | Bldg 703--lab W-6 | COTS | SN-E78751 | Meigen Aronson | BP 000 000 654 | NRTL | 2007-01-25 | 1 | W Rasmussen
Lamp Cambridge instruments | Bldg 703--lab W-6 | COTS | Cat #31-35-28 | Meigen Aronson | BP 000 000 655 | NRTL | 2007-01-25 | 1 | W Rasmussen
Bransonic Ultrasonic cleaner | Bldg 703--lab W-6 | COTS | SN-RLCO10494887D | Meigen Aronson | BP 000 000 656 | NRTL | 2007-01-25 | 1 | W Rasmussen
Buch Rotavapor | Bldg 703--lab W-6 | COTS | SN-4117130450008 | Meigen Aronson | BP 000 000 657 | NRTL | 2007-01-25 | 1 | W Rasmussen
Portable Ohaus Scale | Bldg 703--lab W-6 | COTS | SN-CA7635 | Meigen Aronson | BP 000 000 658 | NRTL | 2007-01-25 | 1 | W Rasmussen
Bransonic Ultrasonic cleaner | Bldg 703--lab W-6 | COTS | SN-RLC020496915D | Meigen Aronson | BP 000 000 659 | NRTL | 2007-01-25 | 1 | W Rasmussen
Barnstead Mixer | Bldg 703--lab W-6 | COTS | SN-268021176288 | Meigen Aronson | BP 000 000 660 | NRTL | 2007-01-25 | 1 | W Rasmussen
IEC HN-SII Centrifuge | Bldg 703--lab W-6 | COTS | SN-235526275 | Meigen Aronson | BP 000 000 661 | A | 2007-01-26 | 1 | W Rasmussen
Fisher Scientific Hot Plate | Bldg 703--lab W-6 | COTS | SN-56001126 | Meigen Aronson | BP 000 000 662 | A | 2007-01-26 | 1 | W Rasmussen
Fisher Scientific Hot Plate | Bldg 703--lab W-6 | COTS | SN-560910469909 | Meigen Aronson | BP 000 000 663 | A | 2007-01-26 | 1 | W Rasmussen
Welch Vacuum Pump | Bldg 703--lab W-6 | COTS | Manf.# Go37 | Meigen Aronson | BP 000 000 664 | NRTL | 2007-01-26 | 1 | W Rasmussen
Controller for Box Oven | Bldg 703--lab W-6 | COTS | SN-S190502204T0 | Meigen Aronson | BP 000 000 665 | A | 2007-01-26 | 1 | W Rasmussen
Controller for Box Oven | Bldg 703--lab W-6 | COTS | SN-O09050139500 | Meigen Aronson | BP 000 000 666 | A | 2007-01-26 | 1 | W Rasmussen
Lamp Telescoping Arm | Bldg 703--lab W-6 | COTS | N/A | Meigen Aronson | BP 000 000 668 | NRTL | 2007-01-26 | 1 | W Rasmussen
Lindberg Tube Oven | Bldg 703--lab W-6 | COTS | SN-889188 | Meigen Aronson | BP 000 000 669 | A | 2007-01-26 | 1 | W Rasmussen
Vacuum Oven | Bldg 703--lab W-6 | COTS | SN-0501293 | Meigen Aronson | BP 000 000 667 | A | 2007-01-26 | 1 | W Rasmussen
Diamond Wheel Saw | Bldg 703--lab W-6 | COTS | Model 650 | Meigen Aronson | BP 000 000 670 | A | 2007-01-26 | 1 | W Rasmussen
Lake Shore Current Source | Bldg 703--lab W-8 | COTS | SN-33297 | Meigen Aronson | BP 000 000 601 | A | 2007-01-29 | 1 | W Rasmussen
Lock in Amplifier model JB4 | Bldg 703--lab W-8 | COTS | SN-402 | Meigen Aronson | BP 000 000 602 | A | 2007-01-29 | 1 | W Rasmussen
Lake Shore Temp. Controller | Bldg 703--lab W-8 | COTS | Model DCR91CA | Meigen Aronson | BP 000 000 603 | A | 2007-01-29 | 1 | W Rasmussen
Stanford Labs Lock in Amplifier | Bldg 703--lab W-8 | COTS | SN-23868 | Meigen Aronson | BP 000 000 604 | A | 2007-01-29 | 1 | W Rasmussen
Cryo-Magnetics Inc. ISP100 Power sys. | Bldg 703--lab W-8 | COTS | 1050 | Meigen Aronson | BP 000 000 594 | CA | 2007-01-29 | 1 | W Rasmussen | Output connections must be covered
Battery Charger | Bldg 703--lab W-8 | COTS | Model1014C | Meigen Aronson | BP 000 000 605 | A | 2007-01-29 | 1 | W Rasmussen
HP Power Supply Model 6218A | Bldg 703--lab W-8 | COTS | 8M1278 | Meigen Aronson | BP 000 000 606 | A | 2007-01-29 | 1 | W Rasmussen
Adjust-a-volt (Variac) | Bldg 703--lab W-8 | COTS | SN-1703 | Meigen Aronson | BP 000 000 607 | A | 2007-01-29 | 1 | W Rasmussen
Oxford Air pump | Bldg 703--lab W-8 | COTS | Model GFS/VPZ0233 | Meigen Aronson | BP 000 000 608 | A | 2007-01-29 | 1 | W Rasmussen
Air Pump | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 609 | A | 2007-01-29 | 1 | W Rasmussen
Lake Shore Temp Controller Model 91C | Bldg 703--lab W-8 | COTS | SN-16705 | Meigen Aronson | BP 000 000 610 | A | 2007-01-29 | 1 | W Rasmussen
Lock-in Amplifier Model SR830DSP | Bldg 703--lab W-8 | COTS | SN-53702 | Meigen Aronson | BP 000 000 611 | A | 2007-01-30 | 1 | W Rasmussen
Keithly DMM Scanner | Bldg 703--lab W-8 | COTS | SN-485365 | Meigen Aronson | BP 000 000 612 | A | 2007-01-30 | 1 | W Rasmussen
Vacuum Guage | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 613 | A | 2007-01-30 | 1 | W Rasmussen
Lock-in Amplifier Model SR530 | Bldg 703--lab W-8 | COTS | SN-7215 | Meigen Aronson | BP 000 000 614 | A | 2007-01-30 | 1 | W Rasmussen
Precision Impedance Bridge | Bldg 703--lab W-8 | COTS | SN-262810-03 | Meigen Aronson | BP 000 000 615 | A | 2007-01-30 | 1 | W Rasmussen
Kepco Power Supply Model SC-15-4M | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 616 | A | 2007-01-30 | 1 | W Rasmussen
Kepco Power Supply 1Amp- 100Volts | Bldg 703--lab W-8 | COTS | SN-72765 | Meigen Aronson | BP 000 000 620 | CA | 2007-01-30 | 1 | W Rasmussen | Output connections must be covered
Proto Board Model 203A | Bldg 703--lab W-8 | COTS | SN-192786 | Meigen Aronson | BP 000 000 617 | A | 2007-01-30 | 1 | W Rasmussen
Liquid Helium Level Model 12 | Bldg 703--lab W-8 | COTS | SN-C621E | Meigen Aronson | BP 000 000 618 | A | 2007-01-30 | 1 | W Rasmussen
Liquid Helium Level Model 1200SH | Bldg 703--lab W-8 | COTS | SN-C1052E | Meigen Aronson | BP 000 000 619 | A | 2007-01-30 | 1 | W Rasmussen
Weller soldering iron | Bldg 703--lab W-8 | COTS | SN-AL0013750 | Meigen Aronson | BP 000 000 621 | NRTL | 2007-01-30 | 1 | W Rasmussen
Desk Lamp | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 622 | NRTL | 2007-01-31 | 1 | W Rasmussen
Micro Hub/Repeater Model 820T | Bldg 703--lab W-8 | COTS | SN-607050816 | Meigen Aronson | BP 000 000 623 | NRTL | 2007-01-31 | 1 | W Rasmussen
AMS Model SAX 4x4-232 | Bldg 703--lab W-8 | COTS | SN-7298-01 | Meigen Aronson | BP 000 000 624 | A | 2007-01-31 | 1 | W Rasmussen
Switching power supply | Bldg 703--lab W-8 | COTS | SN-093-03-07845 | Meigen Aronson | BP 000 000 625 | A | 2007-01-31 | 1 | W Rasmussen
Keithly Micro volt DMM | Bldg 703--lab W-8 | COTS | SN-0550686 | Meigen Aronson | BP 000 000 626 | A | 2007-01-31 | 1 | W Rasmussen
Granville-Phillips Controller | Bldg 703--lab W-8 | COTS | SN-815 | Meigen Aronson | BP 000 000 627 | A | 2007-01-31 | 1 | W Rasmussen
AVS Resistance bridge | Bldg 703--lab W-8 | COTS | SN-910106 | Meigen Aronson | BP 000 000 628 | A | 2007-01-31 | 1 | W Rasmussen
Keithly Micro volt DMM | Bldg 703--lab W-8 | COTS | SN-482205 | Meigen Aronson | BP 000 000 629 | A | 2007-01-31 | 1 | W Rasmussen
Keithly Micro volt DMM | Bldg 703--lab W-8 | COTS | SN-0550685 | Meigen Aronson | BP 000 000 630 | A | 2007-01-31 | 1 | W Rasmussen
Keithly Micro volt DMM | Bldg 703--lab W-8 | COTS | SN-482208 | Meigen Aronson | BP 000 000 631 | A | 2007-01-31 | 1 | W Rasmussen
Granville-Phillips Controller | Bldg 703--lab W-8 | COTS | SN-7154 | Meigen Aronson | BP 000 000 632 | A | 2007-01-31 | 1 | W Rasmussen
Air Cadet Model 420-1901 | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 633 | NRTL | 2007-01-31 | 1 | W Rasmussen
Fiber Lite | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 634 | A | 2007-01-31 | 1 | W Rasmussen
Nikon Scope Model SMZ2T | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 635 | A | 2007-01-31 | 1 | W Rasmussen
Bell Telephone | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 636 | NRTL | 2007-01-31 | 1 | W Rasmussen
Quantum Design Helium-3 Model 825 | Bldg 703--lab W-8 | COTS | SN-HEQ36 | Meigen Aronson | BP 000 000 637 | CA | 2007-01-31 | 1 | W Rasmussen | Secure covers
Varian controller Model 9699507 | Bldg 703--lab W-8 | COTS | SN-213955 | Meigen Aronson | BP 000 000 638 | NRTL | 2007-01-31 | 1 | W Rasmussen
Edwards vacuum pump | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 640 | A | 2007-01-31 | 1 | W Rasmussen
Heat Gun | Bldg 703--lab W-8 | COTS | SN-1290 | Meigen Aronson | BP 000 000 639 | NRTL | 2007-01-31 | 1 | W Rasmussen
Quantum Design MPMS XL | Bldg 703--lab W-8 | COTS | BNL-A048907 | Meigen Aronson | BP 000 000 641 | A | 2007-01-31 | 1 | W Rasmussen
QuantiumPhysical Property Measurment system | Bldg 703--lab W-8 | COTS | BNL-131164 | Meigen Aronson | BP 000 000 642 | A | 2007-01-31 | 1 | W Rasmussen
Quantum Design MPMS XL | Bldg 703--lab W-8 | COTS | BNL-131163 | Meigen Aronson | BP 000 000 643 | A | 2007-01-31 | 1 | W Rasmussen
Defusion pump | Bldg 703--lab W-8 | COTS | N/A | Meigen Aronson | BP 000 000 644 | CA | 2007-01-31 | 1 | W Rasmussen | Repair wiring
Lindberg/Blue box Oven (1) | Bldg 703--lab W-6 | COTS | N/A | Meigen Aronson | BP 000 000 645 | A | 2007-01-31 | 1 | W Rasmussen
Lindberg/Blue box Oven (2) | Bldg 703--lab W-6 | COTS | N/A | Meigen Aronson | BP 000 000 600 | A | 2007-01-31 | 1 | W Rasmussen
Double variacs in rack | Bldg 703--lab W-5 | COTS | N/A | Meigen Aronson | BP 000 000 599 | R | 2007-01-31 | 1 | W Rasmussen | Cover must be built for panel
Sorenson Power Supply | Bldg 703--lab W-5 | COTS | N/A | Meigen Aronson | BP 000 000 598 | R | 2007-01-31 | 1 | W Rasmussen | AC cord must be replaced
Edwards vacuum pump | Bldg 703--lab W-5 | COTS | SN-037799226 | Meigen Aronson | BP 000 000 597 | NRTL | 2007-01-31 | 1 | W Rasmussen
Vacuum Pump | Bldg 703--lab W-5 | COTS | N/A | Meigen Aronson | BP 000 000 596 | NRTL | 2007-01-31 | 1 | W Rasmussen
Defusion pump | Bldg 703--lab W-5 | COTS | N/A | Meigen Aronson | BP 000 000 595 | R | 2007-01-31 | 1 | W Rasmussen | Needs new plug
Note on A4: NSLS: This is the equipment manufacturer, model # and name. If built "in-house", enter the name you assign to the equipment. Example1: Hewlett Packard 8752a Network Analyzer Example2: Experimental Sample Temperature Controller This is entered before inspection by the Equipment Owner
Note on C4: NSLS: This is either COTS (commercial off the shelf) CABS (constructed according to BNL specification) BIHB (built in house by BNL) BBEC (brought by external collaborator) O (other) This is entered before inspection by the Equipment Owner
Note on D4: NSLS: This is either the commercial serial number BNL inventory number, or group assigned number. If none exists, enter NA This is entered before inspection by the Equipment Owner
Note on E4: NSLS: Person or group responsible for equipment This is entered before inspection by the Equipment Owner
Note on F4: NSLS: This is the number on the scan tag. For the Light Source it starts with LS This is entered at time of inspection by the EEI
Note on G4: NSLS: This is either A (approved Lt Grn) CA (Conditionally Approved Lt Yel) NRTL (NRTL Approved Lt Blu) DNA/NIS (Does not Apply/Not in System Lt Vio) R (Rejected Red) This is entered at time of inspection by the EEI
Note on H4: Date inspected This is entered at time of inspection by the EEI
Note on J4: NSLS: This is entered at time of inspection by the EEI
Note on B4: NSLS: This is the equipment location. This is entered before inspection by the Equipment Owner

## EENS
EENS Electrical Equipment Inspection
These columns are entered by the equipment owner | [5 empty] | These columns are entered by the EEI
Total
Manufacturer/Model#/Equipment Name | Model # | Equipment Location | Equipment Origin | Equipment ID Number | Cognizant Group/Person | Equipment Reference Number | Inspection Information | Date | Inspector | 21 | Comments
SPLAT laser (ESR182) |  | Bldg. 815 Rm 1-25 | BNL Built | ESR -182 | EENS | BOP 000514 | R | 2006-11-17 | J. Durnan | 1
Power Switcher |  | Bldg 815 Lab C-1 | BNL Built | #01 | EENS-ASD | BOP 000517 | A | 2007-01-30 | J. Durnan | 1
High Voltage Module |  | Bldg 815 Lab C-1 | BNL Built |  | EENS-ASD | BOP 000516 | DNA | 2007-01-30 | J. Durnan | 1
Ohmite Variable Transformer | Cat # VT8-F | Bldg. 815 Lab D-6 | BNL Built |  | EENS - DDG | BOP 000687 | CA | 2007-03-12 | J. Durnan | 1 | 2-wire w/ metal frame -if not used for system ground isolation, requires frame grounded
Keithley Picoammeter | 6485 | Bldg. 815 Lab D-6 | BNL Built |  | EENS - DDG | BOP 000688 | A | 2007-03-12 | J. Durnan | 1
Keithley Picoammeter/Voltage Source | 6487 | Bldg. 815 Lab D-6 | BNL Built |  | EENS - DDG | BOP 000689 | A | 2007-03-12 | J. Durnan | 1
Doric Scientific Digital Thermometer | 400 | Bldg. 815 Lab D-6 | COTS |  | EENS - DDG | BOP 000696 | R | 2007-03-12 | J. Durnan | 1 | Exposed terminals because of no rear cover
Omega Digital Thermometer | 2176A | Bldg. 815 Lab D-6 | COTS |  | EENS - DDG | BOP 000697 | A | 2007-03-12 | J. Durnan | 1
National Instruments | N1 PX1-1031 | Bldg. 815 Lab D-6 | NRTL |  | EENS - DDG | BOP 000698 | NRTL | 2007-03-12 | J. Durnan | 1
Lambda Power Supply | LH 121B FM | Bldg. 815 Lab D-6 | COTS |  | EENS - DDG | BOP 000699 | A | 2007-03-12 | J. Durnan | 1
Rack mounted NIM Bin |  | Bldg. 815 Lab D-6 | BNL Built |  | EENS - DDG | BOP 000700 | A | 2007-03-12 | J. Durnan | 1 | Rack Assy. Of low voltage NIM Cards
Astrasan Ultrasonic cleaner | 14HT | Bldg. 815 Lab D-6 | COTS |  | EENS - DDG | BOP 000711 | A | 2007-03-12 | J. Durnan | 1
Soicor Ultrasonic Cleaner | SC-211H | Bldg. 815 Lab D-6 | COTS |  | EENS - DDG | BOP 000712 | A | 2007-03-12 | J. Durnan | 1
Buehler Isomet low speed saw | 11-1280-160 | Bldg. 815 Lab D-6 | COTS |  | EENS - EST | BOP 000713 | A | 2007-03-12 | J. Durnan | 1
MTI Corp. Precision Wire Saw | SXJ2 | Bldg. 815 Lab D-6 | COTS |  | EENS - DDG | BOP 000714 | A | 2007-03-12 | J. Durnan | 1
Omega -Thermo-couple meter |  | Bldg. 815 Lab D-6 | BNL Built |  | EENS - DDG | BOP 000715 | R | 2007-03-12 | J. Durnan | 1 | Needs plate over opening in rear cover to restrict access to live parts
LeCroy 500 Mhz Oscilloscope | Waverunner 6050 | Bldg. 815 Lab D-6 | NRTL |  | EENS - DDG | BOP 000716 | NRTL | 2007-03-12 | J. Durnan | 1
Kenwood Regulated Power Supply | PW 18-1.8Q | Bldg. 815 Lab D-6 | COTS |  | EENS - DDG | BOP 000717 | A | 2007-03-12 | J. Durnan | 1
Gas Purification System | GPS-001 | Bldg. 815 Lab D-6 | BNL Built |  | EENS - DDG | BOP 000718 | A | 2007-03-12 | J. Durnan | 1 | approved system built with commercial equipment.
Detector Box | DTB -001 | Bldg. 815 Lab D-6 | BNL Built |  | EENS - DDG | BOP 000719 | A | 2007-03-12 | J. Durnan | 1
Detector Box | DTB -003 | Bldg. 815 Lab D-6 | BNL Built |  | EENS - DDG | BOP 000720 | A | 2007-03-12 | J. Durnan | 1 | For use with Keithley Picoammeter
Note on A4: This is the equipment manufacturer, model # and name. If built "in-house", enter the name you assign to the equipment. Example1: Hewlett Packard 8752a Network Analyzer Example2: Experimental Sample Temperature Controller This is entered before inspection by the Equipment Owner
Note on D4: This is either COTS (commercial off the shelf) CABS (constructed according to BNL specification) BIHB (built in house by BNL) BBEC (brought by external collaborator) O (other) This is entered before inspection by the Equipment Owner
Note on E4: This is either the commercial serial number BNL inventory number, or group assigned number. If none exists, enter NA This is entered before inspection by the Equipment Owner
Note on F4: Person or group responsible for equipment This is entered before inspection by the Equipment Owner
Note on G4: This is the number on the scan tag. This is entered at time of inspection by the EEI
Note on H4: This is either A (approved Lt Grn) CA (Conditionally Approved Lt Yel) NRTL (NRTL Approved Lt Blu) DNA/NIS (Does not Apply/Not in System Lt Vio) R (Rejected Red) This is entered at time of inspection by the EEI
Note on I4: Date inspected This is entered at time of inspection by the EEI
Note on J4: This is entered at time of inspection by the EEI
Note on C4: This is the equipment location. This is entered before inspection by the Equipment Owner

## Chemistry
Chemistry Department EEI
Total
Manufacturer/Model#/Equipment Name | Building | Equipment Location | Equipment Origin | Equipment ID Number | Cognizant Group/Person | EEI BarCode Number | Inspection Information | Date | Inspector | 73 | Comments
stanford research systems DG535 pulse generator | 555 | Room 327 | COTS | S/N 01023 | Chemistry | BP 000300 | Approved | 2006-06-22 | A. Piper | 1 | N/A
stanford research systems DG535 pulse generator | 555 | Room 327 | COTS | S/N 11689 | Chemistry | BP 000301 | Approved |  | A. Piper | 1 | N/A
Nim Crate Assembly | 555 | Room 327 | COTS/BNL Made | N/A | Chemistry | BP 000302 | Approved |  | A. Piper | 2 | Modules: TTL Pulse switcher; 571 pre-amplifier;SR 245 comp. Interface; 301 DC power supply; 305 DC power supply (2)
Marconi 2022D signal generator | 555 | Room 327 | COTS | S/N 11981/007 | Chemistry | BP 000303 | NRTL/Approved |  | A. Piper | 1 | N/A
Mini-circuits T1A-1000-1R8 amplifier | 555 | Room 327 | COTS | S/N N111601-1 | Chemistry | BP 000304 | Approved |  | A. Piper | 1 | N/A
power supply; 24vdc, 1.5A | 555 | Room 327 | BNL made | N/A | Chemistry | BP 000305 | Approved |  | A. Piper | 1 | Initially had improper fuse in place. This deficiency was corrected.
Uniphase 1201-1 laser power supply | 555 | Room 327 | COTS | S/N 6L4060 | Chemistry | BP 000306 | Approved |  | A. Piper | 1 | N/A
Lecroy LT354 wave analyzer | 555 | Room 327 | COTS | S/N 01483 | Chemistry | BP 000307 | NRTL |  | A. Piper | 1 | N/A
Lecroy 9450 oscilloscope | 555 | Room 327 | COTS | S/N 1317 | Chemistry | BP 000308 | Approved |  | A. Piper | 1 | N/A
Burleigh RC43 programmable generator | 555 | Room 327 | COTS | S/N A6540074 | Chemistry | BP 000309 | Approved |  | A. Piper | 1 | N/A
Burleigh WA1500 wavemeter | 555 | Room 327 | COTS | S/N R7145576 | Chemistry | BP 000310 | Approved |  | A. Piper | 1 | N/A
Melcor thermoelectric chiller | 555 | Room 327 | COTS | S/N KG10520245 | Chemistry | BP 000311 | NRTL |  | A. Piper | 1 | N/A
Tektronix 2235 oscilloscope | 555 | Room 327 | COTS | BNL Barcode 56680 | Chemistry | BP 000312 | NRTL |  | A. Piper | 1 | N/A
Hinds instruments PEM90 modulator | 555 | Room 327 | COTS | S/N E-248-C | Chemistry | BP 000313 | Approved |  | A. Piper | 1 | N/A
Burleigh wavemeter | 555 | Room 327 | COTS | S/N N7081025 | Chemistry | BP 000314 | Approved |  | A. Piper | 1 | N/A
Hitachi VC-6025 oscilloscope | 555 | Room 327 | COTS | S/N 0062227 | Chemistry | BP 000315 | Approved |  | A. Piper | 1 | N/A
Burleigh RG91 ramp generator | 555 | Room 327 | COTS | S/N N1329498 | Chemistry | BP 000316 | Approved |  | A. Piper | 1 | N/A
Stanford research systems DS345 function generator | 555 | Room 327 | COTS | S/N 2888 | Chemistry | BP 000317 | Approved |  | A. Piper | 1 | N/A
Stanford research systems SR830DSP amplifier | 555 | Room 327 | COTS | S/N 52449 | Chemistry | BP 000318 | Approved |  | A. Piper | 1 | N/A
General valve 60-19-900 generator | 555 | Room 327 | COTS | S/N 6L4060 | Chemistry | BP 000319 | Approved |  | A. Piper | 1 | N/A
MKS model 247 amplifier | 555 | Room 327 | COTS | S/N 000663720 | Chemistry | BP 000320 | Approved |  | A. Piper | 1 | N/A
Unlabelled generic "black box" | 555 | Room 327 | BNL built | N/A | Chemistry | BP 000321 | Approved |  | A. Piper | 1 | N/A
Stanford research systems D535 generator | 555 | Room 327 | COTS | S/N -1324 | Chemistry | BP 000322 | Approved |  | A. Piper | 1 | N/A
Marconi instruments 2022D signal generator | 555 | Room 327 | COTS | S/N 119215/034 | Chemistry | BP 000323 | NRTL |  | A. Piper | 1 | N/A
New focus high voltage amplifier model 3211 | 555 | Room 327 | COTS | S/N 144 | Chemistry | BP 000324 | Approved |  | A. Piper | 1 | N/A
Tektronics SC502 oscilloscope | 555 | Room 327 | COTS | S/N B038434 | Chemistry | BP 000325 | Approved |  | A. Piper | 1 | N/A
Lecroy DA1822A differential amplifier | 555 | Room 327 | COTS | S/N LDCB90064911 | Chemistry | BP 000326 | Approved |  | A. Piper | 1 | N/A
Burleigh RG-91 ramp generator | 555 | Room 327 | COTS | S/N 006353 | Chemistry | BP 000327 | Approved |  | A. Piper | 1 | N/A
Coherent MBR E-110 laser amplifier | 555 | Room 327 | COTS | S/N MBRE110/54 | Chemistry | BP 000328 | Approved |  | A. Piper | 1 | N/A
Tektronics TDS310 oscilloscope | 555 | Room 327 | COTS | S/N B040209 | Chemistry | BP 000329 | NRTL |  | A. Piper | 1 | N/A
Neutral Density Control Box | 555 | Lab 261 | BNL built |  | TPR Group | BP 000206 | Approved | 2007-04-10 | J. Durnan | 1
Battery Charger (12 V power supply) | 555 | Lab 261 | BNL built |  | TPR Group | BP 000207 | Approved | 2007-04-10 | J. Durnan | 1
Battery Charger | 555 | Lab 261 | DNA |  | TPR Group | BP 000208 | Approved | 2007-04-10 | J. Durnan | 1 | Fed from NRTL power supply 120/12 V charger
Camac Crate - Control Room Rack | 555 | Lab 261 | BNL built |  | TPR Group | BP 000211 | Approved | 2007-04-10 | J. Durnan | 1 | COTS components
PMT base | 555 | Lab 261 | BNL built |  | TPR Group | BP 000212 | Approved | 2007-04-10 | J. Durnan | 10 | provides power to PMT
Shutter power strip control box | 555 | Lab 261 | BNL built |  | TPR Group | BP 000213 | Approved | 2007-04-10 | J. Durnan | 1 | bonding wire added
Infra-red detector power supply | 555 | Lab 261 | BNL built |  | TPR Group | BP 000214 | Approved | 2007-04-10 | J. Durnan | 1
Calex 5 V power supply | 555 | Lab 261 | BNL built |  | TPR Group | BP 000215 | Rejected | 2007-04-10 | J. Durnan | 1 | No line fuse
Supercritical IR cell Heating jacket controller | 555 | Lab 281 | BNL built |  | TPR Group | BP 000216 | Approved | 2007-04-10 | J. Durnan | 1
Flow Switch Controller box | 555 | Lab 283 | BNL built | SS3-26 | TPR Group | BP 000217 | Rejected | 2007-04-10 | J. Durnan | 1 | Ground & bonding of enclosure by press fit of receptacle. Require bonding wire to case.
Degitect Voltmeter Control box | 555 | Lab 283 | BNL built |  | TPR Group | BP 000218 | Approved | 2007-04-10 | J. Durnan | 1
Shutter controller-computer interface | 555 | Lab 283 | BNL built |  | TPR Group | BP 000219 | Approved | 2007-04-10 | J. Durnan | 1
Hi-pressure Micro pump | 555 | Lab 281 | NRTL-modified |  | TPR Group | BP 000220 | Approved | 2007-04-10 | J. Durnan | 1 | Approved receptacle & plug added to cord from controller to pump
Pulse Signal Box | 555 | Lab 261 | BNL built |  | TPR Group | BP 000701 | Approved | 2007-04-10 | J. Durnan | 1 | Used with COTS power supply
Flash Tube power supply | 555 | Lab 261 | BNL built |  | TPR Group | BP 000706 | Approved | 2007-04-10 | J. Durnan | 1
Laser Interlock | 555 | Lab 261 | BNL built |  | TPR Group | BP 000707 | Approved | 2007-04-10 | J. Durnan | 1
Chopper power supply | 555 | Lab 261 | BNL built |  | TPR Group | BP 000708 | Approved | 2007-04-10 | J. Durnan | 1
Shutter Controller on laser bench | 555 | Lab 261 | BNL built |  | TPR Group | BP 000709 | Approved | 2007-04-10 | J. Durnan | 1
Stepping motor power supply | 555 | Lab 261 | BNL built |  | TPR Group | BP 000710 | Approved | 2007-04-10 | J. Durnan | 1
Power Supply box | 555 | Lab 323 | BNL built |  | GPMD Group |  | Rejected | 2007-04-10 | J. Durnan | 1 | Deleting from system- excessing/discarding -cord cut
Microscope Illuminator power supply | 555 | Lab 323 | BNL built |  | GPMD Group |  | Rejected | 2007-04-10 | J. Durnan | 1 | Deleting from system- excessing/discarding -cord cut
Buffer Amp power supply | 555 | Lab 321 | BNL built |  | GPMD Group | BP000221 | Approved | 2007-04-10 | J. Durnan | 1
Laser Shutter door Interlock system | 555 | Lab 321 | BNL built |  | GPMD Group | BP000222 | Approved | 2007-04-10 | J. Durnan | 1 | Inspected similar unit in lab 261
Rotator | 555 | Lab 279 -W | BNL built |  | TPR Group | BP000223 | Approved | 2007-04-10 | J. Durnan | 1
Lamp Trigger | 555 | Lab 261 | BNL built |  | TPR Group | BP000224 | Approved | 2007-04-10 | J. Durnan | 1
Current Controller | 555 | Lab 279 -W | BNL built |  | TPR Group | BP000225 | Approved | 2007-04-10 | J. Durnan | 1
Water Soloniod Controller | 555 | Lab 323 | BNL built |  | GPMD Group | BP000226 | Approved | 2007-04-10 | J. Durnan | 1
Power Supply +5 V +_15 V | 555 | Lab 323 | BNL built |  | GPMD Group | BP000227 | Rejected | 2007-04-10 | J. Durnan | 1 | No fuse; enclosure needs bonding
Chopper Motor | 555 | Lab 323 | BNL built |  | GPMD Group | BP000228 | Approved | 2007-04-10 | J. Durnan | 1
Laser Shutter Door Interlock system | 555 | Lab 323 | BNL built | CH5-23 | GPMD Group | BP000229 | Approved | 2007-04-10 | J. Durnan | 1 | Inspected similar unit in lab 261
Vacuum Pumping Station | 555 | Lab 321 | BNL built |  | GPMD Group | BP000230 | Conditional Approval | 2007-04-10 | J. Durnan | 1 | 2-wire fan motor requires ground wire
Ultra-fast Chemical dynamics experiment | 555 | Lab 353 | BNL built |  | SCD Group | BP000234 | Conditional Approval | 2007-04-10 | J. Durnan | 1 | Ground Cable Tray; no extension cords with RPT's; GFCI's for bakeout heating wires; if use Isolation valve control circuit -use splice box for instrument connector
Capacitance Manometer Readout | 555 | Lab 325 | BNL built |  | GPMD Group | BP000235 | Approved | 2007-04-10 | J. Durnan | 1
Voltage Amplifier | 555 | Lab 221 | BNL built |  | Solar Neutrino Group |  | Rejected | 2007-04-17 | J. Durnan | 1 | Excessing/discarding
Power Supplies | 555 | Lab 233 | BNL built |  | Solar Neutrino Group |  | Rejected | 2007-04-17 | J. Durnan | 2 | Excessing/discarding
Auger Controller | 555 | Lab 127 | BNL built | B | White Group | BP 000231 | Approved | 2007-04-17 | J. Durnan | 1
Power Distribution box | 555 | Lab 127 | BNL built |  | White Group | BP 000232 | Approved | 2007-04-17 | J. Durnan | 1
Power Distribution box | 555 | Lab 125 | BNL built |  | White Group | BP 000233 | Approved | 2007-04-17 | J. Durnan | 1
Amplifier | 555 | Lab 245 | BNL built |  | Catalysis Group | BP 000236 | Approved | 2007-04-17 | J. Durnan | 1
Amplifier | 555 | Lab 245 | BNL built |  | Catalysis Group | BP 000237 | Approved | 2007-04-17 | J. Durnan | 1
Power Supply | 555 | Lab 245 | BNL built |  | Catalysis Group | BP 000238 | Approved | 2007-04-17 | J. Durnan | 1
Heater Distribution box | 555 | Lab 127 | BNL built |  | White Group | BP 000239 | Conditional Approval | 2007-04-17 | J. Durnan | 1 | No bonding jumper to chassis
Heater Distribution box | 555 | Lab 127 | BNL built |  | White Group | BP 000240 | Conditional Approval | 2007-04-17 | J. Durnan | 1
Amplifier | 555 | Lab 245 | BNL built |  | Catalysis Group | BP 000241 | Approved | 2007-04-17 | J. Durnan | 1
Gallium Leakage Alarm | 555 | Lab 225 | BNL built |  | Solar Neutrino Group | BP 000242 | Conditional Approval | 2007-04-17 | J. Durnan | 1 | In service, could not inspect
TTL Trigger | 555 | 179 Chase | BNL built | CH6-5 | TPR Group | BP 000243 | Approved | 2007-04-17 | J. Durnan | 1
Signal Reducer | 555 | 179 Chase | BNL built |  | TPR Group | BP 000244 | Rejected | 2007-04-17 | J. Durnan | 1 | Require hard ground to chassis

## ATF
Accelerator Test Facility EEI
Total
Manufacturer/Model#/Equipment Name | Equipment Location | Equipment Origin | Equipment ID Number | Cognizant Group/Person | Equipment Reference Number | Inspection Information | Date | Inspector | 3 | Comments
ATF Modulator #1 | 820 ATF Control Rm | BNL Built | 129245(BNL Inventory) | A T F/ K. Kushe | BP 000513 | A | 2006-08-24 | J. Durnan | 1 | Installed in ATF Control Rm
X-Band Modulator | 820 -High Bay | BNL Built | 129246(BNL Barcode) | ATF/ K. Kushe | BP 000515 | A | 2007-01-12 | J. Durnan | 1 | Klystron modulator- on floor near chillers
Neslab Instruments Chiller | 820 -High Bay | COT | 82120(BNL Barcode) | ATF/ K. Kushe | BP 000516 | A | 2007-01-12 | J. Durnan | 1 | Bottom chiller in rack
Note on A4: This is the equipment manufacturer, model # and name. If built "in-house", enter the name you assign to the equipment. Example1: Hewlett Packard 8752a Network Analyzer Example2: Experimental Sample Temperature Controller This is entered before inspection by the Equipment Owner
Note on C4: This is either COTS (commercial off the shelf) CABS (constructed according to BNL specification) BIHB (built in house by BNL) BBEC (brought by external collaborator) O (other) This is entered before inspection by the Equipment Owner
Note on D4: This is either the commercial serial number BNL inventory number, or group assigned number. If none exists, enter NA This is entered before inspection by the Equipment Owner
Note on E4: Person or group responsible for equipment This is entered before inspection by the Equipment Owner
Note on F4: This is the number on the scan tag. This is entered at time of inspection by the EEI
Note on G4: This is either A (approved Lt Grn) CA (Conditionally Approved Lt Yel) NRTL (NRTL Approved Lt Blu) DNA/NIS (Does not Apply/Not in System Lt Vio) R (Rejected Red) This is entered at time of inspection by the EEI
Note on H4: Date inspected This is entered at time of inspection by the EEI
Note on I4: This is entered at time of inspection by the EEI
Note on B4: This is the equipment location. This is entered before inspection by the Equipment Owner

## Magnet
Magnet Division EEI
Total
Manufacturer/Model#/Equipment Name | Equipment Location | Equipment Origin | Equipment ID Number | Cognizant Group/Person | Equipment Reference Number | Inspection Information | Date | Inspector | 18 | Comments
Herco/Zero Flux Current X-former #1 | 902 Rm 22A | COTS | 81374(BNL Inventory) | Magnet Division/P. Ribaudo | BP 000533 | A | 2006-05-10 | J. Durnan | 1 | Installed in Vertical PS #1-refurbished Power Supply
Herco/Zero Flux Current X-former #2 | 902 Rm 22A | COTS | D20 BP 060 | Magnet Division/P. Ribaudo | BP 000532 | A | 2006-05-10 | J. Durnan | 1 | Installed in Vertical PS #2-refurbished Power Supply
IGBT -Thermocouple Chasis PS #2 | 902 Rm 22A | BIHB | 1010 | Magnet Division/P. Ribaudo | BP 000526 | A | 2006-05-10 | J. Durnan | 1 | Installed in Vertical PS #2-refurbished Power Supply
IGBT -Thermocouple Chasis PS #1 | 902 Rm 22A | BIHB | 1009 | Magnet Division/P. Ribaudo | BP 000528 | A | 2006-05-10 | J. Durnan | 1 | Installed in Vertical PS #1-refurbished Power Supply
IGBT -Control Bucket PS #2 | 902 Rm 22A | BIHB | 1007 | Magnet Division/P. Ribaudo | BP 000527 | A | 2006-05-10 | J. Durnan | 1 | Installed in Vertical PS #2-refurbished Power Supply
IGBT -Control Bucket PS #1 | 902 Rm 22A | BIHB | 1008 | Magnet Division/P. Ribaudo | BP 000535 | A | 2006-05-10 | J. Durnan | 1 | Installed in Vertical PS #1-refurbished Power Supply
Power Supply Thermocouple Chasis 1&2 | 902 Rm 22A | BIHB | 250 | Magnet Division/P. Ribaudo | BP 000534 | A | 2006-05-10 | J. Durnan | 1 | Installed in Vertical PS #1-refurbished Power Supply
Siemens F.D.S. J13 | 902 Mag. Test Area | COTS | Cat. # BT34160 | Magnet Test Group/J. Cintirino | BP 000531 | A | 2006-05-30 | J. Durnan | 1 | modified door handle for Kirk Key
HP 6177C DC Current sourrce | 902 High bay | COTS | N/A | Magnet Division/P. Ribaudo | BP000506 | A | 2006-02-22 | J. Durnan | 1 | N/A
HP 3478A Multimeter | 902 High bay | COTS | 1008 | Magnet Division/P. Ribaudo | BP000501 | A | 2006-02-22 | J. Durnan | 1 | N/A
HP 3478A Multimeter | 902 High bay | COTS | 1007 | Magnet Division/P. Ribaudo | BP000502 | A | 2006-02-22 | J. Durnan | 1 | N/A
HP 3456A Digital Voltmeter | 902 High bay | COTS | 1029 | Magnet Division/P. Ribaudo | BP000503 | A | 2006-02-22 | J. Durnan | 1 | N/A
Short sample quencher detector | 902 high bay | BNL Built | N/A | Magnet Division/P. Ribaudo | BP000504 | A | 2006-02-22 | J. Durnan | 1 | Used for detection of quench of superconducting magnets
Agilant 6680A system DC power supply | 902 High bay | COTS | BNL 123513 | Magnet Division/P. Ribaudo | BP000505 | NRTL | 2006-02-22 | J. Durnan | 1 | N/A
DC Voltage reference chassis | 902 High bay | BNL Built | N/A | Magnet Division/P. Ribaudo | BP000507 | Conditional Approval | 2006-02-22 | J. Durnan | 1 | Soft ground not checked yet
Isolation buffer chassis | 902 high bay | BNL built | chasis # 202 | Magnet Division/P. Ribaudo | BP000530 | Approved | 2006-04-27 | J. Durnan | 1
Trim quench detector | 902 High bay | BNL built | chassis #235 | Magnet Division/P. Ribaudo | BP000529 | Approved | 2006-04-27 | J. Durnan | 1
BEHRINGER -Power Hacksaw | 902 machine shop | COTS | A050031 | AM T. Dilgen | BP000519 | Approved | 2007-02-02 | J. Durnan | 1
Note on A4: This is the equipment manufacturer, model # and name. If built "in-house", enter the name you assign to the equipment. Example1: Hewlett Packard 8752a Network Analyzer Example2: Experimental Sample Temperature Controller This is entered before inspection by the Equipment Owner
Note on C4: This is either COTS (commercial off the shelf) CABS (constructed according to BNL specification) BIHB (built in house by BNL) BBEC (brought by external collaborator) O (other) This is entered before inspection by the Equipment Owner
Note on D4: This is either the commercial serial number BNL inventory number, or group assigned number. If none exists, enter NA This is entered before inspection by the Equipment Owner
Note on E4: Person or group responsible for equipment This is entered before inspection by the Equipment Owner
Note on F4: This is the number on the scan tag. This is entered at time of inspection by the EEI
Note on G4: This is either A (approved Lt Grn) CA (Conditionally Approved Lt Yel) NRTL (NRTL Approved Lt Blu) DNA/NIS (Does not Apply/Not in System Lt Vio) R (Rejected Red) This is entered at time of inspection by the EEI
Note on H4: Date inspected This is entered at time of inspection by the EEI
Note on I4: This is entered at time of inspection by the EEI
Note on B4: This is the equipment location. This is entered before inspection by the Equipment Owner

## Material-480
Material Science EEI
Total
Manufacturer/Equipment Name | Model no. | Building | Equipment Location | Equipment Origin | Equipment ID Number | Cognizant Group/Person | EEI BarCode Prefix | EEI BarCode Number | Inspection Information | Date | Inspector | 187 | Comments
|  | 480 |  | COTS |  | Material Science | BP | 0 | Approved |  | A. Piper | 1 | N/A
Model 650 low speed diamond wheel saw |  | 480 |  | COTS | 650-1546 | Material Science | BP | 1 | Approved | 2005-10-11 | A. Piper | 1 | N/A
Transformer for microscope light assembly |  | 480 |  | COTS | cat. # 969-001 model # 282 | NC | BP | 2 | NRTL | 2005-10-11 | A. Piper | 1 | N/A
Thermolyne nuova 2 hot plate |  | 480 |  | COTS | S/N 30511122 | NC | BP | 3 | NRTL | 2005-10-11 | A. Piper | 1 | N/A
VCR group dimpler | D500i | 480 |  | COTS | S/N 112699 | NC | BP | 4 | Approved | 2005-10-11 | A. Piper | 1 | N/A
VCR group ion mill | XLA 2000 | 480 |  | COTS | S/N 112599 | NC | BP | 5 | Approved | 2005-10-11 | A. Piper | 1 | N/A
Bueler variable speed grinder-polisher |  | 480 |  | COTS | S/N 606EPN00185 | NC | BP | 6 | Approved | 2005-10-11 | A. Piper | 1 | N/A
Bamstead/thermolyne hot plate | type 1900 | 480 |  | COTS | S/N 64030186325 | NC | BP | 7 | NRTL | 2005-10-11 | A. Piper | 1 | N/A
Nova 2000 high intensity illuminator |  | 480 |  | COTS | S/N 00013625 | NC | BP | 8 | NRTL | 2005-10-11 | A. Piper | 1 | N/A
South bay technology polishing machine | model 910 | 480 |  | COTS | S/N 093099 | NC | BP | 9 | Approved | 2005-10-11 | A. Piper | 1 | N/A
low angle ion mill & polishing machine | Model 1010 | 480 |  | COTS | S/N 105298 | NC | BP | 10 | Approved | 2005-10-11 | A. Piper | 1 | N/A
Coolflow recirculator |  | 480 |  | COTS | S/N 87 NML63120-39 | NC | BP | 11 | Approved | 2005-10-11 | A. Piper | 1 | N/A
Baltec res101 rapid etching system |  | 480 |  | COTS | # 200 | NC | BP | 12 | Approved | 2005-10-11 | A. Piper | 1 | N/A
Ultra sonik 57X parts washer |  | 480 |  | COTS | 9493435 | NC | BP | 13 | NRTL | 2005-10-11 | A. Piper | 1 | N/A
Cole-parmer ultra sonic cleaneer |  | 480 |  | COTS | QGA100389803 | NC | BP | 14 | Approved | 2005-10-11 | A. Piper | 1 | N/A
South bay technology wire saw |  | 480 |  | COTS | S/N 121391 | NC | BP | 15 | Approved | 2005-10-11 | A. Piper | 1 | N/A
Microscope lamp - bausch & laumb |  | 480 |  | COTS | cat. # 31-35-47 | NC | BP | 16 | NRTL | 2005-10-11 | A. Piper | 1 | N/A
|  | 480 | Lab 1-3 | COTS | 2833A-01487 | CMPMSD | BP | 17 | Approved | 2005-10-18 | A. Piper | 1 | N/A
Sunbonder ultrasonic soldering unit |  | 480 | Lab 1-3 | COTS |  | Material Science | BP | 18 | Approved | 2005-11-01 | A. Piper | 1 | Initially failed. Accepted after required repairs were completed.
|  | 480 | Lab 1-3 | COTS | 120134 | CMPMSD | BP | 19 | NRTL | 2005-10-18 | A. Piper | 1 | N/A
Hakkosoldering iron | 926 | 480 | Lab 1-3 | COTS | N/A | CMPMSD | BP | 20 | Approved | 2005-10-18 | A. Piper | 1 | N/A
Keithley multimeter | 2001 | 480 | Lab 1-3 | COTS | 629876 | CMPMSD | BP | 21 | Approved | 2005-10-18 | A. Piper | 1 | N/A
Keithley digital voltmeter | 182 | 480 | Lab 1-3 | COTS | 505007 | CMPMSD | BP | 22 | Approved | 2005-10-18 | A. Piper | 1 | N/A
Keithley high current unit | 238 | 480 | Lab 1-3 | COTS | 658125 | CMPMSD | BP | 23 | Approved | 2005-10-18 | A. Piper | 1 | N/A
Quantum design rack system |  | 480 | Lab 1-3 | COTS | E025 for model 600 contoller (main component in rack) | CMPMSD | BP | 24 | Approved | 2005-10-18 | A. Piper | 1 | N/A
|  | 480 | Lab 1-3 | COTS | BNL property # 55420 | CMPMSD | BP | 25 | Approved | 2005-10-18 | A. Piper | 1 | N/A
|  | 480 |  | COTS |  | Material Science | BP | 26 | Approved |  | A. Piper | 1 | N/A
Branson 200 ultrasonic cleaner |  | 480 |  | COTS | S/N 00014528 | Material Science | BP | 27 | NRTL | 2005-11-01 | A. Piper | 1 | N/A
Thermolyne type 10600 hot plate oven |  | 480 | Lab 1-3 | COTS | S/N 64040702894 | Material Science | BP | 27 | NRTL | 2005-11-01 | A. Piper | 1 | N/A
Lake shore cryotronics constant current power supply |  | 480 |  | COTS | S/N 1189 | Material Science | BP | 28 | Approved | 2005-11-01 | A. Piper | 1 | N/A
Polaroid micro cam camera |  | 480 |  | COTS | BNL barcode 108516 | Material Science | BP | 29 | NRTL | 2005-11-01 | A. Piper | 1 | N/A
Barnant company vacuum pressure pump |  | 480 |  | COTS | S/N M0003615 | Material Science | BP | 30 | NRTL | 2005-11-01 | A. Piper | 1 | N/A
Edwards active gauge controller |  | 480 |  | COTS | S/N 017021708 | Material Science | BP | 31 | Approved | 2005-11-01 | A. Piper | 1 | N/A
Edwards 1.5 two stage vacuum pump |  | 480 |  | COTS | S/N 017212341 | Material Science | BP | 32 | NRTL | 2005-11-01 | A. Piper | 1 | N/A
Veeco TG-70 gauge control |  | 480 |  | COTS | S/N 00-5882-4317 | Material Science | BP | 33 | Approved | 2005-11-01 | A. Piper | 1 | N/A
Veeco TG-7 thermocouple gauge control |  | 480 |  | COTS | S/N 00-5882-7840 | Material Science | BP | 34 | Approved | 2005-11-01 | A. Piper | 1 | N/A
|  | 480 |  | COTS |  | Material Science | BP | 35 | Approved |  | A. Piper | 1 | N/A
Hewlett-packard |  | 480 |  | COTS | S/N 2731A-11386 | Material Science | BP | 36 | Approved | 2005-11-01 | A. Piper | 1 | N/A
Lindberg blue oven unit |  | 480 |  | COTS | N/A | Material Science | BP | 36 | Approved | 2005-11-01 | A. Piper | 1 | N/A
BOC edwards vacuum pump unit |  | 480 |  | COTS | S/N 006832224 | Material Science | BP | 37 | Approved | 2005-11-01 | A. Piper | 1 | N/A
Keithly 182 sensitive digital voltmeter |  | 480 |  | COTS | S/N 0570777 | Material Science | BP | 38 | Approved | 2005-11-01 | A. Piper | 1 | N/A
New wave research microscope system |  | 480 |  | COTS | S/N 4209 (base unit) | Material Science | BP | 39 | Approved | 2005-11-01 | A. Piper | 1 | N/A
Stocker & yale image light |  | 480 |  | COTS | S/N 100443985 | Material Science | BP | 40 | NRTL | 2005-11-01 | A. Piper | 1 | N/A
Stocker & yale image light |  | 480 |  | COTS | S/N 10043988 | Material Science | BP | 41 | NRTL | 2005-11-01 | A. Piper | 1 | N/A
Hakko 936 soldering station |  |  |  | COTS | S/N 9361605609 | Material Science | BP | 42 | Approved | 2005-11-01 | A. Piper | 1 | N/A
Whisper 400 air pump |  | 480 |  | COTS | S/N 091197 | Material Science | BP | 43 | NRTL | 2005-11-01 | A. Piper | 1 | N/A
Blak-Ray longwave uultraviolet lamp |  | 480 |  | COTS | Model B-100A | Material Science | BP | 44 | NRTL | 2000-11-01 | A. Piper | 1 | N/A
Mag-mix hot plate |  | 480 | Lab 1-4 | COTS | S/N 13AC-6 | Material Science | BP | 45 | Approved | 111/22/05 | A. Piper | 1 | N/A
Olympus tokyo microscope |  | 480 | Lab 1-4 | COTS | S/N 532628 (microscope base) | Material Science | BP | 46 | Approved | 2005-11-22 | A. Piper | 1 | N/A
Mettler PC440 scale |  | 480 | Lab 1-4 | COTS | S/N A69938 | Material Science | BP | 47 | Approved | 2005-11-22 | A. Piper | 1 | N/A
Voltalab PGZ-402 |  | 480 | Lab 1-4 | COTS | S/N 674R020N009 | Material Science | BP | 48 | Approved | 2005-11-22 | A. Piper | 1 | N/A
Keithly digital voltmeter |  | 480 | Lab 1-4 | COTS | S/N 60887A | Material Science | BP | 49 | Approved | 2005-11-22 | A. Piper | 1 | N/A
Ministat potentiostat |  | 480 | Lab 1-4 | COTS | S/N 0163-2 | Material Science | BP | 50 | Approved | 2005-11-22 | A. Piper | 1 | N/A
Voltammograph CV-27 |  | 480 | Lab 1-4 | COTS | S/ N 647 | Material Science | BP | 51 | Approved | 2005-11-22 | A. Piper | 1 | N/A
Scribner MMA |  | 480 | Lab 1-4 | COTS | S/N 0415229 | Material Science | BP | 52 | Approved | 2005-12-13 | A. Piper | 1 | N/A
HP 3478A multimeter |  | 480 | Lab 1-4 | COTS | S/N 2301A17510 | Material Science | BP | 53 | Approved | 2005-12-13 | A. Piper | 1 | N/A
Ministat low noise potentiostat |  | 480 | Lab 1-4 | COTS | S/N 0823 | Material Science | BP | 54 | Approved | 2005-12-13 | A. Piper | 1 | N/A
Foster EKE-R |  | 480 | Lab 1-4 | COTS | S/N B1189 | Material Science | BP | 55 | NRTL | 2005-12-13 | A. Piper | 1 | N/A
SRS SR830 amplifier |  | 480 | Lab 1-4 | COTS | S/N 21551 | Material Science | BP | 56 | Approved | 2005-12-13 | A. Piper | 1 | N/A
HP 1740A oscilloscope |  | 480 | Lab 1-4 | COTS | S/N 2306A20121 | Material Science | BP | 57 | Approved | 2005-12-13 | A. Piper | 1 | N/A
Keithly 614 electrometer |  | 480 | Lab 1-4 | COTS | S/N 123885 | Material Science | BP | 58 | Approved | 2005-12-13 | A. Piper | 1 | N/A
HP 1222A oscilloscope |  | 480 | Lab 1-4 | COTS | S/N 1845505777 | Material Science | BP | 59 | Approved | 2005-12-13 | A. Piper | 1 | N/A
Master heat gun |  | 480 | Lab 1-4 | COTS | M691 | Material Science | BP | 60 | NRTL | 2005-12-27 | A. Piper | 1 | N/A
HP 6264B DC power supply |  | 480 | Lab 1-4 | COTS | 2334A-05484 | Material Science | BP | 61 | Approved | 2005-12-27 | A. Piper | 1 | N/A
Polyscience heating/cooling unit |  | 480 | Lab 1-4 | COTS | BNL ID A044859 | Material Science | BP | 62 | Approved | 2005-12-27 | A. Piper | 1 | N/A
Fisher 200M hotplate |  | 480 | Lab 1-4 | COTS | S/N 401 | Material Science | BP | 63 | Approved | 2005-12-27 | A. Piper | 1 | N/A
Voltalab PGZ 100 |  | 480 | Lab 1-4 | COTS | BNL # 117659 | Material Science | BP | 64 | Approved | 2005-12-27 | A. Piper | 1 | N/A
Model 4202 (system) |  | 480 | Lab 1-5 | COTS | BNL barcode 001877 | Material Science | BP | 65 | Approved |  | A. Piper | 1 | N/A
Veeco TG70 controller |  | 480 | Lab 1-5 | COTS | S/N 00-882-3871 | Material Science | BP | 66 | Approved | 2006-01-03 | A. Piper | 1 | N/A
Bausch & lamb microscope |  | 480 | Lab 1-5 | COTS | BNL barcode 003104 | Material Science | BP | 67 | Approved |  | A. Piper | 1 | N/A
Thermolyne nuova 2 stir plate |  | 480 | Lab 1-5 | COTS | S/N 30637008 | Material Science | BP | 68 | NRTL | 2006-01-03 | A. Piper | 1 | N/A
Keithly 614 electrometer |  | 480 | Lab 1-5 | COTS | S/N 137424 | Material Science | BP | 69 | Approved | 2006-01-03 | A. Piper | 1 | N/A
Lambda LE 104-FM power supply |  | 480 | Lab 1-4 | COTS | BNL property # 72310 | Material Science | BP | 70 | Approved | 2006-01-03 | A. Piper | 1 | N/A
|  | 480 | Lab 1-5 | COTS | BNL barcode 003051 | Material Science | BP | 71 | Approved | 2006-01-03 | A. Piper | 1 | N/A
JSM-6400 scanning microscope (system) |  | 480 | Lab 1-5 | COTS | BNL barcode (main section) 87573 | Material Science | BP | 72 | Approved | 2006-02-10 | A. Piper | 1 | Initially received conditional approval. The following repairs were made for full approval: repair 208v line cord strain relief; connect unit to house ground, bond all cabinets to mother unit.
Perkin-elmer PHI600 scanning auger multiprobe (system) |  | 480 | Lab 1-5 | COTS | BNL barcode 100900 | Material Science | BP | 73 | Approved | 2006-02-10 | A. Piper | 1 | Initially received conditional approval. The following repair was made to receive full approval: install cover on the interface module.
Mettler deltarange scale | AT261 | 480 | Lab 1-5 | COTS | BNL barcode 81176 | Material Science | BP | 74 | Approved | 2006-01-20 | A. Piper | 1 | N/A
|  | 480 | Lab 1-5 | COTS | N/A | Material Science | BP | 75 | NRTL | 2006-01-20 | A. Piper | 1 | N/A
Keithly 155 null detector voltmeter |  | 480 | Lab 1-4 | COTS | N/A | Material Science | BP | 76 | Approved | 2006-01-20 | A. Piper | 1 | N/A
Love controls corporation heating/cooling unit |  | 480 | Lab 1-6 | COTS | # 290-281 | Material Science | BP | 77 | Conditional approval | 2006-01-24 | A. Piper | 1 | replace in-line fuse with correct size fuse; ground chassis; fabricate & install rear panel cover
Thermonics e-gun control |  | 480 | Lab 1-6 | COTS | S/N 27629 | Material Science | BP | 78 | Approved | 2006-01-24 | A. Piper | 1 | N/A
Powerstat variable autotransformer |  | 480 | Lab 1-6 | COTS | Type 3PN116C | Material Science | BP | 79 | Approved | 2006-01-24 | A. Piper | 1 | N/A
SPEX 8000 mixer/mill |  | 480 | Lab 1-6 | COTS | N/A | Material Science | BP | 80 | Approved | 2006-01-24 | A. Piper | 1 | N/A
Troemner submersible magnetic stirrer |  | 480 | Lab 1-6 | COTS | Model 1700 | Material Science | BP | 81 | Approved | 1/24//06 | A. Piper | 1 | N/A
TA instuments calorimeter |  | 480 | Lab 1-6 | COTS | BNL barcode # 113995 | Material Science | BP | 82 | Approved | 2006-01-24 | A. Piper | 1 | N/A
TA instruments SDT 2960 |  | 480 | Lab 1-6 | COTS | BNL barcode # 85386 | Material Science | BP | 83 | Approved |  | A. Piper | 1 | N/A
Mettler electronic scale |  | 480 | Lab 1-6 | COTS | BNL property tag # 50522 | Material Science | BP | 84 | Approved | 2006-01-24 | A. Piper | 1 | N/A
Fisher scientific analytical balance |  | 480 | Lab 1-6 | COTS | S/N 30239 | Material Science | BP | 85 | Approved | 2006-01-24 | A. Piper | 1 | N/A
Penguin computing altus 2100 |  | 480 | Lab 1-108 | COTS | S/N P1235120456 | Material Science | BP | 86 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Netgear 16 port gigabit unit |  | 480 | Lab 1-108 | COTS | S/N 1232565A001F7 | Material Science | BP | 87 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Penguin computing altus 1300 |  | 480 | Lab 1-108 | COTS | S/N P1235120458 | Material Science | BP | 88 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Penguin computing altus 1300 |  | 480 | Lab 1-108 | COTS | S/N P1235120459 | Material Science | BP | 89 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Penguin computing altus 1300 |  | 480 | Lab 1-108 | COTS | S/N P1235120460 | Material Science | BP | 90 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Penguin computing altus 1300 |  | 480 | Lab 1-108 | COTS | S/N P1235120457 | Material Science | BP | 91 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
SVA2200RM2U UPS |  | 480 | Lab 1-108 | COTS | S/N JS0532027480 | Material Science | BP | 92 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Nexus 470 system |  | 480 | Lab 1-6 | COTS | BNL barcode # 116849 | Material Science | BP | 93 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Newport power meter |  | 480 | Lab 1-6 | COTS | S/N 1189 | Material Science | BP | 94 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Keithly 2000 multimeter |  | 480 | Lab 1-6 | COTS | S/N 105298 | Material Science | BP | 95 | Approved | 2006-02-10 | A. Piper | 1 | N/A
Keithly 2400-C source meter |  | 480 | Lab 1-6 | COTS | S/N 1007558 | Material Science | BP | 96 | NRTL | 2006-02-10 | A. Piper | 1 | N/A
Lake shore 340 temperature controller |  | 480 | Lab 1-6 | COTS | S/N 341553 | Material Science | BP | 97 | Approved | 2006-02-10 | A. Piper | 1 | N/A
Lindberg/blue box furnace |  | 480 | Lab 1-9 | COTS | Model # BF51732C-1 | Material Science | BP | 98 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Fiber-Lite model 190 |  | 480 | Lab 1-9 | COTS | N/A | Material Science | BP | 99 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Multimode atomic force microscope (system) |  | 480 | Lab 1-9 | COTS | S/N 1371EX | Material Science | BP | 100 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Nikon epihot 200 microscope |  | 480 | Lab 1-9 | COTS | S/N 1102897 | Material Science | BP | 101 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Beckman coulter DM640 spectrophotometer |  | 480 | Lab 1-9 | COTS | BNL barcode # 117266 | Material Science | BP | 102 | NRTL | 2006-03-08 | A. Piper | 1 | N/A
Nexus 670 FT IR |  | 480 | Lab 1-9 | COTS | BNL barcode # 122009 | Material Science | BP | 103 | NRTL | 2006-03-08 | A. Piper | 1 | N/A
Lindberg/blue control console |  | 480 | Lab 1-9 | COTS | Model # CC58114PC-1 | Material Science | BP | 104 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Betham 218F VF chopper |  | 480 | Lab 1-9 | COTS | BNL property # 002787 | Material Science | BP | 105 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Speedline technologies model P6204 coater |  | 480 | Lab 1-9 | COTS | S/N 136204-4 | Material Science | BP | 106 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Duoseal vacuum pump model # 1500 |  | 480 | Lab 1-9 | COTS | S/N 3762 | Material Science | BP | 107 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Kepco bipolar power supply |  | 480 | Lab 1-9 | COTS | Model BOP 20-10M | Material Science | BP | 108 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Eldex VS metering pump |  | 480 | Lab 1-9 | COTS | BNL barcode # 092545 | Material Science | BP | 109 | Approved | 2006-03-08 | A. Piper | 1 | manual was available
Tektroniics 2235 oscillocope |  | 480 | Lab 1-9 | COTS | BNL barcode 49948 | Material Science | BP | 110 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Cimarec 2 hot plate |  | 480 | Lab 1-9 | COTS | Model # HP46825 | Material Science | BP | 111 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Ohaus analytical scale |  | 480 | Lab 1-9 | COTS | BNL barcode # 89277 | Material Science | BP | 112 | Approved | 2006-03-08 | A. Piper | 1 | N/A
Curtis matheson mixer plate |  | 480 | Lab 1-9 | COTS | S/N 280-1372 | Material Science | BP | 113 | Approved | 2006-03-08 | A. Piper | 1 | N/A
mag-mix hot plate |  | 480 | Lab 1-9 | COTS | S/N 19AM-8 | Material Science | BP | 114 | Approved | 3//8/06 | A. Piper | 1 | N/A
Curtis matheson mixer |  | 480 | Lab 1-9 | COTS | S/N 280-1369 | Material Science | BP | 115 | Approved | 2006-03-08 | A. Piper | 1 | N/A
BLH transducer-conditioner, model 5200 |  | 480 | Lab 1-9 | COTS | S/N 3101 | Material Science | BP | 116 | Rejected | 2006-03-24 | A. Piper | 1 | Ground plug removed.
Corman industries pump |  | 480 | Lab 1-9 | COTS | S/N Z11968 | Material Science | BP | 117 | Approved | 2006-03-24 | A. Piper | 1 | N/A
Gra Lab universal timer |  | 480 | Lab 1-10 | COTS | Type 171 | Material Science | BP | 118 | Approved | 2006-03-24 | A. Piper | 1 | N/A
Fisher scientific FS60H |  | 480 | Lab 1-10 | COTS | BNL property # A017505 | Material Science | BP | 119 | Approved | 2006-03-24 | A. Piper | 1 | N/A
Scientific products MV1 |  | 480 | Lab 1-10 | COTS | S/N 79550 | Material Science | BP | 120 | NRTL | 2006-03-24 | A. Piper | 1 | N/A
Buschi 461 water bath |  | 480 | Lab 1-10 | COTS | N/A | Material Science | BP | 121 | Conditional approval | 2006-03-24 | A. Piper | 1 | Unit requires repairs - need further evaluation, case is adequately grounded
VWR scientific products |  | 480 | Lab 1-10 | COTS | S/N 0201301 | Material Science | BP | 122 | NRTL | 2006-03-31 | A. Piper | 1 | N/A
Franklin electric vacuum pump/Welch 1400 |  | 480 | Lab 1-10 | COTS | BNL barcode # 84590 | Material Science | BP | 123 | NRTL | 2006-03-31 | A. Piper | 1 | N/A
Beckman coulter allegra 21 centrifuge |  | 480 | Lab1-10 | COTS | BNL barcode # 117267 | Material Science | BP | 124 | NRTL | 2006-03-31 | A. Piper | 1 | N/A
Corning pinnacle 530 meter |  | 480 | Lab 1-10 | COTS | S/N 301402 | Material Science | BP | 125 | NRTL | 2006-03-31 | A. Piper | 1 | N/A
Corning 240 PH meter |  | 480 | Lab 1-10 | COTS | N/A | Material Science | BP | 126 | Approved | 2006-03-31 | A. Piper | 1 | N/A
Star energy variac model # 3PN1010B |  | 480 | Lab 1-10 | COTS | S/N 0131 (stamped on bottom) | Material Science | BP | 127 | NRTL | 2006-04-10 | A. Piper | 1 | N/A
Sartorius BP61S scale |  | 480 | Lab 1-10 | COTS | S/N 12003097 | Material Science | BP | 128 | Approved | 2006-04-10 | A. Piper | 1 | N/A
Precision 25EG oven |  | 480 | Lab 1-10 | COTS | S/N 600101770 | Material Science | BP | 129 | Approved | 2006-04-10 | A. Piper | 1 | N/A
CLT measurements electrochemical workstation |  | 480 | Lab 1-10 | COTS | BNL barcode 128326 | Material Science | BP | 130 | Approved | 2006-04-10 | A. Piper | 1 | N/A
Depths of the earth material press |  | 480 | Lab 1-128 | COTS | BNL barcode # 128628 | Material Science | BP | 131 | Approved | 2006-03-23 | A. Piper | 1 | additional house ground and "danger HV" label was added to be accepted.
PMC 5 position stirrer |  | 480 | Lab 1-10 | COTS | BNL property tag A017230 | Material Science | BP | 132 | NRTL | 2006-04-10 | A. Piper | 1 | N/A
Fisher ther-mix |  | 480 | Lab 1-10 | COTS | N/A | Material Science | BP | 133 | NRTL | 2006-04-10 | A. Piper | 1 | N/A
Therolyne cimarec 2 hot plate model # HP46825 |  | 480 | Lab 1-10 | COTS | S/N 6399604717787 | Material Science | BP | 134 | NRTL | 2006-04-10 | A. Piper | 1 | N/A
Thermolyne cimarec 2 hot plate |  | 480 | Lab 1-10 | COTS | S/N 639950249070 | Material Science | BP | 135 | NRTL | 2006-04-10 | A. Piper | 1 | N/A
Thermolyne nuova 2 hot plate |  | 480 | Lab 1-10 | COTS | S/N 30639537 | Material Science | BP | 136 | Conditional approval | 2006-04-10 | A. Piper | 1 | line cord requires replacement. Insulation jacket at the unit body is deteriorated.
Panalytical xpert pro x-ray diffraction sytem |  | 480 | Lab 1-7 | COTS | ID 12NC943003040601 | Material Science | BP | 137 | NRTL | 2006-04-10 | A. Piper | 1 | N/A
Haskro heat exchanger model # R100 |  | 480 | Lab 1-7 | COTS | S/N HB16692 | Material Science | BP | 138 | Approved | 2006-04-10 | A. Piper | 1 | N/A
Fisher scientific cantrific model 228 centrifuge |  | 480 | Lab 1-10 | COTS | S/N 312N0988 | Material Science | BP | 139 | NRTL | 2006-04-29 | A. Piper | 1 | N/A
Buchi rotovapor R-300 |  | 480 | Lab 1-10 | COTS | N/A | Material Science | BP | 140 | NRTL | 2006-04-29 | A. Piper | 1 | N/A
Electrothermal MC242X1 heater |  | 480 | Lab 1-10 | COTS | S/N 10598605 | Material Science | BP | 141 | Approved | 2006-04-29 | A. Piper | 1 | N/A
Fisher scientific 550 sonic dismembrator |  | 480 | Lab 1-10 | COTS | S/N F2996 | Material Science | BP | 142 | NRTL | 2006-04-29 | A. Piper | 1 | N/A
Lindberg /blue HP53016A hot plate |  | 480 | Lab 1-10 | COTS | S/N W06F-311618-WF | Material Science | BP | 143 | Approved | 2006-04-29 | A. Piper | 1 | N/A
Thermolyne cimarec 2 hot plate |  | 480 | Lab 1-10 | COTS | S/N 639950249088 | Material Science | BP | 144 | NRTL | 2006-04-29 | A. Piper | 1 | N/A
Corning stirrir/hot plate model # PC-620 |  | 480 | Lab 1-10 | COTS | S/N 370802351969 | Material Science | BP | 145 | NRTL | 2006-04-29 | A. Piper | 1 | N/A
Varian 9210066 pump control unnit |  | 480 | Lab 1-123 | COTS | S/N H1G3894 | Material Science | BP | 146 | Approved | 2006-04-29 | A. Piper | 1 | N/A
Granville-phillips series 271 gauge controller |  | 480 | Lab 1-123 | COTS | S/N 123 | Material Science | BP | 147 | Approved | 4//29/06 | A. Piper | 1 | N/A
Powerstat variable autotransformer |  | 480 | Lab 1-123 | COTS | type 3PN116C | Material Science | BP | 148 | Approved | 2006-04-29 | A. Piper | 1 | N/A
Variian model 922-0043 control unit |  | 480 | Lab 1-123 | COTS | S/N H1K31005 | Material Science | BP | 149 | Approved | 2006-05-12 | A. Piper | 1 | N/A
Sorenson model QRD20-4 power supply |  | 480 | Lab 1-123 | COTS | S/N 2948 | Material Science | BP | 150 | Rejected | 2006-05-12 | A. Piper | 1 | Line cord in poor condition, requires replacement
Bell 610 gaussmeter |  | 480 | Lab 1-123 | COTS | S/N 213279 | Material Science | BP | 151 | Approved | 2006-05-12 | A. Piper | 1 | N/A
Bransonic 12 ultrasonic cleaner |  | 480 | Lab 1-123 | COTS | N/A | Material Science | BP | 152 | Approved | 2006-05-12 | A. Piper | 1 | N/A
Corning PC-351 hot plate stirrer |  | 480 | Lab 1-123 | COTS | N/A | Material Science | BP | 153 | NRTL | 2006-05-12 | A. Piper | 1 | N/A
Power distribution rack (located in N/W corner) |  | 480 | Lab 1-123 | BNL built | N/A | Material Science | BP | 154 | Approved | 2006-05-12 | A. Piper | 1 | initially received a conditional approval. Repairs were performed and system was subsequently approved.
Experimental rack system (see notes for individual components) |  | 480 | Lab 1-123 | COTS | N/A | Material Science | BP |  | Approved | 2006-05-12 | A. Piper | 1 | 1)telemark model 850 thickness monitor 2)varian dual ion pump controller 3) Varian sublimation controller 4) thermonics PS-10 ion pump power supply 5) thermonics xysc controller 6) thermonics xys controller 7) thermonics seb3 emission gun supply
Electronic measurements TCR power supply |  | 480 | Lab 1-123 | COTS | S/N 85E3025 | Material Science | BP | 156 | Conditional approval | 2006-05-12 | A. Piper | 1 | Output terminals on rear of chassis need a fixed guard installed.
Bench-top sputtering system (see notes for individual components) |  | 480 | Lab 1-7 | COTS | 107321 - (BNL property tag # for power supply component) | Material Science | BP | 157 | Approved | 2006-05-12 | A. Piper | 1 | 1) US inc DC100 power supply 2) BNL fabricated sputtering system 3) generic (NRTL) vacuum pump
Thermolyne HP2305B hot plate |  | 480 | Lab 1-7 | COTS | S/N 242 | Material Science | BP | 158 | Approved | 2006-05-12 | A. Piper | 1 | N/A
MPMS-5 (system) |  | 480 | Lab 1-122 | COTS | 58073/74/76 | Material Science | BP | 159 | Approved | 2006-06-01 | A. Piper | 1 | N/A
Blue M. electric oven model 0V12A |  | 480 | Room 1-128 | COTS | S/N 2A3046 | Material Science | BP | 160 | NRTL | 2006-06-16 | A. Piper | 1 | N/A
Fiber-lite high intensity illuminator |  | 480 | Room 1-128 | COTS | N/A | Material Science | BP | 161 | NRTL | 2006-06-16 | A. Piper | 1 | N/A
SPEX 8000D mixer/mill |  | 480 | Room 1-128 | COTS | S/N 04000 | Material Science | BP | 162 | Approved | 2006-06-16 | A. Piper | 1 | N/A
Sputtering system (multiple components) |  | 480 | Room 1-128 | COTS | BNL barcodes 092556/093024 | Material Science | BP | 163 | Approved | 2006-06-16 | A. Piper | 1 | N/A
Labonco protector glove box |  | 480 | Room 1-128 | COTS | BNL property barcode # 123103 | Material Science | BP | 164 | Approved | 2006-06-16 | A. Piper | 1 | N/A
Vacuum/purge station |  | 480 | Room 1-128 | BNL built | BNL barcode # 002968 | Material Science | BP | 165 | Approved | 2006-06-16 | A. Piper | 1 | N/A
Lindberg type 59544-E5 tube furnace |  | 480 | Room 1-128 | COTS | S/N 878078-L | Material Science | BP | 166 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Spin melter system |  | 480 | Room 129 | COTS | BNL property barcode 123694 | Material Science | BP | 167 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Gero NR683/689 tube furnace "system" |  | 480 | Room 129 | COTS | 81564/91067 | Material Science | BP | 168 | Approved | 2006-06-26 | A. Piper | 1 | N/A
whirlpool ice system |  | 480 | Room 129 | COTS | BNL Property Tag # 104852 | Material Science | BP | 169 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Lindberg Tube Furnace |  | 480 | Room 129 | COTS | S/N 919098 | Material Science | BP | 170 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Blue tube furnace model HTF55342C |  | 480 | Room 129 | COTS | S/N 010E-198861-OE | Material Science | BP | 171 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Veeco TG-70 TC Gauge Controller |  | 480 | Room 129 | COTS | S/N 00-5882-01999 | Material Science | BP | 172 | Approved | 2006-06-26 | A. Piper | 1 | N/A
MKS Type 254A flow rate controller |  | 480 | Room 129 | COTS | BNL Property Barcode | Material Science | BP | 173 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Powerstat variable autotransformer type 3PN116B |  | 480 | Room 129 | COTS | N/A | Material Science | BP | 174 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Powerstat variable autotransformer type 3PN116B |  | 480 | Room 129 | COTS | N/A | Material Science | BP | 175 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Powerstat variable autotransformer type 116 |  | 480 | Room 129 | COTS | N/A | Material Science | BP | 176 | Approved | 2006-06-26 | A. Piper | 1 | N/A
Standford Research Systems | Model SR 560 | 480 | Lab 113 | COTS | serial 81495 | Laser Group | BP | 677 | Approved | 2007-03-01 | J. Durnan | 1 | N/A
Standford Research Systems | Model SR 830 | 480 | Lab 113 | COTS | serial 76872 | Laser Group | BP | 678 | Approved | 2007-03-01 | J. Durnan | 1 | N/A
DL Instruments -Current Amp | Model 1211 | 480 | Lab 113 | COTS | Serial 87639 | Laser Group | BP | 679 | Approved | 2007-03-01 | J. Durnan | 1 | N/A
Coherent Synchro-Lock AP | V 18 | 480 | Lab 113 | COTS | serial EM 064900002 | Laser Group | BP | 521 | Approved | 2007-03-01 | J. Durnan | 1 | Master Clock
Coherent Verdi -laser |  | 480 | Lab 113 | COTS | CV063800036 | Laser Group | BP | 680 | Approved | 2007-03-01 | J. Durnan | 1
MIRA laser controller | 900-F | 480 | Lab 113 | COTS | EM064400022/ EM064800003 | Laser Group | BP | 521 & 524 | Approved | 2007-03-01 | J. Durnan | 2
REG A Controller for laser Amp | 9050 | 480 | Lab 113 | COTS | EM064700001 | Laser Group | BP | 523 | Approved | 2007-03-01 | J. Durnan | 1