# Exercise:	Preparation of reagents and their quality control

- Introduction (on flipchart): 0.5 hr

- various stains used for ZN and fluorescence

ZN hot, original:

basic fuchsin 10 g

phenol 50g

alcohol (denatured or methanol) 150 ml

dissolve fuchsin in heated phenol + alcohol

water up to 1000 ml

ZN cold, Kinyoun stain:

basic fuchsin 25 g

phenol 50 g

alcohol (denatured or methanol) 125 ml

dissolve fuchsin in heated phenol + alcohol

water 625 ml

ZN 10% “saturated” basic fuchsin stock in alcohol

basic fuchsin 10 g

dissolve in alcohol 100 ml

auramine 0.1%:

auramine 1 g (1% stock)

dissolve in in alcohol (denatured or methanol) 100 ml

for 100 ml working solution

stocksolution 10 ml

+ stocksolution phenol 33.3 g in water 1000 ml

sulphuric acid ~20%

water 800 ml (cooled)

slowly add 200 ml sulphuric acid 96%

acid alcohol 1% for auramine staining

alcohol (denatured) 700 ml

hydrochloric acid (fuming) 10 ml

water 290 ml

acid alcohol 3% for ZN

use 1% acid alcohol, 500 ml

add hydrochloric acid (fuming) 10 ml

methylene blue 0.1%

dissolve methylene blue 1 g in 1000 ml water

Gabbett destaining and counterstaining solution

alcohol (denatured) 150 ml

dissolve methylene blue 5 g

water 250 ml

sulphuric acid 100 ml

permanganate 0.1%

stock 5g permangate in 1000 ml water

dilute to 0.1% same day

- water quality

distilled water not absolutely needed

importance of bacteriologically pure water: ultra-filtration

- labeling

each bottle

at least name of reagent and date of preparation

- QC of stains and logbook

as explained in modules

- Prepare stains in 3 groups of 2 persons; 1.5 hr

- demonstration:		prepare “saturated” stock basic fuchsin in alcohol

show amount of sediment left

- first batches: 		carbolfuchsin Kinyoun (at least 100 ml)

compare sediment after dissolving in phenol + alcohol

carbolfuchsin ZN 1 % (at least 1000 ml)

compare sediment after dissolving in phenol + alcohol

auramin stock in alcohol: 100 ml

let stand in dark bottle

separately prepare stock phenol 3.33% in water

- second batches:	sulphuric acid ~20% (500 ml)

Gabbett (500 ml)

acid alcohol 1% (at least 1000 ml) + use half to make 3%

methylene blue 0.1% (1000 ml)

- QC of stains prepared:  1 hr

only staining, reading on day 3

each separately one set

Group 1: hot Ziehl Neelsen

- CF 1% 10 minutes + sulphuric acid 20% 5 minutes + MBlue 0.1% 1/2 to 1 min

- CF 1% 10 minutes + acid alcohol 3% 3 minutes + MBlue 0.1% 1/2 to 1 min

Group 2: cold staining

- Kinyoun 3 minutes + Gabbett 1 minute (Tan Thiam Hok staining)

- Kinyoun 5 minutes + acid alcohol 3% 3 minutes + MBlue 0.1% 1/2 to 1 min

Group 3: fluorescence staining, NO HEATING

- auramin  15 minutes + acid alcohol 1% 3 minutes + MBlue 0.1% 1 minute

- auramin  15 minutes + acid alcohol 1% 3 minutes + 0.1% permanganate 1 minute

each set 2 positive and 2 negative controls

label slides: POS or NEG & STAINING TYPE

examples: 	POS / ZN- Sulph

NEG / KIN- Gab

POS / AUR -Mbl

stain POS controls only once

3 staining cycles for NEG controls

put smears away in boxes

# Exercise:	Use of the microscope (SEARO booklet)

- Brief description of microscope types and how they work: 15 minutes

- brightfield

- fluorescence

- brightfield microscope parts

- Adjustment of a microscope with practice: 45 minutes

- condenser

- diaphragm

- interpupillary distance

- dioptric adjustment

- changing objectives / slides

- lighting (mirror) and filters

- focusing

- Demonstration of fluorescence microscopy: 0.5 hrs

- high positive

- low positive

- negative

- artefacts to be distinguished from AFB: characteristics of AFB in flu

- Reading smears ordinary microscope (ZN): 1 hr

- 1 panel of smears per participants

- for the more experienced: try to quantify, IUATLD/WHO scale