# NOVA Extrusion Pressure Tester

# And First Results

Hans Jostlein								November 2, 2007

### Introduction:

We have completed a new testing device for NOVA extrusions.

It uses water to pressurize 6 inch long samples of the 16-cell NOVA PVC extrusion to failure.

The tester uses 4 hydraulic cylinders to seal the extrusion against rubber mats, and  an electric water pump.  It can measure an extrusion in about 8 minutes time.

### History

We have recently measured 11 extrusion samples with a primitive manual setup.

Here is a picture of the setup:

Test Setup

We used pressurized water to keep the stored energy extremely low.

As result there is not only just a minimal safety risk, but also the extrusions survive the failure event nearly intact, allowing meaningful failure analysis by inspection.

Figure 1 is a drawing of the test device.  A NOνA extrusion RPVC sample, approximately six (6) inches long, was placed between two large aluminum bars, measuring 2 7/8” H × 3 ¾” W × 32” L .  The bottom face of the top aluminum bar was machined creating a channel to allow pressurized water to fill each of the sixteen cells of the sample.  At one end of the channel, a threaded fill hole was drilled through the bar.  At the other end of the channel, a second threaded hole was drilled to accommodate a 0–200 psi pressure gauge and a bleed-off value.

To ensure the ability to maintain internal pressure, two large rectangular gaskets were manufactured.  The gaskets were drilled with sixteen (16) 3/4” holes with positions corresponding to the center of the individual cells (See Figure 2.).  The gaskets were placed between the sample’s top and bottom and the corresponding aluminum bar.  To enable the test device to maintain pressure, bolts, washer, and 1/2” all-thread rods were used to clamp the top and bottom sections securely together.  A picture of the test device is included (See Figure 3.)

Water was placed in the sample cells through a hose connected to a nearby sink.  Air trapped in the cells was bled off.  The test device was then pressurized by a hand pump until failure.  Each test average forty-five (45) minutes in length.  Figure 4 illustrates test results.

The Gasket:

### Example of a failed extrusion:

### Drawbacks and Limitations of the  Manual Test

The manual test proved the concept, but was cumbersome and time consuming.

Several tons of force are required to seal the ends of the extrusion against a rubber mat.

Using threaded rod to provide the clamping force worked, but was quite slow.

The hand pump used for pressurizing the water was supplying only a small rate of flow, preventing it from pressurizing if there were even small leaks.

Lastly, at each sample failure all the water ended up on the floor and had to be mopped up.

It took about an hour for each test, and a large amount of frustration.

### The New Tester

We have completed construction of a  more efficient, pressure test device.

The device is equipped with

--hydraulic cylinders to compress and seal the rubber gaskets

--springs to release the sample after the test

--a closed circuit water system, using an electric pump and a water tank

--valving to allow quick water filling,

followed by  slow and controlled pressurizing through a separate needle valve

### Here is an Assembly Drawing:

### Description of  the New  Tester:

The extrusion is clamped between a top  aluminum bars against the bottom I-beam, using  rubber 1/2” thick gaskets , reinforced with fiberglass mats.

For gasket compression, we use four hydraulic cylinders, all pressurized from the same hand pump.

This assures sufficient force and uniform force distribution over the length of the bar.

The hydraulic pressure is  limited to 5000 psi, and the force on the frame to 40,000#f.

The current pressure gauge has a range to 4000 psi, which provides a maximum force of  32,000 #  f, which appears adequate to achieve  a good seal.

The cylinder forces are resisted by a steel frame made from  steel I-beam of 5 1/2 flange width, 8 3/8 height, with a moment of inertia of 58.7 in^4.

The maximum stress in the I-beam at 40,000 #f  is  11,600 psi.

The steel frame has been  proof loaded to 4000 psi (32,000 #f)  by operating the hydraulic cylinders against cribbing material.

The steel frame sits in an open tank made form 1/16” stainless sheet.

The tank is 36” long by 18” wide by 36” high, and has a hinged half-inch thick plastic cover.

The tank has a bottom drain which connects to a  small circulation pump, which is located below the tank.  The circulation pumps provides the positive pressure required by the high pressure test pump.

The high pressure pump is a Rice EL-1A pump rated at 3 gpm at 500 psi.

The pump outlet (the high pressure hose) connects to the top aluminum bar.

The second hose connection at the top aluminum bar is used for  air bleeding and pressurization control, as follows:

A 10 f t high pressure water hose goes to the operator valve  box, mounted on the operator’s control panel.  (this is not shown on the above diagram)

The operator valve  box has a bleed valve which sends the air (and water, as the case may be) back to the tank sump.

A pressure gauge is included to read the maximum pressure before extrusion failure.

The pump is powered from a GFI  outlet in an outdoor type outlet box.

To pressurize, the operator turns on the water pump while the outlet valve from the extrusion is open.  This flushes the air out of the test extrusion and  fills it  with water.

When no air exits the extrusion (as seen through the transparent drain hose),  the operator   closes the main input valve, and then the output drain valve to the extrusion.   Pressurized water enters the test extrusion at low flow volume through the needle valve,   causing  the test pressure to rise  in the extrusion.  The rate of rise is controlled by this valve.

After the extrusion has failed, the operator turns the water pump off,opens the bleed valve,  and releases the hydraulic seal pressure  in preparation for the next sample.

The springs push the top bar up to free the failed extrusion and to allow insertion of the next sample.

The operator uses the hydraulic hand pump to seal the new sample, closes the tank lid, and proceeds to the operator control box for the next cycle.

### Here are a few pictures of the Pressure tester:

## Initial Results

## We have tested  seven “vertical” extrusion that were  made from NOVA 27-material  in August 2007.  All failed by having the end cell split at mid-level.  This is the highest stress point as per Ang Lee’s FEA calculations, thus indicating no flaws in material, knitting, or shapes.

## The pressure at failure was

## Sample #4	330 psi

## Sample #5	340 psi,

## Sample #6	320 psi,  time to test  4 minutes

## Sample #7	350 psi,  time to test  7 minutes

## Sample #8	380 psi,  time to test  6 minutes

## Sample #9	320 psi,  time to test  6  minutes

## Sample #10	320 psi,  time to test  5 minutes

The pressures at failure were quite similar between samples.

We will next test other samples , including horizontal samples and samples made from problematic PVC runs, to verify that this device can tell the difference between sound and flawed extrusions.

## Appendix:  Details

### The Pump

Rice Hydro EL-1A Electric Hydrostatic Test Pump 3 GPM, 500 PSI SKU EL-1A

| View larger photo | Sale Price: $1,102.00 Original Price: $1,160.00 Usually ships within 1 - 2 weeks Check stock Purchasing options & Accessories below |
| --- | --- |
|  | Picture Gallery: Click on a thumbnail below to view View super-sized pic 1 View super-sized pic 2 |

Features: • Hydrostatic Test Pump, designed for testing sprinkler systems, irrigation lines, pressure vessels, plant process piping, and small diameter water lines • 1 HP Electric Motor, TEFC, 110/220 Volt, 50/60 HZ • Automatic Pressure regulator with inlet bypass • Stainless Steel liquid-filled gauge ensures accurate readings • Rugged roll cage construction, protects pump from damage • Dimensions: 22x14x14, 1 Year Warranty from Manufacture Defect The EL-series is a low volume medium pressure line of hydrostatic test pump designed for testing any small water line or pressure system. This lightweight compact design combines the versatility of indoor/outdoor use along with simple operation. The EL-series test pumps are ideal for testing sprinkler systems, irrigation lines, plant piping and pressure vessels. Standard power sources include, 1 HP, 50/60 HZ, 115 volt, 15 amp electric motor. Thanks to its stainless steel vibration dampening liquid-filled gauge, smooth and accurate readings become the rule, not the exception. Electric motors are available in varieties of phase, voltages and explosion proof models. Specifications 3 GPM 500 PSI Triplex Ceramic Plunger Pump High Pressure Outlet Hose 1/2 HP, Single Phase, TEFC, 115 volt, 50/60 HZ Standard UPS and Fedex transit times: *Upon customer request only. Truck Deliveries Most of our trucking carriers offer Curb Side Delivery to customers location. Curbed Side Delivery means it is the responsibility of the customer who is receiving the shipment to have the means to unload by (fork lift or loading dock) Please note larger items that have free shipping on ground delivery (Cement Mixers, Panel saws, log splitters) Toolfetch pays the shipping to a commercial/business address assuming the location has a fork lift, loading dock or man power to remove the product off the truck. Toolfetch does not pay for Residential fees or lift gate service, this is responsibility of the customer. One of our shipping personal will contact you prior to processing your order with (Residential fees, or liftgate fee if required). Toolfetch offers different options for truck deliveries: Commercial/Business Address: If Toolfetch ships to a commercial address (assuming customer has a forklift or loading dock) Toolfetch will pay the shipping. If liftgate service is required additional fees will be added depending on the freight carrier. A Toolfetch shipping representative will contact you and ask you if you need this additional service. Toolfetch is not responsible if the product cannot be unloaded off the truck. Truck drivers are not deemed responsible to help unload the product off the truck. Please note if you operate your business out of residential address you will be charged for the additional costs of delivering to a residential address Residential address: (Larger Items such as Cement Mixers, Panel Saws, Generators etc..) Trucking carriers impose additional shipping costs to a residential address. The additional fee is due to shipping carriers not being able to use 18 wheelers down residential streets and need to switch to smaller trucks. Toolfetch is not responsible for additional shipping costs to a residential address. Additional costs for residential delivery can range from $30-65 depending on shipping carriers.. Larger items that are heavy over 70 lb. customers should request lift gate service. Liftgate service charges can range from $40 - 100 depending on carrier. A Toolfetch representative will call prior to processing the order and discuss these additional costs If the customer does not request lift gate service Toolfetch is not responsible if the product cannot be removed from the truck. Please note truck drivers are not deemed responsible to help unload the product off the truck .If the driver or customer cannot unload the product it will be brought back to the shipping depot. The customer is responsible for making a 2nd delivery attempt and will be charged accordingly Pickup from Terminal: An alternative to shipping to a residential or commercial/business address is shipping the unit to a shipping terminal. If the customer requests to ship to a terminal we can make arrangements with our shipping carriers for the customer to pickup the unit from the terminal. Toolfetch will determine the closest terminal to the customers location prior to processing the order. Please note you must bring proper identification to the terminal. Products that have the icon free shipping on ground deliveries: Items under 70 lb. (lasers, carts, powertools, buckets, etc...) that have free shipping on ground delivery will be shipped regular UPS ground or FedEx Ground. Items over 70 Ib. (Cement Mixers, Log Splitters, Generators, kettles, rail saws etc...) that have free shipping on ground delivery will be shipped by freight truck. Please note these items can only ship ground service. Toolfetch has free shipping to commercial/business addresses assuming the location has a loading dock or fork lift present to remove the item off the truck. Toolfetch does not pay for additional residential fees or lift gate service (please see residential address section for additional costs). These additional costs are the responsibility of the customer. A Toolfetch representative will call prior to processing the order and ask if the customer is shipping the product to a residential address or commercial address and if you require lift gate service. UPS and FedEx: Most items under 70 pounds ship by UPS or FedEx Please note UPS and FedEx cannot ship to a PO Box. If shipping to Hawaii or Alaska the shipment will be mailed Second day air only Back Orders: If an item is purchased from Toolfetch that is backordered you will be notified via email immediately. We will keep you updated on the backorder status of the product or we will offer substitute products if the product cannot ship in a timely fashion. International Shipping: If you are shipping outside of the United States please contact our international dept at international@toolfetch.com and a Toolfetch representative will contact you back with information. We can also ship to a freight forwarder within the United States that can forward your freight to your location. (Please note freight forwarders cannot be setup by Toolfetch, it is the customers responsibility for this information. Product Reviews Summary (Powered by PowerReviews.com) Be the first to review this product. Write a Review

| Item | Price | Qty. |
| --- | --- | --- |
| Rice Hydro EL-1A Electric Hydrostatic Test Pump 3 GPM, 500 PSI SKU EL-1A | $1,102.00 $1,160.00 |  |

### Hydraulic system

Hand pump:  Choose  P-392 with pressure gauge

| P series | Lightweight Hand Pumps Lightweight and compact design Durable glass-filled nylon reservoir and nylon encapsulated aluminium pump base Two-speed operation to reduce handle strokes Integral 4-way valve on P-842 for operation of double-acting cylinders Large oil capacities to power a wide variety of cylinders and tools Internal pressure relief valve for overload protection Download Catalog page pdf |
| --- | --- |

| Pump Typex | Usable Oil Capacity | Model Number | Pressure Rating** | Oil Displacement/Stroke | Max. Handle Effort | Piston Stroke | Weight |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | bar | cm3 |  |  |  |
|  | cm3 |  | 1st stage | 2nd stage | 1st stage | 2nd stage | kg | mm | kg |
| Single- Speed | 327 | P-141 | - | 700 | - | 0,90 | 32,7 | 12,7 | 2,0 |
|  | 901 | P-391 | - | 700 | - | 2,47 | 38,6 | 25,4 | 4,1 |
| Two- Speed | 327 | P-142* | 13 | 700 | 3,62 | 0,90 | 35,4 | 12,7 | 2,0 |
|  | 901 | P-202 | 13 | 700 | 3,62 | 0,90 | 28,6 | 12,7 | 3,4 |
|  | 901 | P-392* | 13 | 700 | 11,26 | 2,47 | 42,2 | 25,4 | 4,1 |
|  | 2540 | P-802 | 27 | 700 | 39,33 | 2,47 | 43,1 | 25,4 | 8,2 |
|  | 2540 | P-842 | 27 | 700 | 39,33 | 2,47 | 43,1 | 25,4 | 10,0 |

* Available as sets (Cylinder, gauge, couplers, hose and pump)

Choose Cylinders:  RC-102

| Cylinder Capacity | Stroke | Model Number | Cylinder Effective Area | Oil Capacity | Coll. Height | Weight |
| --- | --- | --- | --- | --- | --- | --- |
| ton (kN) | mm |  | cm2 | cm3 | mm | kg |
| 5 (45) | 16 | RC-50 | 6,5 | 10 | 41 | 1 |
|  | 25 | RC-51 | 6,5 | 16 | 110 | 1 |
|  | 76 | RC-53 | 6,5 | 50 | 165 | 1,5 |
|  | 127 | RC-55* | 6,5 | 83 | 215 | 1,9 |
|  | 177 | RC-57 | 6,5 | 115 | 273 | 2,4 |
|  | 232 | RC-59 | 6,5 | 151 | 323 | 2,8 |
| 10(101) | 26 | RC-101 | 14,5 | 38 | 89 | 1,8 |
|  | 54 | RC-102* | 14,5 | 78 | 121 |  |

# Text boxes

Gaskets

Aluminum Bars

Fig. 1: Test Device

Water Fill Line

PVC Sample

Pressure Gauge

Fill Holes for Cells

Holes for 1/2”

All-thread Rod

Test Sample