## Channel Summary Overview
Serial No. | Hybrid_Initial | Hybrid LTT | post fanout | modulization | thermal cycling | Module LTT | General Comments
202200402 | bad channels | bad channels | bad channels | bad channels | bad channels | bad channels
00008 | 0 | Warm:0 Cold:0 | 0 |  |  |  | All tests passed okay, but had a minor NO reading during LTT Warm.
00009 | 0 | Warm:0 Cold:0 | 5 |  |  |  | Hy_In masked 80 channels, still not understood. 4 channels shorted during fanout.
00010 | 1 | Warm:1 Cold:1 | 1 |  |  |  | Channel 542 manually masked due to high offset, but generally okay.
00011 | 0 | Warm:0 Cold:0 | 0 |  |  |  | Chip 11 replaced due to high noise, and significant NO reading during LTT
00012 | 0 | Warm:0 Cold:0 | 0 |  |  |  | All tests passed okay, but 3 inner-row fanout bonds appear popped off.
00013 | 0 | Warm:0 Cold: |  |  |  |  | LTT Warm fails all three digital tests at 33 degrees, sent to Santa Cruz.
00014 | 0 | Warm:0 Cold:0 | 0 |  |  |  | High gain/low offset persist, and a handful of channels have noise above 750.
00015 | 0 | Warm:1 Cold:1 | 1 |  |  |  | Non-uniform gain, 17 "partbonded" on chip 8, minor NO reading during LTT Warm.
00016 | 0 | Warm:0 Cold:0 | 0 |  |  |  | Irregular gain persists, had extremely minor NO readings during LTT Warm.
00017 | 0 | Warm:0 Cold:0
00018 | 0 | Warm:0 Cold:0
00019 | 1 | Warm:2 Cold:1
00020 | 128 | Warm: Cold: |  |  |  |  | TOKEN defect chip1 prevents rest from functioning. When bypassed the rest work.
00021 | 128 | Warm: Cold: |  |  |  |  | Chip 8 shows high and irregular gain and offset. Before trim also vt50 is irregular
00022 | 128 | Warm: Cold: |  |  |  |  | Pre-trim chip10 shows hi offset, increasing NO. If bypassed all chips are normal.
00023 | 1 | Warm:1 Cold:1
00024 | 0 | Warm:1 Cold:0
00025 | 1 | Warm:2 Cold:2 |  |  |  |  | Channel 699 shows higher gain that would deform NO results. Masked and retest
00026 |  | Warm: Cold:
00027 | 0 | Warm: Cold:
00028 |  | Warm: Cold:
00029 | 0 | Warm: Cold:
00030 | 1 | Warm: Cold:
00031 |  | Warm: Cold:
00032 |  | Warm: Cold:
00033 |  | Warm: Cold:
00034 |  | Warm: Cold:
00035 | 2 | Warm: Cold: |  |  |  |  | Fails TimeWalk.
00036 |  | Warm: Cold:
00037 |  | Warm: Cold:
00038 |  | Warm: Cold:
00039 |  | Warm: Cold:
00040 |  | Warm: Cold:
00041 |  | Warm: Cold:
00042 |  | Warm: Cold:
00043 |  | Warm: Cold:
00044 |  | Warm: Cold:
00045 |  | Warm: Cold:
00046 |  | Warm: Cold:
00047 |  | Warm: Cold:
00048 |  | Warm: Cold:
00049 |  | Warm: Cold:
00050 |  | Warm: Cold:
00051 |  | Warm: Cold:
00052 |  | Warm: Cold:
00053 |  | Warm: Cold:
00054 |  | Warm: Cold:
00055 |  | Warm: Cold:
00056 |  | Warm: Cold:
00057 |  | Warm: Cold:
00058 |  | Warm: Cold:
00059 |  | Warm: Cold:
00060 |  | Warm: Cold:
00061 |  | Warm: Cold:
00062 |  | Warm: Cold:
00063 |  | Warm: Cold:
00064 |  | Warm: Cold:
00065 |  | Warm: Cold:
00066 |  | Warm: Cold:
00067 |  | Warm: Cold:

## Channel Map
Stream 0
Chip 0 | channels 0 - 127
Chip 1 | channels 128 - 255
Chip 2 | channels 256 - 383
Chip 3 | channels 384 - 511
Chip 4 | channels 512 - 639
Chip 5 | channels 640 - 767
Stream 1
Chip 6 | channels 768 - 895
Chip 7 | channels 896 - 1023
Chip 8 | channels 1024 - 1151
Chip 9 | channels 1152 - 1279
Chip 10 | channels 1280 - 1407
Chip 11 | channels 1408 - 1535

## Wafer Lot Info
20220040200008 |  |  | cap correction | 20220040200009 |  |  | cap correction | 20220040200010 |  |  | cap correction
Chip 0 | Z40859-W11-1 | 16 | 1.044 | Chip 0 | Z40859-W11-1 | 136 | 1.044 | Chip 0 | Z40800-W13-1 | 73 | 1.002
Chip 1 | Z40859-W11-1 | 32 | 1.044 | Chip 1 | Z40859-W11-1 | 140 | 1.044 | Chip 1 | Z40800-W13-1 | 77 | 1.002
Chip 2 | Z40859-W11-1 | 48 | 1.044 | Chip 2 | Z40859-W11-1 | 142 | 1.044 | Chip 2 | Z40800-W13-1 | 91 | 1.002
Chip 3 | Z40859-W11-1 | 51 | 1.044 | Chip 3 | Z40859-W11-1 | 143 | 1.044 | Chip 3 | Z40800-W13-1 | 94 | 1.002
Chip 4 | Z40859-W11-1 | 54 | 1.044 | Chip 4 | Z40859-W11-1 | 147 | 1.044 | Chip 4 | Z40800-W13-1 | 128 | 1.002
Chip 5 | Z40859-W11-1 | 67 | 1.044 | Chip 5 | Z40859-W11-1 | 161 | 1.044 | Chip 5 | Z40800-W13-1 | 143 | 1.002
Chip 6 | Z40859-W11-1 | 80 | 1.044 | Chip 6 | Z40859-W11-1 | 163 | 1.044 | Chip 6 | Z40800-W13-1 | 168 | 1.002
Chip 7 | Z40859-W11-1 | 91 | 1.044 | Chip 7 | Z40859-W11-1 | 165 | 1.044 | Chip 7 | Z40800-W13-1 | 189 | 1.002
Chip 8 | Z40859-W11-1 | 108 | 1.044 | Chip 8 | Z40859-W11-1 | 178 | 1.044 | Chip 8 | Z40800-W13-1 | 193 | 1.002
Chip 9 | Z40859-W11-1 | 112 | 1.044 | Chip 9 | Z40859-W11-1 | 189 | 1.044 | Chip 9 | Z40800-W13-1 | 194 | 1.002
Chip10 | Z40859-W11-1 | 116 | 1.044 | Chip10 | Z40859-W11-1 | 192 | 1.044 | Chip10 | Z40800-W13-1 | 195 | 1.002
Chip 11 | Z40859-W11-1 | 126 | 1.044 | Chip 11 | Z40859-W11-1 | 194 | 1.044 | Chip 11 | Z40800-W13-1 | 196 | 1.002
20220040200011 |  |  | cap correction | 20220040200012 |  |  | cap correction | 20220040200013 |  |  | cap correction
Chip 0 | Z40859-W11-1 | 195 | 1.044 | Chip 0 | Z40800-W13-1 | 201 | 1.002 | Chip 0 | Z40859-W02-1 | 12 | 1.044
Chip 1 | Z40859-W11-1 | 197 | 1.044 | Chip 1 | Z40800-W13-1 | 209 | 1.002 | Chip 1 | Z40859-W02-1 | 18 | 1.044
Chip 2 | Z40859-W11-1 | 205 | 1.044 | Chip 2 | Z40800-W13-1 | 211 | 1.002 | Chip 2 | Z40859-W02-1 | 20 | 1.044
Chip 3 | Z40859-W11-1 | 212 | 1.044 | Chip 3 | Z40800-W13-1 | 212 | 1.002 | Chip 3 | Z40859-W02-1 | 28 | 1.044
Chip 4 | Z40859-W11-1 | 223 | 1.044 | Chip 4 | Z40800-W13-1 | 218 | 1.002 | Chip 4 | Z40859-W02-1 | 32 | 1.044
Chip 5 | Z40859-W11-1 | 229 | 1.044 | Chip 5 | Z40800-W13-1 | 228 | 1.002 | Chip 5 | Z40859-W02-1 | 38 | 1.044
Chip 6 | Z40859-W11-1 | 231 | 1.044 | Chip 6 | Z40800-W13-1 | 233 | 1.002 | Chip 6 | Z40800-W13-1 | 255 | 1.044
Chip 7 | Z40859-W11-1 | 232 | 1.044 | Chip 7 | Z40800-W13-1 | 234 | 1.002 | Chip 7 | Z40859-W02-1 | 53 | 1.044
Chip 8 | Z40859-W11-1 | 233 | 1.044 | Chip 8 | Z40800-W13-1 | 248 | 1.002 | Chip 8 | Z40859-W02-1 | 56 | 1.044
Chip 9 | Z40859-W11-1 | 242 | 1.044 | Chip 9 | Z40800-W13-1 | 250 | 1.002 | Chip 9 | Z40859-W02-1 | 57 | 1.044
Chip10 | Z40859-W11-1 | 244 | 1.044 | Chip10 | Z40800-W13-1 | 252 | 1.002 | Chip10 | Z40859-W02-1 | 73 | 1.044
Chip 11 | Z40859-W11-1 | 253 | 1.044 | Chip 11 | Z40800-W13-1 | 253 | 1.002 | Chip 11 | Z40859-W02-1 | 83 | 1.044
20220040200014 |  |  | cap correction | 20220040200015 |  |  | cap correction | 20220040200016 |  |  | cap correction
Chip 0 | Z40859-W02-1 | 84 | 1.044 | Chip 0 | Z40859-W02-2 | 218 | 1.044 | Chip 0 | Z40859-W02-1 | 181 | 1.044
Chip 1 | Z40859-W02-1 | 90 | 1.044 | Chip 1 | Z40859-W02-2 | 220 | 1.044 | Chip 1 | Z40859-W02-1 | 183 | 1.044
Chip 2 | Z40859-W02-1 | 93 | 1.044 | Chip 2 | Z40859-W02-2 | 222 | 1.044 | Chip 2 | Z40859-W02-1 | 188 | 1.044
Chip 3 | Z40859-W02-1 | 98 | 1.044 | Chip 3 | Z40859-W02-2 | 228 | 1.044 | Chip 3 | Z40859-W02-1 | 191 | 1.044
Chip 4 | Z40859-W02-1 | 100 | 1.044 | Chip 4 | Z40859-W02-2 | 229 | 1.044 | Chip 4 | Z40859-W02-1 | 193 | 1.044
Chip 5 | Z40859-W02-1 | 101 | 1.044 | Chip 5 | Z40859-W02-2 | 231 | 1.044 | Chip 5 | Z40859-W02-1 | 196 | 1.044
Chip 6 | Z40859-W02-1 | 103 | 1.044 | Chip 6 | Z40859-W02-2 | 232 | 1.044 | Chip 6 | Z40859-W02-1 | 201 | 1.044
Chip 7 | Z40859-W02-1 | 105 | 1.044 | Chip 7 | Z40859-W02-2 | 234 | 1.044 | Chip 7 | Z40859-W02-1 | 204 | 1.044
Chip 8 | Z40859-W02-1 | 112 | 1.044 | Chip 8 | Z40859-W02-2 | 235 | 1.044 | Chip 8 | Z40859-W02-1 | 205 | 1.044
Chip 9 | Z40859-W02-1 | 123 | 1.044 | Chip 9 | Z40859-W02-2 | 236 | 1.044 | Chip 9 | Z40859-W02-1 | 208 | 1.044
Chip10 | Z40859-W02-1 | 127 | 1.044 | Chip10 | Z40859-W02-2 | 237 | 1.044 | Chip10 | Z40859-W02-1 | 216 | 1.044
Chip 11 | Z40859-W02-1 | 130 | 1.044 | Chip 11 | Z40859-W02-2 | 239 | 1.044 | Chip 11 | Z40859-W02-1 | 217 | 1.044
20220040200017 |  |  | cap correction | 20220040200018 |  |  | cap correction | 20220040200019 |  |  | cap correction
Chip 0 | Z40859-W02-1 | 138 | 1.044 | Chip 0 | Z40859-W09-1 | 10 | 1.044 | Chip 0 | Z40859-W09-1 | 72 | 1.044
Chip 1 | Z40859-W02-1 | 143 | 1.044 | Chip 1 | Z40859-W09-1 | 21 | 1.044 | Chip 1 | Z40859-W09-1 | 73 | 1.044
Chip 2 | Z40859-W02-1 | 146 | 1.044 | Chip 2 | Z40859-W09-1 | 33 | 1.044 | Chip 2 | Z40859-W09-1 | 74 | 1.044
Chip 3 | Z40859-W02-1 | 148 | 1.044 | Chip 3 | Z40859-W09-1 | 36 | 1.044 | Chip 3 | Z40859-W09-1 | 75 | 1.044
Chip 4 | Z40859-W02-1 | 149 | 1.044 | Chip 4 | Z40859-W09-1 | 38 | 1.044 | Chip 4 | Z40859-W09-1 | 81 | 1.044
Chip 5 | Z40859-W02-1 | 159 | 1.044 | Chip 5 | Z40859-W09-1 | 40 | 1.044 | Chip 5 | Z40859-W09-1 | 82 | 1.044
Chip 6 | Z40859-W02-1 | 162 | 1.044 | Chip 6 | Z40859-W09-1 | 44 | 1.044 | Chip 6 | Z40859-W09-1 | 84 | 1.044
Chip 7 | Z40859-W02-1 | 165 | 1.044 | Chip 7 | Z40859-W09-1 | 61 | 1.044 | Chip 7 | Z40859-W09-1 | 88 | 1.044
Chip 8 | Z40859-W02-1 | 166 | 1.044 | Chip 8 | Z40859-W09-1 | 64 | 1.044 | Chip 8 | Z40859-W09-1 | 89 | 1.044
Chip 9 | Z40859-W02-1 | 176 | 1.044 | Chip 9 | Z40859-W09-1 | 65 | 1.044 | Chip 9 | Z40859-W09-1 | 91 | 1.044
Chip10 | Z40859-W02-1 | 178 | 1.044 | Chip10 | Z40859-W09-1 | 68 | 1.044 | Chip10 | Z40859-W09-1 | 94 | 1.044
Chip 11 | Z40859-W02-1 | 179 | 1.044 | Chip 11 | Z40859-W09-1 | 69 | 1.044 | Chip 11 | Z40859-W09-1 | 96 | 1.044
20220040200020 |  |  | cap correction | 20220040200021 |  |  | cap correction | 20220040200022 |  |  | cap correction
Chip 0 | Z40859-W014-1 | 2 | 1.044 | Chip 0 | Z40859-W09-1 | 100 | 1.044 | Chip 0 | Z40859-W09-1 | 161 | 1.044
Chip 1 | Z40859-W02-2 | 254 | 1.044 | Chip 1 | Z40859-W09-1 | 115 | 1.044 | Chip 1 | Z40859-W09-1 | 165 | 1.044
Chip 2 | Z40859-W014-1 | 7 | 1.044 | Chip 2 | Z40859-W09-1 | 116 | 1.044 | Chip 2 | Z40859-W09-1 | 167 | 1.044
Chip 3 | Z40859-W014-1 | 8 | 1.044 | Chip 3 | Z40859-W09-1 | 119 | 1.044 | Chip 3 | Z40859-W09-1 | 172 | 1.044
Chip 4 | Z40859-W014-1 | 9 | 1.044 | Chip 4 | Z40859-W09-1 | 120 | 1.044 | Chip 4 | Z40859-W09-1 | 175 | 1.044
Chip 5 | Z40859-W014-1 | 14 | 1.044 | Chip 5 | Z40859-W09-1 | 129 | 1.044 | Chip 5 | Z40859-W09-1 | 185 | 1.044
Chip 6 | Z40859-W014-1 | 29 | 1.044 | Chip 6 | Z40859-W09-1 | 130 | 1.044 | Chip 6 | Z40859-W09-1 | 187 | 1.044
Chip 7 | Z40859-W014-1 | 30 | 1.044 | Chip 7 | Z40859-W09-1 | 132 | 1.044 | Chip 7 | Z40859-W09-1 | 192 | 1.044
Chip 8 | Z40859-W014-1 | 35 | 1.044 | Chip 8 | Z40859-W09-1 | 136 | 1.044 | Chip 8 | Z40859-W09-1 | 194 | 1.044
Chip 9 | Z40859-W014-1 | 73 | 1.044 | Chip 9 | Z40859-W09-1 | 141 | 1.044 | Chip 9 | Z40859-W09-1 | 195 | 1.044
Chip10 | Z40859-W014-1 | 96 | 1.044 | Chip10 | Z40859-W09-1 | 142 | 1.044 | Chip10 | Z40859-W09-1 | 196 | 1.044
Chip 11 | Z40859-W014-1 | 102 | 1.044 | Chip 11 | Z40859-W09-1 | 159 | 1.044 | Chip 11 | Z40859-W09-1 | 197 | 1.044
20220040200023 |  |  | cap correction | 20220040200024 |  |  | cap correction | 20220040200025 |  |  | cap correction
Chip 0 | Z40859-W14-1 | 233 | 1.044 | Chip 0 | Z40859-W14-1 | 192 | 1.044 | Chip 0 | Z40859-W14-1 | 103 | 1.044
Chip 1 | Z40859-W14-1 | 234 | 1.044 | Chip 1 | Z40859-W14-1 | 195 | 1.044 | Chip 1 | Z40859-W14-1 | 107 | 1.044
Chip 2 | Z40859-W14-1 | 238 | 1.044 | Chip 2 | Z40859-W14-1 | 198 | 1.044 | Chip 2 | Z40859-W14-1 | 111 | 1.044
Chip 3 | Z40859-W14-1 | 239 | 1.044 | Chip 3 | Z40859-W14-1 | 200 | 1.044 | Chip 3 | Z40859-W14-1 | 112 | 1.044
Chip 4 | Z40859-W14-1 | 241 | 1.044 | Chip 4 | Z40859-W14-1 | 207 | 1.044 | Chip 4 | Z40859-W14-1 | 130 | 1.044
Chip 5 | Z40859-W14-1 | 244 | 1.044 | Chip 5 | Z40859-W14-1 | 209 | 1.044 | Chip 5 | Z40859-W14-1 | 134 | 1.044
Chip 6 | Z40859-W14-1 | 245 | 1.044 | Chip 6 | Z40859-W14-1 | 212 | 1.044 | Chip 6 | Z40859-W14-1 | 143 | 1.044
Chip 7 | Z40859-W14-1 | 248 | 1.044 | Chip 7 | Z40859-W14-1 | 220 | 1.044 | Chip 7 | Z40859-W14-1 | 150 | 1.044
Chip 8 | Z40859-W14-1 | 249 | 1.044 | Chip 8 | Z40859-W14-1 | 221 | 1.044 | Chip 8 | Z40859-W14-1 | 161 | 1.044
Chip 9 | Z40859-W14-1 | 250 | 1.044 | Chip 9 | Z40859-W14-1 | 228 | 1.044 | Chip 9 | Z40859-W14-1 | 166 | 1.044
Chip10 | Z40859-W14-1 | 253 | 1.044 | Chip10 | Z40859-W14-1 | 231 | 1.044 | Chip10 | Z40859-W14-1 | 178 | 1.044
Chip 11 | Z40859-W14-1 | 255 | 1.044 | Chip 11 | Z40859-W14-1 | 232 | 1.044 | Chip 11 | Z40859-W14-1 | 183 | 1.044
20220040200026 |  |  | cap correction | 20220040200027 |  |  | cap correction | 20220040200028 |  |  | cap correction
Chip 0 | Z40859-W03-2 | 223 | 1.044 | Chip 0 | Z40859-W01-1 | 66 | 1.044 | Chip 0 | Z40859-W03-1 | 51 | 1.044
Chip 1 | Z40859-W03-2 | 231 | 1.044 | Chip 1 | Z40859-W01-1 | 67 | 1.044 | Chip 1 | Z40859-W03-1 | 53 | 1.044
Chip 2 | Z40859-W03-2 | 236 | 1.044 | Chip 2 | Z40859-W01-1 | 69 | 1.044 | Chip 2 | Z40859-W03-1 | 54 | 1.044
Chip 3 | Z40859-W03-2 | 239 | 1.044 | Chip 3 | Z40859-W01-1 | 73 | 1.044 | Chip 3 | Z40859-W03-1 | 57 | 1.044
Chip 4 | Z40859-W03-2 | 242 | 1.044 | Chip 4 | Z40859-W01-1 | 77 | 1.044 | Chip 4 | Z40859-W03-1 | 69 | 1.044
Chip 5 | Z40859-W03-2 | 244 | 1.044 | Chip 5 | Z40859-W01-1 | 84 | 1.044 | Chip 5 | Z40859-W03-1 | 78 | 1.044
Chip 6 | Z40859-W03-2 | 246 | 1.044 | Chip 6 | Z40859-W01-1 | 88 | 1.044 | Chip 6 | Z40859-W03-1 | 84 | 1.044
Chip 7 | Z40859-W03-2 | 247 | 1.044 | Chip 7 | Z40859-W01-1 | 94 | 1.044 | Chip 7 | Z40859-W03-1 | 87 | 1.044
Chip 8 | Z40859-W03-2 | 251 | 1.044 | Chip 8 | Z40859-W01-1 | 101 | 1.044 | Chip 8 | Z40859-W03-1 | 89 | 1.044
Chip 9 | Z40859-W03-2 | 253 | 1.044 | Chip 9 | Z40859-W01-1 | 112 | 1.044 | Chip 9 | Z40859-W03-1 | 94 | 1.044
Chip10 | Z40859-W03-2 | 254 | 1.044 | Chip10 | Z40859-W01-1 | 121 | 1.044 | Chip10 | Z40859-W03-1 | 101 | 1.044
Chip 11 | Z40859-W03-2 | 255 | 1.044 | Chip 11 | Z40859-W01-1 | 124 | 1.044 | Chip 11 | Z40859-W03-1 | 106 | 1.044
20220040200029 |  |  | cap correction | 20220040200030 |  |  | cap correction | 20220040200031 |  |  | cap correction
Chip 0 | Z40859-W01-1 | 196 | 1.044 | Chip 0 | Z40859-W03-1 | 112 | 1.044 | Chip 0 | Z40859-W03-1 | 173 | 1.044
Chip 1 | Z40859-W01-1 | 202 | 1.044 | Chip 1 | Z40859-W03-1 | 113 | 1.044 | Chip 1 | Z40859-W03-1 | 178 | 1.044
Chip 2 | Z40859-W01-1 | 205 | 1.044 | Chip 2 | Z40859-W03-1 | 124 | 1.044 | Chip 2 | Z40859-W03-1 | 187 | 1.044
Chip 3 | Z40859-W01-1 | 210 | 1.044 | Chip 3 | Z40859-W03-1 | 128 | 1.044 | Chip 3 | Z40859-W03-1 | 190 | 1.044
Chip 4 | Z40859-W01-1 | 217 | 1.044 | Chip 4 | Z40859-W03-1 | 129 | 1.044 | Chip 4 | Z40859-W03-1 | 193 | 1.044
Chip 5 | Z40859-W01-1 | 222 | 1.044 | Chip 5 | Z40859-W03-1 | 133 | 1.044 | Chip 5 | Z40859-W03-1 | 195 | 1.044
Chip 6 | Z40859-W01-1 | 226 | 1.044 | Chip 6 | Z40859-W03-1 | 142 | 1.044 | Chip 6 | Z40859-W03-1 | 197 | 1.044
Chip 7 | Z40859-W01-1 | 229 | 1.044 | Chip 7 | Z40859-W03-1 | 156 | 1.044 | Chip 7 | Z40859-W03-1 | 204 | 1.044
Chip 8 | Z40859-W01-1 | 230 | 1.044 | Chip 8 | Z40859-W03-1 | 158 | 1.044 | Chip 8 | Z40859-W03-1 | 205 | 1.044
Chip 9 | Z40859-W01-1 | 233 | 1.044 | Chip 9 | Z40859-W03-1 | 159 | 1.044 | Chip 9 | Z40859-W03-1 | 210 | 1.044
Chip10 | Z40859-W01-1 | 237 | 1.044 | Chip10 | Z40859-W03-1 | 160 | 1.044 | Chip10 | Z40859-W03-1 | 215 | 1.044
Chip 11 | Z40859-W01-1 | 242 | 1.044 | Chip 11 | Z40859-W03-1 | 168 | 1.044 | Chip 11 | Z40859-W03-1 | 217 | 1.044
20220040200032 |  |  | cap correction | 20220040200033 |  |  | cap correction | 20220040200034 |  |  | cap correction
Chip 0 | Z40859-W04-1 | 78 | 1.044 | Chip 0 | Z40859-W02-2 | 250 | 1.044 | Chip 0 | Z40859-W04-1 | 25 | 1.044
Chip 1 | Z40859-W04-1 | 83 | 1.044 | Chip 1 | Z40859-W01-1 | 3 | 1.044 | Chip 1 | Z40859-W02-2 | 247 | 1.044
Chip 2 | Z40859-W04-1 | 84 | 1.044 | Chip 2 | Z40859-W01-1 | 9 | 1.044 | Chip 2 | Z40859-W04-1 | 28 | 1.044
Chip 3 | Z40859-W04-1 | 87 | 1.044 | Chip 3 | Z40859-W01-1 | 24 | 1.044 | Chip 3 | Z40859-W04-1 | 31 | 1.044
Chip 4 | Z40859-W04-1 | 90 | 1.044 | Chip 4 | Z40859-W01-1 | 25 | 1.044 | Chip 4 | Z40859-W04-1 | 35 | 1.044
Chip 5 | Z40859-W04-1 | 99 | 1.044 | Chip 5 | Z40859-W01-1 | 28 | 1.044 | Chip 5 | Z40859-W04-1 | 38 | 1.044
Chip 6 | Z40859-W04-1 | 100 | 1.044 | Chip 6 | Z40859-W01-1 | 31 | 1.044 | Chip 6 | Z40859-W04-1 | 46 | 1.044
Chip 7 | Z40859-W04-1 | 102 | 1.044 | Chip 7 | Z40859-W01-1 | 37 | 1.044 | Chip 7 | Z40859-W04-1 | 54 | 1.044
Chip 8 | Z40859-W04-1 | 109 | 1.044 | Chip 8 | Z40859-W01-1 | 40 | 1.044 | Chip 8 | Z40859-W04-1 | 56 | 1.044
Chip 9 | Z40859-W04-1 | 112 | 1.044 | Chip 9 | Z40859-W01-1 | 43 | 1.044 | Chip 9 | Z40859-W04-1 | 70 | 1.044
Chip10 | Z40859-W04-1 | 114 | 1.044 | Chip10 | Z40859-W01-1 | 50 | 1.044 | Chip10 | Z40859-W04-1 | 73 | 1.044
Chip 11 | Z40859-W04-1 | 120 | 1.044 | Chip 11 | Z40859-W01-1 | 56 | 1.044 | Chip 11 | Z40859-W04-1 | 74 | 1.044
20220040200035 |  |  | cap correction | 20220040200036 |  |  | cap correction | 20220040200037 |  |  | cap correction
Chip 0 | Z40859-W01-1 | 127 | 1.044 | Chip 0 | Z40859-W03-1 | 6 | 1.044 | Chip 0 | Z40859-W04-1 | 165 | 1.044
Chip 1 | Z40859-W01-1 | 140 | 1.044 | Chip 1 | Z40859-W03-1 | 11 | 1.044 | Chip 1 | Z40859-W04-1 | 177 | 1.044
Chip 2 | Z40859-W01-1 | 142 | 1.044 | Chip 2 | Z40859-W03-1 | 15 | 1.044 | Chip 2 | Z40859-W04-1 | 182 | 1.044
Chip 3 | Z40859-W01-1 | 148 | 1.044 | Chip 3 | Z40859-W03-1 | 17 | 1.044 | Chip 3 | Z40859-W04-1 | 185 | 1.044
Chip 4 | Z40859-W01-1 | 156 | 1.044 | Chip 4 | Z40859-W03-1 | 21 | 1.044 | Chip 4 | Z40859-W04-1 | 191 | 1.044
Chip 5 | Z40859-W01-1 | 164 | 1.044 | Chip 5 | Z40859-W03-1 | 25 | 1.044 | Chip 5 | Z40859-W04-1 | 192 | 1.044
Chip 6 | Z40859-W01-1 | 171 | 1.044 | Chip 6 | Z40859-W03-1 | 27 | 1.044 | Chip 6 | Z40859-W04-1 | 194 | 1.044
Chip 7 | Z40859-W01-1 | 173 | 1.044 | Chip 7 | Z40859-W02-2 | 251 | 1.044 | Chip 7 | Z40859-W04-1 | 195 | 1.044
Chip 8 | Z40859-W01-1 | 177 | 1.044 | Chip 8 | Z40859-W03-1 | 32 | 1.044 | Chip 8 | Z40859-W04-1 | 201 | 1.044
Chip 9 | Z40859-W02-2 | 253 | 1.044 | Chip 9 | Z40859-W03-1 | 41 | 1.044 | Chip 9 | Z40859-W04-1 | 203 | 1.044
Chip10 | Z40859-W01-1 | 191 | 1.044 | Chip10 | Z40859-W03-1 | 42 | 1.044 | Chip10 | Z40859-W04-1 | 204 | 1.044
Chip 11 | Z40859-W01-1 | 194 | 1.044 | Chip 11 | Z40859-W03-1 | 44 | 1.044 | Chip 11 | Z40859-W04-1 | 205 | 1.044
20220040200038 |  |  | cap correction | 20220040200039 |  |  | cap correction | 20220040200040 |  |  | cap correction
Chip 0 | Z40859-W04-1 | 123 | 1.044 | Chip 0 | Z40859-W04-2 | 207 | 1.044 | Chip 0 | Z40859-W01-2 | 243 | 1.044
Chip 1 | Z40859-W04-1 | 127 | 1.044 | Chip 1 | Z40859-W04-2 | 209 | 1.044 | Chip 1 | Z40859-W01-2 | 244 | 1.044
Chip 2 | Z40859-W04-1 | 128 | 1.044 | Chip 2 | Z40859-W04-2 | 223 | 1.044 | Chip 2 | Z40859-W01-2 | 245 | 1.044
Chip 3 | Z40859-W04-1 | 134 | 1.044 | Chip 3 | Z40859-W04-2 | 232 | 1.044 | Chip 3 | Z40859-W01-2 | 246 | 1.044
Chip 4 | Z40859-W04-1 | 136 | 1.044 | Chip 4 | Z40859-W04-2 | 236 | 1.044 | Chip 4 | Z40859-W01-2 | 247 | 1.044
Chip 5 | Z40859-W04-1 | 140 | 1.044 | Chip 5 | Z40859-W04-2 | 237 | 1.044 | Chip 5 | Z40859-W01-2 | 250 | 1.044
Chip 6 | Z40859-W04-1 | 144 | 1.044 | Chip 6 | Z40859-W04-2 | 240 | 1.044 | Chip 6 | Z40859-W01-2 | 251 | 1.044
Chip 7 | Z40859-W04-1 | 146 | 1.044 | Chip 7 | Z40859-W04-2 | 255 | 1.044 | Chip 7 | Z40859-W01-2 | 256 | 1.044
Chip 8 | Z40859-W04-1 | 160 | 1.044 | Chip 8 | Z40859-W09-2 | 204 | 1.044 | Chip 8 | Z40859-W02-2 | 240 | 1.044
Chip 9 | Z40859-W04-1 | 161 | 1.044 | Chip 9 | Z40859-W04-2 | 207 | 1.044 | Chip 9 | Z40859-W02-2 | 242 | 1.044
Chip10 | Z40859-W04-1 | 164 | 1.044 | Chip10 | Z40859-W04-2 | 210 | 1.044 | Chip10 | Z40859-W02-2 | 246 | 1.044
Chip 11 | Z40859-W04-2 | 228 | 1.044 | Chip 11 | Z40859-W04-2 | 212 | 1.044 | Chip 11 | Z40859-W02-2 | 256 | 1.044
20220040200041 |  |  | cap correction | 20220040200042 |  |  | cap correction | 20220040200043 |  |  | cap correction
Chip 0 | Z40802-W09-1 | 23 | 1.0038 | Chip 0 | Z40802-W09-1 | 96 | 1.0038 | Chip 0 | Z40802-W14-1 | 11 | 1.0038
Chip 1 | Z40802-W09-1 | 28 | 1.0038 | Chip 1 | Z40802-W09-1 | 103 | 1.0038 | Chip 1 | Z40802-W14-1 | 18 | 1.0038
Chip 2 | Z40802-W09-1 | 30 | 1.0038 | Chip 2 | Z40802-W09-1 | 107 | 1.0038 | Chip 2 | Z40802-W14-1 | 38 | 1.0038
Chip 3 | Z40802-W09-1 | 40 | 1.0038 | Chip 3 | Z40802-W09-1 | 118 | 1.0038 | Chip 3 | Z40802-W14-1 | 40 | 1.0038
Chip 4 | Z40802-W09-1 | 41 | 1.0038 | Chip 4 | Z40802-W09-1 | 120 | 1.0038 | Chip 4 | Z40802-W14-1 | 41 | 1.0038
Chip 5 | Z40802-W09-1 | 54 | 1.0038 | Chip 5 | Z40802-W09-1 | 134 | 1.0038 | Chip 5 | Z40802-W14-1 | 43 | 1.0038
Chip 6 | Z40802-W09-1 | 58 | 1.0038 | Chip 6 | Z40802-W09-1 | 139 | 1.0038 | Chip 6 | Z40802-W14-1 | 63 | 1.0038
Chip 7 | Z40802-W09-1 | 59 | 1.0038 | Chip 7 | Z40802-W09-1 | 141 | 1.0038 | Chip 7 | Z40802-W14-1 | 67 | 1.0038
Chip 8 | Z40802-W09-1 | 67 | 1.0038 | Chip 8 | Z40802-W09-1 | 159 | 1.0038 | Chip 8 | Z40802-W14-1 | 80 | 1.0038
Chip 9 | Z40802-W09-1 | 68 | 1.0038 | Chip 9 | Z40802-W09-1 | 169 | 1.0038 | Chip 9 | Z40802-W14-1 | 95 | 1.0038
Chip10 | Z40802-W09-1 | 71 | 1.0038 | Chip10 | Z40802-W09-1 | 173 | 1.0038 | Chip10 | Z40802-W14-1 | 104 | 1.0038
Chip 11 | Z40802-W09-1 | 95 | 1.0038 | Chip 11 | Z40802-W09-1 | 176 | 1.0038 | Chip 11 | Z40802-W14-1 | 126 | 1.0038
20220040200044 |  |  | cap correction | 20220040200045 |  |  | cap correction | 20220040200046 |  |  | cap correction
Chip 0 | Z40802-W14-1 | 19 | 1.0038 | Chip 0 | Z40802-w14-1 | 204 | 1.0038 | Chip 0
Chip 1 | Z40802-W14-1 | 115 | 1.0038 | Chip 1 | Z40802-w14-1 | 205 | 1.0038 | Chip 1
Chip 2 | Z40802-W14-1 | 132 | 1.0038 | Chip 2 | Z40802-w14-1 | 210 | 1.0038 | Chip 2
Chip 3 | Z40802-W14-1 | 134 | 1.0038 | Chip 3 | Z40802-w14-1 | 212 | 1.0038 | Chip 3
Chip 4 | Z40802-W14-1 | 136 | 1.0038 | Chip 4 | Z40802-w14-1 | 234 | 1.0038 | Chip 4
Chip 5 | Z40802-W14-1 | 141 | 1.0038 | Chip 5 | Z40802-w14-1 | 238 | 1.0038 | Chip 5
Chip 6 | Z40802-W14-1 | 153 | 1.0038 | Chip 6 | Z40802-w14-1 | 256 | 1.0038 | Chip 6
Chip 7 | Z40802-W14-1 | 159 | 1.0038 | Chip 7 | Z40802-W09-1 | 115 | 1.0038 | Chip 7
Chip 8 | Z40802-W14-1 | 174 | 1.0038 | Chip 8 | Z40802-W09-1 | 140 | 1.0038 | Chip 8
Chip 9 | Z40802-W14-1 | 178 | 1.0038 | Chip 9 | Z40802-W09-1 | 204 | 1.0038 | Chip 9
Chip10 | Z40802-W14-1 | 194 | 1.0038 | Chip10 | Z40802-W09-1 | 206 | 1.0038 | Chip10
Chip 11 | Z40802-W14-1 | 196 | 1.0038 | Chip 11 | Z40802-W09-1 | 232 | 1.0038 | Chip 11
20220040200047 |  |  | cap correction | 20220040200048 |  |  | cap correction | 20220040200049 |  |  | cap correction
Chip 0 | Z40862-W11-1 | 10 | 1.0561 | Chip 0 |  |  |  | Chip 0 | Z40862-W11-1 | 69 | 1.0561
Chip 1 | Z40862-W11-1 | 18 | 1.0561 | Chip 1 |  |  |  | Chip 1 | Z40862-W11-1 | 73 | 1.0561
Chip 2 | Z40862-W11-1 | 20 | 1.0561 | Chip 2 |  |  |  | Chip 2 | Z40862-W11-1 | 88 | 1.0561
Chip 3 | Z40862-W11-1 | 31 | 1.0561 | Chip 3 |  |  |  | Chip 3 | Z40862-W11-1 | 101 | 1.0561
Chip 4 | Z40862-W11-1 | 39 | 1.0561 | Chip 4 |  |  |  | Chip 4 | Z40862-W11-1 | 102 | 1.0561
Chip 5 | Z40862-W11-1 | 41 | 1.0561 | Chip 5 |  |  |  | Chip 5 | Z40862-W11-1 | 111 | 1.0561
Chip 6 | Z40862-W11-1 | 42 | 1.0561 | Chip 6 |  |  |  | Chip 6 | Z40862-W11-1 | 124 | 1.0561
Chip 7 | Z40862-W11-1 | 45 | 1.0561 | Chip 7 |  |  |  | Chip 7 | Z40862-W11-1 | 126 | 1.0561
Chip 8 | Z40862-W11-1 | 46 | 1.0561 | Chip 8 |  |  |  | Chip 8 | Z40862-W11-1 | 130 | 1.0561
Chip 9 | Z40862-W11-1 | 51 | 1.0561 | Chip 9 |  |  |  | Chip 9 | Z40862-W11-1 | 131 | 1.0561
Chip10 | Z40862-W11-1 | 55 | 1.0561 | Chip10 |  |  |  | Chip10 | Z40862-W11-1 | 134 | 1.0561
Chip 11 | Z40862-W11-1 | 68 | 1.0561 | Chip 11 |  |  |  | Chip 11 | Z40862-W11-1 | 138 | 1.0561

## Summary
| Serial # | ok | Chip repl | Channels Noisy/dead (Hybrid) | Negative Offset masked | High Offset (>50) masked | LTT |  | ok | channels lost after fanout | Total bad channels (Hybrid) | ok | Module | ok | max LC at 150C 500V before TC (mA) | max LC at 150C 500V after TC + Metr (mA) | max LC at 00C 500V after LTT (mA) | Noisy/dead | Unbonded | Total bad channels | TESTS Completed
[7 empty] | cold | warm
1 | 20220040200008 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1 | P2 | 1 | 0.963 | 0.922 | 0.092 | 2 | 9 | 11 | 1
1 | 20220040200009 | 1 |  | 0 |  | 1 | 0 | 0 | 1 | 4 | 5 | 1 | P4 | 1 | 0.256 | 0.266 | 0.072 | 7 | 7 | 14 | 1
1 | 20220040200010 | 1 |  | 0 |  | 1 | 0 | 0 | 1 | 0 | 1 | 1 | P3 | 1 | 0.287 | 0.297 | 0.092 | 1 | 1 | 2 | 1
1 | 20220040200011 | 1 | 1 | 0 |  | 1 | 0 | 0 | 1 | 0 | 1 | 1 | P6 | 1 | 1.917 | 0.758 | 0.307 | 1 | 21 | 22 | 1
1 | 20220040200012 | 1 |  | 0 |  |  | 0 | 0 | 1 | 3 | 3 | 1 | P7 | 1 | 0.307 |  |  | 0 | 16 | 16 | 1
1 | 20220040200013 | 1 | 1 | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1
1 | 20220040200014 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1 | P8 | 1 | 0.738 | 0.656 | 0.348 | 0 | 16 | 16 | 1
1 | 20220040200015 | 1 |  | 1 |  |  | 0 | 0 | 1 | 0 | 1 | 1 | P14 | 1 | 0.482 |  |  | 1 | 21 | 22
1 | 20220040200016 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1 | P9 | 1 | 0.287 |  |  |  | 20 | 20 | 1
1 | 20220040200017 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1 | P10 | 1 | 0.43 |  |  | 1 | 14 | 15
1 | 20220040200018 | 1 |  | 0 |  |  | 0 | 0 | 1 | 2 | 2 | 1 | P11 | 1 | 0.328 |  |  | 0 | 22 | 22
1 | 20220040200019 | 1 |  | 1 |  |  | 0 | 0 | 1 | 0 | 1 | 1 | P12 | 1 | 0.605 |  |  | 2 | 9 | 11
1 | 20220040200020 | 1 | 1 | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1 | P15 | 1 | 0.307 |  |  | 22 |  | 22
1 | 20220040200021 |  | 1 | 0
1 | 20220040200022 |  | 1 | 0 | [8 empty] | study
1 | 20220040202223 | 1 |  | 1 |  |  | 0 | 0 | 1 | 0 | 1 | 1 | P13 | 1 | 0.4 |  |  | 3 | 20 | 23
1 | 20220040200024 | 1 |  | 0 |  |  | 0 | 1 | 1 | 0 | 1 | 1
1 | 20220040200025 | 1 |  | 1 |  | 1 | 1 | 0 | 1 | 0 | 3 | 1
1 | 20220040200026 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200027 |  |  | 0 |  |  | 0 | 0 |  | 0 | 0 |  | study
1 | 20220040200028 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1
1 | 20220040200029 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1
1 | 20220040200030 | 1 |  | 1 |  |  | 0 | 0 | 1 | 0 |  | 1
1 | 20220040200031 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0
1 | 20220040200032 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0 |  | 1
1 | 20220040200033 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200034 | 1 |  | 0 | 1 |  | 0 | 0 | 1
1 | 20220040200035 |  | 1 | [9 empty] | chip-rep
1 | 20220040200036 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200037 | 1 |  | 1 |  |  | -1 | 0 | 1 | 0 | 0 | 1 | P16 | 1 | 0.451 |  |  |  | 6 | 6
1 | 20220040200038 | 1 |  | 0 |  |  | 1 | 0 | 1 | 0 |  | 1
1 | 20220040200039 | 1 |  | 0 |  |  | 3 | 0 | 1
1 | 20220040200040 | 1 |  | 0 |  |  | 0 | 0 | 1 | 0 | 0 | 1
1 | 20220040200041 | 1 |  | 1 |  |  | 0 | 0 | 1
1 | 20220040200042 | 1 |  | 0 |  |  | 1 | 0 | 1
1 | 20220040200043 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200044 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200045 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200047 |  | 1 | [9 empty] | study
1 | 20220040200049 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200050 |  |  | 0 |  |  | 1 | 1 |  |  |  |  | study
| Serial # |  | Chip repl | Noisy/dead (Hybrid) | Negative Offset masked | High Offset (>50) masked | LTT |  |  | #lost after fanout | Total bad channels (Hybrid) |  | Module# |  | max I at 150C 500V before TC (mA) | max LC at 150C 500V after TC + Metr (mA) |  | Noisy/dead | Unbonded | Total bad channels | TESTS Completed
[7 empty] | cold | warm
1 | 20220040200051 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200052 | 1 |  | 0 | 1
1 | 20220040200053 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200057
1 | 20220040200061 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200062 | 1 |  | 0 |  |  | 0 | 0
1 | 20220040200063 | 1 |  | 0 |  | 2 | 0 | 0 | 1
1 | 20220040200064 | 1 |  | 0 |  |  | 1 | 0
1 | 20220040200065 | 1 |  | 0 |  |  | 0 | 0 | 1
1 | 20220040200066 | 1 |  | 0 |  | 1
1 | 20220040200067 | 1 |  | 1
1 | 20220040200068 | 1 |  | 1
1 | 20220040200069 | 1 |  | 0
1 | 20220040200070 | 1 |  | 0
1 | 20220040200071 | 1 |  | 0
1 | 20220040200072 | 1 |  | 0
1 | 20220040200073
1 | 20220040200074 | 1
20220040200075
20220040200076
20220040200077
Hybrids Built |  | Hybrids OK | [5 empty] | Hybrids up to fanout |  |  | Hybrids to module |  |  | Modules | [5 empty] | Module Test Completed
59 |  | 51 | [6 empty] | 40 |  |  | 23 |  | 14 | [6 empty] | 7

## 008-030
20220040200008 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Intitial P2 | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 | 0 | 9 | 11 | 11
Channel No.(Defect)
Comments | FullByassTest failed during Characterization. Manually retested, passed. Chip 10 has unusually high gain. | Performed after fanout due to lack of insulating box when first tested | NO reading of 8.75 E-007 |  | channels: unbonded 60,472,564,1168,1268, 1270,1272 high noise 530, 1074 dead 952,953 1074 has low gain, high noise, high offset, low vt50 | unbonded: 60,472,564,1168,1246,1250,1268,1270,1272 (2 more channel unbonded compared to module initial) Leakage 0.922 microA at 500V | unbonded: 60,472,564,1168,1246,1250,1268,1270,1272 (2 more channel unbonded compared to module initial) dead: 952,953 partbonded: 1083 high noise: 46
2022040200009 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Intitial P4 | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 5 | 0 | 5 | 13 | 14 | 14
Channel No.(Defect) |  | 981(Noisy) 982(Noisy) 993(Noisy) 994(Noisy) 1402(high offset) |  | 981(Noisy) 982(Noisy) 993(Noisy) 994(Noisy) 1402(high offset)
Comments | 80 channels masked on range -1 (noisy) but only 2 defects actually found. Scurve chip 10 (ch#1402) shows higher vt50 Chip 9-10 higher gain Chip11 lower gain | Performed after fanout due to lack of insulating box when first tested. Noisy channels appear due to possible shorting. NO reading of 5 E-004 until channel manually masked. Scurves link 1 show higher vt50 for chip 10 (and irregular shapes). Charact. Retest with ch#1402 manually masked - scurves are fine | NO reading of 2.44 E-007 | These noisy channels appeared after the fan-out bonding. Possibly due to a short | unbonded 130, 136, 348, 1016, 1170, 1174 high noise 688 dead 981,982, 993,994, defect on silicon(dead): 1197, 1198, 1199 | unbonded 130, 136, 348, 1016, 1170, 1174 + 1282 dead 981,982, 993,994, defect on silicon(dead): 1197, 1198, 1199 | unbonded 130, 136, 348, 1016, 1170, 1174 + 1282 dead 981,982, 993,994, defect on silicon(dead): 1197, 1198, 1199 chip0 and chip 7 several channels with high noise
20220040200010 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Intitial P3 | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 1 | 1 | 1 | 2 | 2 | 2
Channel No.(Defect) | 542 (high offset) | 542 (masked) | 542 (masked) | 542 (masked) | 542 (masked) |  | 542 (masked)
Comments | retested after masked manually due to high offset of this channel during RC test. Channel 687 has 15% higher gain than average, but offset is unaffected. Charact. Retest with ch 542 manualy masked. | Cold LTT was manually stopped after 15 h (we realized that cold-warm button are swapped in software). LTT test (redone) and Charact. test done together with hybrids 12-16. NO reading of 5E-004 until channel manually masked. | channel 687 has 15% higher gain than average. |  | unbonded 1404 |  | unbonded 1404
20220040200011 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P6 | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 1 | 1 | 1 | 22 | 22
Channel No.(Defect) | 1369 | 1369 | 1369 | 1369 | 1369 high offset 21 channels unbonded: 15, 83, 628, 298, 769, 789, 791, 795, 799, 801, 811, 816, 900, 924, 932, 949, 969, 1023, 1381, 1444, 1491 | same
Comments | Chip 11 large gain spread - replaced on 8/02/02. Chips 1 and 8 have unusually high gain, channel 1369 has high offset (Chip10) but it wasn't masked | final Characterization manually retested due to software crash before NO. NO reading of 5.76E-005 (because of the shifted level due to the high offset channel | NO reading of 6.47E-004
20220040200012 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P7 | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 | 0 | 16
Channel No.(Defect) |  |  |  |  | unbonded: 16, 62, 68, 96, 614, 790, 906, 940, 948, 1066, 1175, 1189, 1352, 1376, 1438, 1529
Comments |  |  | done in two parts due to corrupt default/system | 3 inner-row fanout bonds popped off, will show up in module tests
20220040200013 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments |  | chip 4 has lower gain towards the last channels and higher noise for the same channels 5 defects (617 620 628-633 635 637-638) one more test at cold where 13 is switched with 62 to study possible failure of any hybrid in module 2 shows 13 failing but the reason is being in module 2 | done in two parts due to corrupt default/system. Chips 2-5 fail all 3 digital tests at around 33 deg, sent to Santa Cruz. And then back… we'll probably replace chip 2
20220040200014 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P8 | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 | 0 | 15 | 16 | 16
Channel No.(Defect) |  |  |  |  | unbonded: 83, 88, 116, 168, 179, 247, 257, 427, 583, 617, 634, 662, 680, 886, 1317 | same 15 unbonded + ch 52 partbonded | unbonded: 52, 83, 88, 116, 168, 179, 247, 257, 427, 583, 617, 634, 662, 680, 886, 1317
Comments | Chip 2 and Chip 8 to 10 show higher gain and lower offset,and gain is non-uniform from chip to chip. 1 defect: Channel 899 listed as partbonded has low gain and higher noise. | non-uniform gain persists | done in two parts due to corrupt default/system, non-uniform gain persists. Up to 11 partbonded (higher noise) channels. |  |  |  | Chip 0 and Chip 1 have several channels with high noise (2nd half encreasingly)
20220040200015 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P14 | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 1 | 1 | 1 | 22
Channel No.(Defect) | 1451(high gain high vt50) | 1451 (masked) | 1451 (masked) | 1451 | 1451 (high gain) 18 ch unbonded: 8, 41, 49, 93, 134, 151, 339, 367, 466, 688, 733, 778, 780, 845, 873, 1382, 1425, 1482 3 ch partbonded: 703, 1206, 1430
Comments | …00015.det file with modified FEShp. created due to non-uniform gain. Gain gradually decreases by 25% from chip0 to chip2. Same on stream 1 and chp 11 has higher gain and close to 30% difference in gain between chip 10 and 11. Characterization test was redone with new settings. and at this point ch 1451 got masked (high gain and high vt50). ch 1451 has high gain and the entire chip 11 has very high gain. | chip 8 had 17 channels with partbonded defect due to noise exceeding 750 | done in two parts due to corrupt default/system. NO reading of 8.14E-008 | Retest with default.det. Channel 1451 manually masked. Data is OK.
20220040200016 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P9 | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 | 0 | 20
Channel No.(Defect) |  |  |  |  | unbonded: 16, 102, 168, 185, 377, 408, 526, 568, 615, 723, 777, 932, 956, 989, 1029, 1087, 1145, 1311, 1347, 1354
Comments | Irregular Gain link 0 | Oscillating and higher gain on chip 2,6,8,10,11. NO plots show a blank spot. It's probably due to a memory problem (demonstrated with warm retest) | done in two parts due to corrupt default/system; has minor detector noise readings. Warm Chrac.test redone individually and NO plots show good results (previously showed blank spots as in cold test)
20220040200017 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P10 | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 |  | 15
Channel No.(Defect) |  |  |  |  | 14 Unbonded: 237,264,266,565,728,745,865,867,1051,1060,1121, 1487,1501,1503 1 high noise: 1473
Comments | 10 Channels part-bonded. | Several channels on chip 5 have high gain, but only after trim range | 8 channels part-bonded | 6 channels partbonded.
20220040200018 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P11 | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 |  | 22
Channel No.(Defect) |  |  |  |  | unbonded:6 8 162 325 406 434 581 727 729 737 739 740 741 743 749 753 1238 1248 1250 1252 1265 1432
Comments |  | Several channels on chip 4 have higher gain, but only after trim
20220040200019 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P12 | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 1 | 1 |  | 11
Channel No.(Defect) | 910 (chip 7) | 910(dead or stuck) | 910(dead or stuck) |  | 712 713 dead (low gain) 9 unbonded: 43 536 582 703 705 742 744 746 1129
Comments | channel 910 dead after trim | chip 4 has spikes in the gain and a irregular offset after the trim | 3 part bonded channels. 910 is masked.
20220040200020 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | chip 1 | 0 | 0 | 0 | 22
Channel No.(Defect) |  |  |  |  | 29 49 162 497 507 514 -515 544 720 723 726 729 733-735 743 788 1006 1308 1400 1512 1517 1528
Comments | Chip 1 TOKEN failure - replaced.TOKEN defect on chip 1 prevents the rest of the hybrid from functioning. When chip1 is bypassed the rest of the hybrid works. Chip 1 is replaced |  |  |  | chip 7 High gain
20220040200021 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | chip 8
Channel No.(Defect)
Comments | Overall chip 8 shows high and irregular gain and offset. Before trim also vt50 is irregular. 9 channels partbonded. Chip 8 large gain spread - it will be replaced (UCSC)
20220040200022 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | chip 10
Channel No.(Defect)
Comments | Before trim chip 10 shows high offset that moves up NO (after trim). If chip10 is bypassed all chips are overall normal. Probably need to be replaced
20220040200023 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial P13 | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 1 | 1 | 0 | 23
Channel No.(Defect) | 1137 (chip 8) | 1137 | 1137 |  | 1136, 1137 1138 (noisy) unbonded: 413 423 435 439 451 492 607 736 950 1012 1023 1051 1117 1236 1334 1512 partbonded: 9 994 1009 1476
Comments | channel 1137 has high noise (825) | Chip 3, 8, 9 spiky gain after trim range. Retest at room temp OK
20220040200024 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 1 | 0
Channel No.(Defect) |  |  | 708(untrimmable) | 0
Comments | 1 channel partbonded (1373) | chip 7, 8, 9, 11 high gain for several channels after trim. Retest at room temp OK | TrimRangeScan fails. Chip 10 shows slightly higher noise/low gain. | 1360 partbonded
20220040200025 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 2 | 2 | 2
Channel No.(Defect) | 699 (chip 5) | 0 (noisy) 699(high gain) | 591(dead), 699(dead)
Comments | Channel 699 shows higher gain that would deform NO results. Masked and retest. NO shape good, but high, after retest ch 591 has very high offset. Need retest before module because of ch 591 high offset | chip 2, 3,6,8, and 11 high/irreg gain. retest at room temp OK | Apparently one channel was gained as another was lost. | Retest with channel 591 manually masked. Channel 0 gets masked by the TrimRange.
20220040200026 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments
20220040200027 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments |  | failed full bypass test, but no defects. Need retest? | channel 445 is partbonded. Final: 447 and 454 part bonded. Fullbypass test is ok here
20220040200028 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 | 0
Channel No.(Defect)
Comments | A couple of channels show peaking high gain which remains also after trim chip 0 channel 97 chip 4 channel 629 60 channels before trim partbonded after trim only chip 10 (15 channels) partbonded |  |  | ch 97 and 629 higher gain
20220040200029 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 | 0
Channel No.(Defect)
Comments | chi 1 and 4 higher gain chip 10 high noise | chips 1, 4, 7, 8, 11 have spikes in the gain after trim. No room temperature retest | 32 channels part bonded on chip 10. | chip 10 higher noise (low gain)
20220040200030 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 1 | 1 | 0
Channel No.(Defect) | 1207(high noise, manually masked) | 1207 | 1207
Comments | Didn't retest, will simply mask 1207 during burn-in. | channel 360 has high offset. After trim, all other channels in stream 0 are raised to that level. Chip 2, 6, 7,8,9, and 11 have spikes in the gain after trim. Errors that occur in the warm test are as follows: First, channel 1207 is manually masked due to high noise. The remaining errors are mostly high noise channels labeled partbonded on chip 04. There are a small number of partbonded errors on chip 05. Neither increases with time.

## 031-050
20220040200031 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments |  | chip 10 and 11 have spikes in the gain after trim. | 3 channels part-bonded in three-point gain. 1 part bonded after trim.
20220040200032 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 | 0
Channel No.(Defect)
Comments | dip in vt50 around chanel 520. Gone after trim | spikes on the gain on chip 8 only after trim. |  | ch 520 (about) low vt50
20220040200033 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments | dip in vt50 around channel 170. Gone after trim.` | dip in vt50 around channel 170. Gone after trim. Chips 0,3,4,7,10 have spikes in the gain, only after trim.
20220040200034 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 1 | 1
Channel No.(Defect) | 400 | 400 | 400
Comments | manually masked after trim range scan - low vt50 | channel 400 has low vt50 before trim and low gain and high offset after trim. | channel 400 has low vt50 before trim and low gain and high offset after trim.
20220040200035 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 2
Channel No.(Defect) | 767(untrimmable), 994(high offset:manually masked)
Comments | Chip 6 Time Walk failure and High gain - to be replaced
20220040200036 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments | good, save a small spike in the vt50 and gain around channel 1400. The largest spike occurs at precisely channel 1423. | Channels 756, 757, 758, 759, 760, and 762 show partbonded defects on the three pont gain test. High gain (and spread) in chip #7 . Channels 760, 766, 768 show high noise in the response curve test. | Channel 1423 shows same spike as noted earlier.
20220040200037 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 0 | 0 | 0 | 5
Channel No.(Defect) | 935 (untrimmable) | channel 935 is good now |  |  | unbonded 311 1059 1105 1188 1280 (partbonded) 1313
Comments | Trim range scan FAILS on hybrid initial test. Channels 355, 357, and 1189 show partbonded defects in the response curve. Noise occupance shows non-100% for low threshold values. | chan 935 untrimmable on hybrid initial is ok after LTT test. Channel 96 shows partbonded defect in 3 point gain. High gain in chip 10 and large spread (also in chip 8 and 9) (8fC effect) on the response curve plot. | Channel 198 shows partbonded error on initial 3 point gain test. Channel 231 shows partbonded defect on final 3-point gain plot. | ch 139 227 247 855 879 1203 with higher noise
20220040200038 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 1 | 0 | 0
Channel No.(Defect) |  | 1015 (dead)
Comments | Channels 383, 838, 890, 1198, and 1243 show partbonded defects on 3-point gain. | Channels 1239 and 1015 show partbonded defects on initial 3-point gain. Channels 1255 and 1015 show partbonded defects on final 3-point gain test. Gain spikes/spread on chip 0, 5 and chip 11 is also high. | cold LTT masks one channel as noisy/dead but warm LTT not. Initial 3-point gain shows chip #9 as noisy with 18 partbonded defects. Final 3-point gain show similar results. | ch 1197 1234 high noise
20220040200039 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 3 | 0
Channel No.(Defect) |  | 179, 1360 untrimmable 787 negative offset after trim
Comments | Large spread in gain on chip 2. | Chip 2 and Chip 6 very large gain spread and channel 787 shows negative offset even after trim. NO bump due to chip 2 | Chip 6 variation on gain chip2 gain spread lost 3 channels on LTT cold, OK at warm | one chip found chipped (chip 10) but chip 2 is actually showing irregular gain and noise
20220040200040 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0 | 0
Channel No.(Defect)
Comments | ch 475 (partbonded)
20220040200041 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1 | 1 | 1
Channel No.(Defect) | 1443 (noisy) | 1443 (noisy) | 1443 (noisy)
Comments | 234 (low vt50 neg offset) after trim chip 2 large spread gain. Channel 309 gain=0 and noise=0 | after trim chip2 very large gain spread | no gain spread
20220040200042 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 1 | 0
Channel No.(Defect) |  | 985 (untrimmable)
Comments |  | from the plots it looks ok! | one chan. untrimmable at cold but ok at warm
20220040200043 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments | channel 636 has very low vt50 and negative offset but NO is ok | same | same
20220040200044 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments
20220040200045 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments
20220040200047 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | needs study
Channel No.(Defect)
Comments | chip 0 possiby bad
20220040200049 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments |  | gain spread on chip 0
20220040200050 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 1 | 1
Channel No.(Defect) |  | 1301 (Noisy) | 1301
Comments |  | Also Redundancy test fails | Redundancy test fails same behaviour as in cold -->need study but final characterization ok

## 051-070
20220040200051 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments
20220040200052 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1
Channel No.(Defect) | 481 (manually masked)
Comments | ch 481 has very low vt50 and negative offset. chip 7 has large gain spread. Retested and without 8fC point. OK
20220040200053 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments
20220040200057 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 2
Channel No.(Defect) | 149 796 low vt50 and negative offset ch 510 untrimmable.
Comments | needs study besides ch 149 and 796 a few other channels with high/low vt50 high offset/negative offset, NO strange Retested with ch 149 and 796 masked. Sort of ok.
20220040200061 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments |  | scurves first 6 chips especially chip 0 show efficiency > 100 % at low threshold | ch 957 negative offset 1304 and 1352 partbonded
20220040200062 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 0 | 0
Channel No.(Defect)
Comments | chip 9 has higher gain and every channel has negative offset. After trim OK | chip 9 has higher gain and every channel has negative offset. After trim OK Redundancy test fails on chip 1 (com1). Stream 0 vt50 for several channels is 0. Need study | chip 9 has higher gain and every channel has negative offset. After trim OK same as in cold. Need study. Later on switched to fixture 0 and confirmation at cold OK no failure
20220040200063 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 2 | 2 | 2
Channel No.(Defect) | 1004 (offset 54.8) 1229 (offset 82.7) | same high offset channels. | same high offset channels.
Comments | Spikes in the vt50 and high offset at channels 1004(offset 54.8) 1229 ( 82.7) | same high offset channels. NO high. Need to mask channels if NO high at module | same high offset channels. NO high. Need to mask channels if NO high at module
20220040200064 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0 | 1 | 0
Channel No.(Defect) |  | 100 (high noise)
Comments | Channel 416 has high gain | 416 has high gain chan 100 is mask at trim and fails Redundancy Test. Need study . Characterization retest later date -> no failure | channel 416 high gain ch 100 is ok
20220040200065 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 2 | 0 | 0
Channel No.(Defect) | 149 (dead) 796 (dead)
Comments | channels 149 and 796 are dead. It looks like chip 1 and 6 were not good to start with (1 bad ch chips) later on comparison with wafer data shows that those chips were perfect. It turned out to be a wrong mask. OK | chan 149,796 that appeared as dead in hybrid initial are ok here. OK
20220040200066 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1
Channel No.(Defect) | 309 (high offset)
Comments | channel 309 has high offset needs to be manually masked. High NO Retest with Channel 309 masked. Looks good.
20220040200067 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1
Channel No.(Defect) | 371 (untrimmable)
Comments
20220040200068 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 1
Channel No.(Defect) | 203 (untrimmable)
Comments | chip 1 has large gain spread also after trim and one chan (203) untrimmable
20220040200069 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0
Channel No.(Defect)
Comments
20220040200070 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0
Channel No.(Defect)
Comments | Chip 11 has lower noise (and higher gain) than the rest. 29 channels are partbonded. A majority of those high noise channels are on chip 10.

## 071-090
20220040200071 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0
Channel No.(Defect)
Comments
20220040200072 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0
Channel No.(Defect)
Comments | NO plot need attention
20220040200073 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | chip 0
Channel No.(Defect)
Comments | chip 0 fails Strobe Delay
20220040200074 | Hybrid_Initial | Hybrid LTT Cold | Hybrid LTT Warm | Post Fan-Out | Module_Initial | Post Therm. Cyc. | Module LTT
No. bad channels | 0
Channel No.(Defect)
Comments