### Library Creation Process SOP:

### Qpix Colony Picker Protocol

Version Number:		1.0

Start Production Date:		x

Author:			Amy Geotina, Doug Smith

Reviewed by:			Nancy Hammon

SOP Change Document:	x

# Summary

This SOP describes the use of the Qpix colony picker as it applies to the Library Creation group at the JGI.

# Materials & Reagents

Materials/Reagents/Equipment			Vendor		Stock Number

##### Disposables

##### LB Amp 100 X-gal Plates (254mm)			Teknova	L133-A100X

##### Nunc 384 Well Plate, Cell Culture			Applied	Scientific

##### Stock Solutions

LB Broth with 7.5% Glycerol				Teknova	0181-G7.5

##### Reagents

##### Equipment

Qpix

# Procedure

Set up:

- To power on the Qpix turn the red lever on the right side of the picker from a horizontal to a vertical position.  Press the red button marked reset on the front of the Qpix then the button on the lower right of the computer to power on the computer.

- Check that the air compressor on the left side of the Qpix is turned on and the release valve is shut.  To turn on the air compressor twist the red button on the top of the pump counter clockwise.

- Launch the software program from the desktop icon Q-pix.

- To install the picking head click diagnostics from the toolbar at the bottom of the picking software’s window.  Move the actuator into an accessible position above by right clicking the mouse on a position above the picking bed and selecting move actuator to from the drop down menu that appears.  Once the head is in the desired position open the door and insert the picking head into the actuator from the right side.  Be sure and inspect the head first for any bent pins.  Secure the head in the actuator by inserting the thumb screw into the hole on the left side of the actuator and tightening.

- Prepare the ethanol bath for picking by inserting the scrub brushes into the rectangular metal well and place the well into the front most well position on the right side of the picking bed.  Fill the well with 70% ethanol so the level of the liquid just covers the top of the brushes.  Check the ethanol level during use and add 100% ethanol if  the liquid level drops below the brush tips.

Calibration:

- The Qpix does not require daily calibration as long as the plates being imaged and picked are checked for accurate picking and proper screening of the blue colonies.  In general calibration once a week is sufficient.

- Refer to the Qpix manual located on the shelves behind the machine for calibration instructions.

- In general the size and shape parameters rarely need to be manipulated but the threshold values can change from plate to plate but tend to be consistent through an entire library.

- Tips for adjusting the threshold setting not mentioned in the manual:  Make note of the different threshold values that are tried.  There really seems to be no rule as how to best set the threshold values.  The best settings are usually found by trying a few different values and selecting the value that gives you the highest number of picks.  When comparing threshold values, be sure that the number of picks being compared is for an entire bioassay plate since the Qpix images different regions of the plate differently.

Run:

- Check agar volume is set to 300 ml.  This can be done from the main menu.

- Set the max. plates to use to the desired number of output plates.  The maximum that can be set up in one run is 15 destination plates.  It is ok to set the plate number higher than the amount of colonies are available because the picker will indicate what plates it inoculated.

- Load the destination plate starting with the front right position and placing each additional plate in the column before moving onto next.

- Load plate with well A1 oriented towards the upper right corner.  The plates should be snug in to corner.  Negligence in this step will lead to picking head crashes and damage the head and needles.

- Check that the plexy glass is clean on the picking bed.

- Check that the bioassay plate is placed snuggly onto the bed.  If the plate is loose it will likely shift during the run and become inaccurate.

- Once source and destination plates are properly placed in the picking bed select full from the bottom toolbar to image the plate.  Partial may be selected to selectively image certain regions of the plate.

- Watch the computer monitor while the qpix is imaging to check each area of the bioassay is being processed correctly.

- If the imaging looked ok click no on the display image data prompt.  Otherwise view the image again.  If adjusting the imaging parameters refer to the manual located behind the qpix.

- After a plate has been picked a dialogue box will alert you to the completion.  From here, another box comes up asking whether to keep plate and position settings.  Only if this was the last plate picked for a library should No be selected.  Selecting yes will allow the picking for the next plate to be imaged to begin where picking just left off or for you to see what the last plate inoculated was and how many wells were inoculated.

Shutting down:

- Remove the sterilization bath and pour the remaining ethanol into the waste container.  Remove the brush and rinse it with ethanol and remove any debris.

- Enter the diagnostic screen and position the actuator towards the center of the picking bed.  Remove the picking head by loosening the screw to the left of the head and sliding the head out of the actuator while taking care to ensure that is does not drop.  DO NOT place your hand under the needles.  Place the head on the bed with the pins facing upward.

- Close the front door and click on UV light from the bottom of the toolbar.  Allow the light to be on for 10 minutes.

- When the UV light is done select done in the bottom right of the Qpix software’s window to exit the program.  Once the program has been exited shut down the computer from the start button.  After the computer has been given adequate time to shut down turn the red switch on the right hand side of the Qpix.  Turn off power to the air compressor by pressing the red button on top then open the bleed valve to release any pressure.

# Reagent/Stock Preparation

How to prepare reagents and stocks