An Assessment of the High-gain Streckeisen STS2 Seismometer for Routine Earthquake Monitoring in the United States

By D. E. McNamara, R.P. Buland and H.M. Benz

Report Series XXXXXXX

U.S. Department of the Interior

U.S. Geological Survey

U.S. Department of the Interior

Gale A. Norton, Secretary

U.S. Geological Survey

P. Patrick Leahy, Acting Director

U.S. Geological Survey, Reston, Virginia 200x
Revised and reprinted: 200x

For product and ordering information:
World Wide Web: http://www.usgs.gov/pubprod
Telephone: 1-888-ASK-USGS

For more information on the USGS—the Federal source for science about the Earth,
its natural and living resources, natural hazards, and the environment:
World Wide Web:  http://www.usgs.gov
Telephone:  1-888-ASK-USGS

Suggested citation:
Author1, F.N., Author2, Firstname, 2001, Title of the publication: Place of publication 
(unless it is a corporate entity), Publisher, number or volume, page numbers; information 
on how to obtain if it’s not from the group above.

Any use of trade, product, or firm names is for descriptive purposes only and does not imply 
endorsement by the U.S. Government.

Although this report is in the public domain, permission must be secured from the individual 
copyright owners to reproduce any copyrighted material contained within this report.

Contents

First-Order Heading (Heading 1)	Error! Bookmark not defined.

Second-Order Heading (Heading 2)	Error! Bookmark not defined.

Third-Order Heading (Heading 3)	Error! Bookmark not defined.

Fourth-Order Heading (Heading 4)	Error! Bookmark not defined.

Fifth-Order Heading (Heading 5)	Error! Bookmark not defined.

References Cited	Error! Bookmark not defined.

Glossary	Error! Bookmark not defined.

Figures

1. Figure caption with (A) multiple parts (B) that are divided by the (C) MultipartFigCap style.	Error! Bookmark not defined.

Tables

1. This is a table title.	Error! Bookmark not defined.

Conversion Factors

Inch/Pound to SI

| Multiply | By | To obtain |
| --- | --- | --- |
| Length |
| inch (in.) | 2.54 | centimeter (cm) |
| inch (in.) | 25.4 | millimeter (mm) |
| foot (ft) | 0.3048 | meter (m) |
| mile (mi) | 1.609 | kilometer (km) |
| mile, nautical (nmi) | 1.852 | kilometer (km) |
| yard (yd) | 0.9144 | meter (m) |
| Area |
| acre | 4,047 | square meter (m2) |
| acre | 0.4047 | hectare (ha) |
| acre | 0.4047 | square hectometer (hm2) |
| acre | 0.004047 | square kilometer (km2) |
| square foot (ft2) | 929.0 | square centimeter (cm2) |
| square foot (ft2) | 0.09290 | square meter (m2) |
| square inch (in2) | 6.452 | square centimeter (cm2) |
| section (640 acres or 1 square mile) | 259.0 | square hectometer (hm2) |
| square mile (mi2) | 259.0 | hectare (ha) |
| square mile (mi2) | 2.590 | square kilometer (km2) |
| Volume |
| barrel (bbl), (petroleum, 1 barrel=42 gal) | 0.1590 | cubic meter (m3) |
| ounce, fluid (fl. oz) | 0.02957 | liter (L) |
| pint (pt) | 0.4732 | liter (L) |
| quart (qt) | 0.9464 | liter (L) |
| gallon (gal) | 3.785 | liter (L) |
| gallon (gal) | 0.003785 | cubic meter (m3) |
| gallon (gal) | 3.785 | cubic decimeter (dm3) |
| million gallons (Mgal) | 3,785 | cubic meter (m3) |
| cubic inch (in3) | 16.39 | cubic centimeter (cm3) |
| cubic inch (in3) | 0.01639 | cubic decimeter (dm3) |
| cubic inch (in3) | 0.01639 | liter (L) |
| cubic foot (ft3) | 28.32 | cubic decimeter (dm3) |
| cubic foot (ft3) | 0.02832 | cubic meter (m3) |
| cubic yard (yd3) | 0.7646 | cubic meter (m3) |
| cubic mile (mi3) | 4.168 | cubic kilometer (km3) |
| acre-foot (acre-ft) | 1,233 | cubic meter (m3) |
| acre-foot (acre-ft) | 0.001233 | cubic hectometer (hm3) |
| Flow rate |
| acre-foot per day (acre-ft/d) | 0.01427 | cubic meter per second (m3/s) |
| acre-foot per year (acre-ft/yr) | 1,233 | cubic meter per year (m3/yr) |
| acre-foot per year (acre-ft/yr) | 0.001233 | cubic hectometer per year (hm3/yr) |
| foot per second (ft/s) | 0.3048 | meter per second (m/s) |
| foot per minute (ft/min) | 0.3048 | meter per minute (m/min) |
| foot per hour (ft/hr) | 0.3048 | meter per hour (m/hr) |
| foot per day (ft/d) | 0.3048 | meter per day (m/d) |
| foot per year (ft/yr) | 0.3048 | meter per year (m/yr) |
| cubic foot per second (ft3/s) | 0.02832 | cubic meter per second (m3/s) |
| cubic foot per second per square mile [(ft3/s)/mi2] | 0.01093 | cubic meter per second per square kilometer [(m3/s)/km2] |
| cubic foot per day (ft3/d) | 0.02832 | cubic meter per day (m3/d) |
| gallon per minute (gal/min) | 0.06309 | liter per second (L/s) |
| gallon per day (gal/d) | 0.003785 | cubic meter per day (m3/d) |
| gallon per day per square mile [(gal/d)/mi2] | 0.001461 | cubic meter per day per square kilometer [(m3/d)/km2] |
| million gallons per day (Mgal/d) | 0.04381 | cubic meter per second (m3/s) |
| million gallons per day per square mile [(Mgal/d)/mi2] | 1,461 | cubic meter per day per square kilometer [(m3/d)/km2] |
| inch per hour (in/h) | 0 .0254 | meter per hour (m/h) |
| inch per year (in/yr) | 25.4 | millimeter per year (mm/yr) |
| mile per hour (mi/h) | 1.609 | kilometer per hour (km/h) |
| Mass |
| ounce, avoirdupois (oz) | 28.35 | gram (g) |
| pound, avoirdupois (lb) | 0.4536 | kilogram (kg) |
| ton, short (2,000 lb) | 0.9072 | megagram (Mg) |
| ton, long (2,240 lb) | 1.016 | megagram (Mg) |
| ton per day (ton/d) | 0.9072 | metric ton per day |
| ton per day (ton/d) | 0.9072 | megagram per day (Mg/d) |
| ton per day per square mile [(ton/d)/mi2] | 0.3503 | megagram per day per square kilometer [(Mg/d)/km2] |
| ton per year (ton/yr) | 0.9072 | megagram per year (Mg/yr) |
| ton per year (ton/yr) | 0.9072 | metric ton per year |
| Pressure |
| atmosphere, standard (atm) | 101.3 | kilopascal (kPa) |
| bar | 100 | kilopascal (kPa) |
| inch of mercury at 60ºF (in Hg) | 3.377 | kilopascal (kPa) |
| pound-force per square inch (lbf/in2) | 6.895 | kilopascal (kPa) |
| pound per square foot (lb/ft2) | 0.04788 | kilopascal (kPa) |
| pound per square inch (lb/in2) | 6.895 | kilopascal (kPa) |
| Density |
| pound per cubic foot (lb/ft3) | 16.02 | kilogram per cubic meter (kg/m3) |
| pound per cubic foot (lb/ft3) | 0.01602 | gram per cubic centimeter (g/cm3) |
| Energy |
| kilowatthour (kWh) | 3,600,000 | joule (J) |
| Radioactivity |
| picocurie per liter (pCi/L) | 0.037 | becquerel per liter (Bq/L) |
| Specific capacity |
| gallon per minute per foot [(gal/min)/ft)] | 0.2070 | liter per second per meter [(L/s)/m] |
| Hydraulic conductivity |
| foot per day (ft/d) | 0.3048 | meter per day (m/d) |
| Hydraulic gradient |
| foot per mile (ft/mi) | 0.1894 | meter per kilometer (m/km) |
| Transmissivity* |
| foot squared per day (ft2/d) | 0.09290 | meter squared per day (m2/d) |
| Application rate |
| pounds per acre per year [(lb/acre)/yr] | 1.121 | kilograms per hectare per year [(kg/ha)/yr] |
| Leakance |
| foot per day per foot [(ft/d)/ft] | 1 | meter per day per meter |
| inch per year per foot [(in/yr)/ft] | 83.33 | millimeter per year per meter [(mm/yr)/m] |

Temperature in degrees Celsius (°C) may be converted to degrees Fahrenheit (°F) as follows:

°F=(1.8×°C)+32

Temperature in degrees Fahrenheit (°F) may be converted to degrees  Celsius (°C) as follows:

°C=(°F-32)/1.8

Vertical coordinate information is referenced to the insert datum name (and abbreviation) here for instance, “North American Vertical Datum of 1988 (NAVD 88).”

Horizontal coordinate information is referenced to the insert datum name (and abbreviation) here for instance, “North American Datum of 1983 (NAD 83).”

Altitude, as used in this report, refers to distance above the vertical datum.

*Transmissivity: The standard unit for transmissivity is cubic foot per day per square foot times foot of aquifer thickness [(ft3/d)/ft2]ft. In this report, the mathematically reduced form, foot squared per day (ft2/d), is used for convenience.

Specific conductance is given in microsiemens per centimeter at 25 degrees Celsius (µS/cm at 25 °C).

Concentrations of chemical constituents in water are given either in milligrams per liter (mg/L) or micrograms per liter (µg/L).

NOTE TO USGS USERS: Use of hectare (ha) as an alternative name for square hectometer (hm2) is restricted to the measurement of small land or water areas. Use of liter (L) as a special name for cubic decimeter (dm3) is restricted to the measurement of liquids and gases. No prefix other than milli should be used with liter. Metric ton (t) as a name for megagram (Mg) should be restricted to commercial usage, and no prefixes should be used with it.

SI to Inch/Pound

| Multiply | By | To obtain |
| --- | --- | --- |
| Length |
| centimeter (cm) | 0.3937 | inch (in.) |
| millimeter (mm) | 0.03937 | inch (in.) |
| meter (m) | 3.281 | foot (ft) |
| kilometer (km) | 0.6214 | mile (mi) |
| kilometer (km) | 0.5400 | mile, nautical (nmi) |
| meter (m) | 1.094 | yard (yd) |
| Area |
| square meter (m2) | 0.0002471 | acre |
| hectare (ha) | 2.471 | acre |
| square hectometer (hm2) | 2.471 | acre |
| square kilometer (km2) | 247.1 | acre |
| square centimeter (cm2) | 0.001076 | square foot (ft2) |
| square meter (m2) | 10.76 | square foot (ft2) |
| square centimeter (cm2) | 0.1550 | square inch (ft2) |
| square hectometer (hm2) | 0.003861 | section (640 acres or 1 square mile) |
| hectare (ha) | 0.003861 | square mile (mi2) |
| square kilometer (km2) | 0.3861 | square mile (mi2) |
| Volume |
| cubic meter (m3) | 6.290 | barrel (petroleum, 1 barrel = 42 gal) |
| liter (L) | 33.82 | ounce, fluid (fl. oz) |
| liter (L) | 2.113 | pint (pt) |
| liter (L) | 1.057 | quart (qt) |
| liter (L) | 0.2642 | gallon (gal) |
| cubic meter (m3) | 264.2 | gallon (gal) |
| cubic decimeter (dm3) | 0.2642 | gallon (gal) |
| cubic meter (m3) | 0.0002642 | million gallons (Mgal) |
| cubic centimeter (cm3) | 0.06102 | cubic inch (in3) |
| cubic decimeter (dm3) | 61.02 | cubic inch (in3) |
| liter (L) | 61.02 | cubic inch (in3) |
| cubic decimeter (dm3) | 0.03531 | cubic foot (ft3) |
| cubic meter (m3) | 35.31 | cubic foot (ft3) |
| cubic meter (m3) | 1.308 | cubic yard (yd3) |
| cubic kilometer (km3) | 0.2399 | cubic mile (mi3) |
| cubic meter (m3) | 0.0008107 | acre-foot (acre-ft) |
| cubic hectometer (hm3) | 810.7 | acre-foot (acre-ft) |
| Flow rate |
| cubic meter per second (m3/s) | 70.07 | acre-foot per day (acre-ft/d) |
| cubic meter per year (m3/yr) | 0.000811 | acre-foot per year (acre-ft/yr) |
| cubic hectometer per year (hm3/yr) | 811.03 | acre-foot per year (acre-ft/yr) |
| meter per second (m/s) | 3.281 | foot per second (ft/s) |
| meter per minute (m/min) | 3.281 | foot per minute (ft/min) |
| meter per hour (m/hr) | 3.281 | foot per hour (ft/hr) |
| meter per day (m/d) | 3.281 | foot per day (ft/d) |
| meter per year (m/yr) | 3.281 | foot per year ft/yr) |
| cubic meter per second (m3/s) | 35.31 | cubic foot per second (ft3/s) |
| cubic meter per second per square kilometer [(m3/s)/km2] | 91.49 | cubic foot per second per square mile [(ft3/s)/mi2] |
| cubic meter per day (m3/d) | 35.31 | cubic foot per day (ft3/d) |
| liter per second (L/s) | 15.85 | gallon per minute (gal/min) |
| cubic meter per day (m3/d) | 264.2 | gallon per day (gal/d) |
| cubic meter per day per square kilometer [(m3/d)/km2] | 684.28 | gallon per day per square mile [(gal/d)/mi2] |
| cubic meter per second (m3/s) | 22.83 | million gallons per day (Mgal/d) |
| cubic meter per day per square kilometer [(m3/d)/km2] | 0.0006844 | million gallons per day per square mile [(Mgal/d)/mi2] |
| cubic meter per hour (m3/h) | 39.37 | inch per hour (in/h) |
| millimeter per year (mm/yr) | 0.03937 | inch per year (in/yr) |
| kilometer per hour (km/h) | 0.6214 | mile per hour (mi/h) |
| Mass |
| gram (g) | 0.03527 | ounce, avoirdupois (oz) |
| kilogram (kg) | 2.205 | pound avoirdupois (lb) |
| megagram (Mg) | 1.102 | ton, short (2,000 lb) |
| megagram (Mg) | 0.9842 | ton, long (2,240 lb) |
| metric ton per day | 1.102 | ton per day (ton/d) |
| megagram per day (Mg/d) | 1.102 | ton per day (ton/d) |
| megagram per day per square kilometer [(Mg/d)/km2] | 2.8547 | ton per day per square mile [(ton/d)/mi2] |
| megagram per year (Mg/yr) | 1.102 | ton per year (ton/yr) |
| metric ton per year | 1.102 | ton per year (ton/yr) |
| Pressure |
| kilopascal (kPa) | 0.009869 | atmosphere, standard (atm) |
| kilopascal (kPa) | 0.01 | bar |
| kilopascal (kPa) | 0.2961 | inch of mercury at 60°F (in Hg) |
| kilopascal (kPa) | 0.1450 | pound-force per inch (lbf/in) |
| kilopascal (kPa) | 20.88 | pound per square foot (lb/ft2) |
| kilopascal (kPa) | 0.1450 | pound per square inch (lb/ft2) |
| Density |
| kilogram per cubic meter (kg/m3) | 0.06242 | pound per cubic foot (lb/ft3) |
| gram per cubic centimeter (g/cm3) | 62.4220 | pound per cubic foot (lb/ft3) |
| Energy |
| joule (J) | 0.0000002 | kilowatthour (kWh) |
| Radioactivity |
| becquerel per liter (Bq/L) | 27.027 | picocurie per liter (pCi/L) |
| Specific capacity |
| liter per second per meter [(L/s)/m] | 4.831 | gallon per minute per foot [(gal/min)/ft] |
| Hydraulic conductivity |
| meter per day (m/d) | 3.281 | foot per day (ft/d) |
| Hydraulic gradient |
| meter per kilometer (m/km) | 5.27983 | foot per mile (ft/mi) |
| Transmissivity* |
| meter squared per day (m2/d) | 10.76 | foot squared per day (ft2/d) |
| Application rate |
| kilograms per hectare per year [(kg/ha)/yr] | 0.8921 | pounds per acre per year [(lb/acre)/yr] |
| Leakance |
| meter per day per meter [(m/d)/m] | 1 | foot per day per foot [(ft/d)/ft] |
| millimeter per year per meter [(mm/yr)/m] | 0.012 | inch per year per foot [(in/yr)/ft] |

Temperature in degrees Celsius (°C) may be converted to degrees Fahrenheit (°F) as follows:

°F=(1.8×°C)+32

Temperature in degrees Fahrenheit (°F) may be converted to degrees  Celsius (°C) as follows:

°C=(°F-32)/1.8

Vertical coordinate information is referenced to the insert datum name (and abbreviation) here, for instance, “North American Vertical Datum of 1988 (NAVD 88)”

Horizontal coordinate information is referenced to the insert datum name (and abbreviation) here, for instance, “North American Datum of 1983 (NAD 83)”

Altitude, as used in this report, refers to distance above the vertical datum.

*Transmissivity: The standard unit for transmissivity is cubic foot per day per square foot times foot of aquifer thickness [(ft3/d)/ft2]ft. In this report, the mathematically reduced form, foot squared per day (ft2/d), is used for convenience.

Specific conductance is given in microsiemens per centimeter at 25 degrees Celsius (µS/cm at 25°C).

Concentrations of chemical constituents in water are given either in milligrams per liter (mg/L) or micrograms per liter (µg/L).

NOTE TO USGS USERS: Use of hectare (ha) as an alternative name for square hectometer (hm2) is restricted to the measurement of small land or water areas. Use of liter (L) as a special name for cubic decimeter (dm3) is restricted to the measurement of liquids and gases. No prefix other than milli should be used with liter. Metric ton (t) as a name for megagram (Mg) should be restricted to commercial usage, and no prefixes should be used with it.

An Assessment of the High-gain Streckeisen STS2 Seismometer for Routine Earthquake Monitoring in the United States

By D. E. McNamara, R.P. Buland, and H.M. Benz

# Abstract

Objective. In this document, we report the results of a study to determine if the Streckeisen STS2  high-gain seismometer is appropriate for use by the United States Geological Survey’s (USGS) Advanced National Seismic System (ANSS) for routine earthquake monitoring in the United States (US).

Issue. At issue is whether the high-gain STS2, with a sensitivity of 20,000 volts/meter/sec, can produce onscale recordings of earthquake activity within the continental US, or whether the low-gain STS2, with a sensitivity of 1500 volts/meter/sec, is more appropriate for earthquake monitoring due to the relatively lower input velocity required for clipping the recording system. The test ANSS backbone station configuration, considered in this study, consists of an STS2 broadband seismometer coupled to a Quanterra Q330 digitizer. The Q330 has a channel sensitivity of 4.19e5 counts/volt and a clip level of 8.38e6 = 8388608 counts (2e23) (20 volts). Therefore, an input ground velocity of only 0.001 m/sec will clip the high-gain system while an input ground velocity of 0.013 m/sec, a factor of 13 larger, is required to clip the lower-gain configuration.

Methods. In this report, we use three different methods to examine the levels of input ground velocity expected during routine earthquake monitoring within the continental US. These methods include:

- Computing peak ground velocity (PGV) expected at each ANSS backbone station from the USGS’s 10 percent probability of exceedance in 30 years 1Hz spectral acceleration maps (fig. 1).

- Modeling the amplitude of PGV at regional distances using Brune source scaling (fig. 2).

- Analysis of the September 2004, M6 Parkfield earthquake at four broadband seismic stations within 1000km of the epicenter (figs. 2 and 3).

Recommendation: Based on the analysis of the three methods discussed above, it is our recommendation that the low-gain STS2 is more appropriate for routine earthquake monitoring in the US.  Method 1 indicates that if the high-gain configuration is used at all ANSS backbone stations, all can be expected to clip given the ground-shaking levels expected in the next 30 years. Methods 2 and 3 indicate that an M6 will clip stations within approximately 875 km of the epicenter. The high-gain STS2 is more useful for recording teleseismic events on a global scale but not suitable for onscale recording of moderate to large earthquakes within the continental US.

Methods:

I.  PGV computed from 1Hz Spectral Acceleration Map.

# In this analysis we use the 1 Hz spectral acceleration 10% probability of exceedance in 30 years to calculate the expected PGV at each ANSS backbone station within the continental US (Frankel et al., 2002). The method includes the following procedures:

- For each ANSS station, determine %g from the 1 Hz spectral acceleration map (fig. 1) and convert to g  (g = 0.01* %g) where 1g=980cm/s/s.

- Compute PGV in m/sec after (Newmark and Hall, 1982) where

PGV = 594.8 * g / (2*).

- Compare PGV to clipping velocity of low and high-gain recording systems.

Results from this analysis are summarized in table 1 and in figure 1. All high-gain STS2’s will clip the Q330 digitizer given the input PGV computed from the 10 percent probability of exceedance spectral acceleration in the next 30 years. In contrast, 70 percent of the low-gain STS2’s will clip in the regions with the highest expected input ground velocity. Stations in the midcontinent are not expected to clip (fig. 1). The high level of clipped stations for the low and high gain STS2 seismometers indicate the importance of colocated low gain accelerometers if onscale recording of all large motion is desired.

# II. Brune Source Modeling

In this analysis, we calculate expected shear-wave amplitude using the source scaling models of Brune (1970) and Brune (1971) for a range of distances (25-1200 km) and magnitudes (M5-7). We then compare the modeled amplitude to the velocity required to clip the high and low gain recording systems. (fig. 2). The method includes the following procedures.

1. For each Mw, determine Mo from: Mw = 0.667 log(Mo) – 10.7 	(Kanimori, 1977).

2. For each Mo, determine fault dimension from: Mo = 2.29(r3 dyne-cm

(Brune 1970, 1971)  where  ( = 1.0*108 effective stress.

3. For each fault dimension, determine source corner frequency from:

fault dimension in cm (Brune 1970, 1971).

where	( = shear wave velocity = 3.5km/s

fc = source corner frequency

4. Find angular frequency (fm), by bisection, where Brune source spectrum is a maximum given ANSS short-period detection filters (0.5-12 Hz).

5. For each source-receiver distance (from 25-1100 km), calculate Shear-wave PGV amplitude (As) in m/sec from:

(Brune 1970, 1971).

where	( = rigidity,

fc = source corner frequency,

fm = effective frequency where amplitude is maximum,

( = shear-wave velocity (3.5km/s), and

 = source-receiver distance.

5. Apply attenuation to Shear-wave amplitude (As).

where ( = geometric spreading (0.5)

Q = Quality factor.

Northern California (Erickson et al. 2004).

Q = 105fm 0.67

6.  Multiply the modeled amplitude by the total sensitivity of the high-gain recording system (8.38e9 counts/m/s) and plot corrected PGV as a function of epicentral distance.

Results from this analysis are shown in figure 2 along with the input clipping velocities for the high and low-gain systems.  From figure 2, we see that an earthquake with a magnitude of 6.0 will clip the high-gain system at distances less than about 875 km while the low-gain system will clip at distances less than 150 km. A magnitude 5.0 will clip the high-gain system at distances less than 175 km.

# III.  M6.0 Parkfield Earthquake, September 2004

To verify the amplitude modeling results calculated in Method II we analyzed the amplitudes of the September 2004, M6 Parkfield earthquake at four broadband seismic stations within 1100 km of the epicenter. Four stations that cooperate with the ANSS from the US, BRK and IU networks were used in the analysis. In each case, the M6.0 Parkfield earthquake did not clip any of the broadband stations. The stations used in the analysis include:

HOPS: Hopland, Mendicino County California. BRK network.

Digitizer: Quanterra Q380 recording at 20 samples/second.

seismometer : STS1.

COR: 	Corvallis Oregon, US/IU network.

Digitizer: Quanterra Q380 recording at 20 samples/second.

Seismometer: STS1(sens 2.51e3 counts/volt)

WDC:	Whiskeytown Dam, California, US/BRK network.

Digitizer: Quanterra Q380 recordinga t 20 samples/second.

Seismometer:  Low-gain STS2 (sens 1500 counts/volt).

MOD:	Modoc County California, BRK network.

Digitizer: Quanterra QX-80 recording at 20 samples/second.

Seismometer: STS1 (sens 2240 counts/volt).

The individual station instrument responses (Appendix 1) were deconvolved from the data and converted to displacement. We then convolved the data from each station with total response of the Q330 and high-gain STS2 (Appendix 1). For each station, peak amplitude (PGV) was measured from the largest arrivals on the vertical component of motion (BHZ), which at regional distances is dominated the multiply reflected crustal shear-wave, Lg. PGV is then plotted as a function of epicentral distance (fig. 2) (Table 2).

From figure 2, it is apparent that the three stations closest to the Parkfield earthquake (HOPS, WDC, MOD) would have clipped if the high-gain recording systems were in use. In contrast, the stations did not actually clip with their current instrumentation and would not have clipped if the low-gain STS2 were deployed (fig. 2). Station COR, at 1003 km, would not have clipped either the low or high-gain system. Figure 3 shows the amplitude of the Parkfield M6.0 recorded at MOD at 667 km for three different instrumentation configurations. MOD currently has an STS1 (black line, fig. 3) and did not clip on the Parkfield earthquake. Figure 3 also demonstrates that a high-gain STS2 recording system would have been clipped by the Parkfield earthquake at MOD at a distance of 667 km.

#### Conclusions

Based on the results from this analysis, it is apparent that the high-gain STS2 is not optimal for routine earthquake monitoring in the US. The gain level is sufficiently high that moderate size (M6.0) earthquakes would clip the Quanterra Q330 digitizer at distances less than 875 km. This is unacceptable if we hope to record these events onscale. Instead, the high-gain STS2 is more appropriate for remote seismic stations where weaker ground motion is expected.

# References

Brune, J., 1970, Tectonic stress and the spectra of seismic shear waves from earthquakes, J. Geophys. Res., 75, 4997-5009.

Brune, J., 1971, Correction, J. Geophys. Res., 76, 5002.

Erickson, D., D. McNamara, H. Benz, 2004, Frequency-dependent Lg Q within the Continental United States, Seism. Soc. Am. Bull, 94, 1630-1643.

Frankel, A., Petersen, M., Mueller, C., Haller, K., Wheeler, R., Leyendecker, E., Wesson, R., Harmsen, S., Cramer, C., Perkins, D., and Rukstales, K., 2002, Documentation for the 2002 update of the national seismic hazard maps, U.S. Geological Survey Open-File Report 02-420, 39 pages.

Kanimori, 1977, The energy release in great earthquakes, J. Geophys. Res, 82, 2881-2987.

Newmark, N.M. and Hall, W.J., 1982, "Earthquake Spectra and Design" EERI monograph.

Table 1. Percent Clip for standard and high-gain STS2 from Method I.

Sta  latitude  longitude  %g       g        PGV      SG %clip HG %clip

(m/s)

AAM 42.300999 -83.656998 1.330970 0.013310 0.0125997 0.944757 12.6018

ACSO 40.231998 -82.983002 1.736430 0.017364 0.016438 1.23256 16.4407

AHID 42.764999 -111.099998 13.802000 0.138020 0.130657 9.79702 130.679

ANMO 34.945999 -106.457001 4.496410 0.044964 0.0425654 3.19167 42.5725

BINY 42.199001 -75.986000 1.559670 0.015597 0.0147647 1.10709 14.7671

BLA 37.210999 -80.420998 2.192460 0.021925 0.020755 1.55626 20.7584

BLO 39.172001 -86.522003 2.566570 0.025666 0.0242965 1.82182 24.3006

BMN 40.431000 -117.222000 8.432590 0.084326 0.0798274 5.98567 79.8407

BOZ 45.599998 -111.633003 9.916620 0.099166 0.093876 7.03907 93.8916

BW06 42.778000 -109.556000 4.586710 0.045867 0.0434202 3.25577 43.4275

CBKS 38.813999 -99.737000 1.145350 0.011454 0.0108425 0.812999 10.8443

CBN 38.205002 -77.373001 1.408630 0.014086 0.0133349 0.999882 13.3371

CCM 38.056000 -91.245003 3.163960 0.031640 0.0299517 2.24586 29.9567

CMB 38.035000 -120.385002 9.530080 0.095301 0.0902169 6.7647 90.2318

COR 44.585999 -123.303001 9.505110 0.095051 0.0899805 6.74697 89.9954

DAC 36.277000 -117.594002 16.641701 0.166417 0.157539 11.8127 157.565

DUG 40.195999 -112.816002 5.698430 0.056984 0.0539444 4.04489 53.9533

DWPF 28.110001 -81.432999 0.653747 0.006537 0.00618872 0.464047 6.18975

ELK 40.744999 -115.238998 4.964200 0.049642 0.0469938 3.52372 47.0016

EYMN 47.945999 -91.495003 0.310897 0.003109 0.00294312 0.220683 2.94361

GOGA 33.410999 -83.467003 2.422660 0.024227 0.0229342 1.71967 22.938

HAWA 46.393002 -119.532997 5.147710 0.051477 0.048731 3.65398 48.7391

HKT 29.950001 -95.833000 0.816093 0.008161 0.00772557 0.579284 7.72685

HLID 43.561001 -114.418999 5.659060 0.056591 0.0535717 4.01695 53.5806

HOPS 38.993999 -123.071999 36.925400 0.369254 0.349556 26.2106 349.614

HRV 42.506001 -71.557999 1.748700 0.017487 0.0165541 1.24127 16.5569

HWUT 41.606998 -111.565002 9.699410 0.096994 0.0918198 6.88489 91.835

ISCO 39.799999 -105.612999 1.654140 0.016541 0.015659 1.17415 15.6616

JCT 30.479000 -99.802002 0.503462 0.005035 0.00476604 0.35737 4.76683

JFWS 42.914001 -90.248001 1.112710 0.011127 0.0105335 0.78983 10.5353

KNB 37.016998 -112.821999 4.696350 0.046964 0.0444582 3.33359 44.4655

Sta  latitude  longitude  %g       g        PGV      SG %clip HG %clip

LBNH 44.240002 -71.926003 2.280530 0.022805 0.0215887 1.61878 21.5923

LKWY 44.564999 -110.400002 12.156100 0.121561 0.115076 8.62871 115.095

LRAL 33.035000 -86.998001 2.224120 0.022241 0.0210547 1.57874 21.0582

LSCT 41.678001 -73.223999 1.604300 0.016043 0.0151872 1.13877 15.1897

LTX 29.334000 -103.667000 2.222410 0.022224 0.0210385 1.57752 21.042

MCWV 39.658001 -79.846001 1.584410 0.015844 0.0149989 1.12465 15.0014

MIAR 34.546001 -93.572998 2.096850 0.020968 0.0198499 1.4884 19.8532

MNV 38.432999 -118.153000 17.172800 0.171728 0.162567 12.1897 162.594

MOD 41.903000 -120.306000 8.412420 0.084124 0.0796365 5.97135 79.6497

MVU 38.505001 -112.210999 7.899620 0.078996 0.074782 5.60736 74.7944

MYNC 35.074001 -84.127998 3.033970 0.030340 0.0287212 2.15359 28.726

NCB 43.971001 -74.223999 2.356310 0.023563 0.0223061 1.67257 22.3098

NEW 48.263000 -117.120003 3.619540 0.036195 0.0342645 2.56924 34.2702

NHSC 33.106998 -80.178001 3.245180 0.032452 0.0307206 2.30351 30.7257

OCWA 47.749001 -124.178001 16.503599 0.165036 0.156232 11.7147 156.258

OXF 34.512001 -89.408997 3.180920 0.031809 0.0301123 2.2579 30.1173

PAS 34.147999 -118.171997 37.812199 0.378122 0.357951 26.8401 358.01

PFO 33.609001 -116.455002 33.764099 0.337641 0.319629 23.9666 319.682

PLAL 34.981998 -88.075996 3.189070 0.031891 0.0301894 2.26368 30.1945

RSSD 44.119999 -104.036003 1.247570 0.012476 0.0118102 0.885558 11.8121

SAO 36.764999 -121.445000 43.401501 0.434015 0.410862 30.8075 410.93

SDCO 37.745998 -105.500999 2.562140 0.025621 0.0242546 1.81867 24.2586

SLM 38.636002 -90.236000 3.590330 0.035903 0.033988 2.54851 33.9936

SSPA 40.636002 -77.888000 1.393160 0.013932 0.0131884 0.988901 13.1906

TPH 38.075001 -117.223000 9.866430 0.098664 0.0934009 7.00345 93.4164

TPNV 36.948002 -116.249001 8.793650 0.087936 0.0832454 6.24196 83.2592

TUC 32.310001 -110.783997 2.316950 0.023170 0.0219335 1.64463 21.9371

WCI 38.229000 -86.293999 2.863830 0.028638 0.0271106 2.03282 27.115

WDC 40.580002 -122.540001 12.858200 0.128582 0.121723 9.12708 121.743

WMOK 34.737999 -98.780998 1.342820 0.013428 0.0127119 0.953168 12.714

WUAZ 35.516998 -111.374001 3.457260 0.034573 0.0327283 2.45405 32.7337

WVOR 42.433998 -118.637001 4.970230 0.049702 0.0470509 3.528 47.0587

WVT 36.130001 -87.830002 3.815320 0.038153 0.0361179 2.70821 36.1239

EGMT 47.950001 -109.779999 1.126660 0.011267 0.0106656 0.799733 10.6673

DGMT 48.580002 -104.199997 0.691208 0.006912 0.00654335 0.490638 6.54444

Sta  latitude  longitude  %g       g        PGV      SG %clip HG %clip

VKMN 48.220001 -96.410004 0.356017 0.003560 0.00337025 0.25271 3.37081

COWI 46.060001 -89.260002 0.589685 0.005897 0.00558227 0.418573 5.5832

GRMI 44.660000 -84.720001 0.939180 0.009392 0.00889078 0.666654 8.89226

OGNE 41.130001 -101.720001 1.013730 0.010137 0.00959651 0.719572 9.5981

SCIA 42.020000 -93.160004 1.104300 0.011043 0.0104539 0.783861 10.4556

PDNY 44.669998 -74.980003 2.698170 0.026982 0.0255423 1.91523 25.5466

MIME 45.660000 -68.709999 2.144410 0.021444 0.0203001 1.52216 20.3035

MONC 35.200001 -78.070000 1.583770 0.015838 0.0149928 1.1242 14.9953

ATAL 31.020000 -87.489998 1.479430 0.014794 0.0140051 1.05014 14.0074

GHTX 27.170000 -98.120003 0.288432 0.002884 0.00273045 0.204736 2.7309

SNSD 45.060001 -99.510002 0.750744 0.007507 0.00710694 0.532897 7.10812

ANND 47.880001 -100.239998 0.509268 0.005093 0.004821 0.361491 4.8218

AMTX 35.180000 -101.870003 0.955765 0.009558 0.00904778 0.678426 9.04928

BNMN 46.360001 -94.199997 0.442334 0.004423 0.00418737 0.31398 4.18806

BRMT 45.299999 -108.910004 3.247320 0.032473 0.0307409 2.30503 30.746

KCCO 38.759998 -102.790001 1.054920 0.010549 0.00998644 0.74881 9.9881

ERPA 42.130001 -80.089996 1.453010 0.014530 0.013755 1.03138 13.7573

SEKY 36.599998 -83.720001 2.942140 0.029421 0.0278519 2.08841 27.8565

CLKS 37.029999 -97.610001 1.406600 0.014066 0.0133156 0.998441 13.3178

NNNM 36.889999 -109.690002 1.343560 0.013436 0.0127189 0.953694 12.721

MNTX 31.697001 -105.382004 2.613080 0.026131 0.0247368 1.85483 24.7409

Table 2. Parkfield PGV amplitude convolved with high-gain system response.

Sta		latitude	longitude	distance 	PGV

(km)		(m/s)

MOD		41.9033	-120.305	667 		0.00183

HOPS		38.9939	-123.071	421 		0.00305

WDC		40.5800	-122.539	558 		0.00190

COR		44.5857	-123.303	1003		0.00080

Appendix 1: Instrument responses

High-gain STS2 and Q330:

ZEROS   3

0.  0.

0.  0.

0.  0.

POLES   5

-3.70237E-02    3.70244E-02

-3.70237E-02   -3.70244E-02

-251.327  0.

-118.634    423.065

-118.634   -423.065

CONSTANT     4.05950E+17

HOPS:

ZEROS   3

0.  0.

0.  0.

0.  0.

POLES   4

-1.23400E-02    1.23400E-02

-1.23400E-02   -1.23400E-02

-39.1800    49.1200

-39.1800   -49.1200

CONSTANT     3.88355E+12

WDC:

ZEROS   3

0.  0.

0.  0.

0.  0.

POLES   5

-3.70237E-02    3.70244E-02

-3.70237E-02   -3.70244E-02

-251.327  0.

-118.634    423.065

-118.634   -423.065

CONSTANT     2.92868E+16

MOD:

ZEROS   3

0.  0.

0.  0.

0.  0.

POLES   4

-1.23400E-02    1.23400E-02

-1.23400E-02   -1.23400E-02

-19.5900    24.5600

-19.5900   -24.5600

CONSTANT     9.85368E+11

COR:

ZEROS   3

0.  0.

0.  0.

0.  0.

POLES   4

-1.23400E-02    1.23400E-02

-1.23400E-02   -1.23400E-02

-39.1800    49.1200

-39.1800   -49.1200

CONSTANT     4.18246E+12

Figure 1: 10 percent in 30 years 1 Hz spectral acceleration for the United States.

- Assuming Low-gain STS2 instrumentation (red triangles = clipped stations).

- Assuming high-gain STS2 instrumentation (red triangles = clipped stations).

Figure 2: Brune source modeled PGV as a function of epicentral distance. Also plotted is the PGV measured at four stations from the September 2004 Parkfield M6.0 earthquake.

Figure 3: Seismogram of the Parkfield M6.0 recorded at MOD. Black line is the original seismogram using an STS1 seismometer. Magenta line is the waveform convolved with the low-gain recording system. Green line is the waveform convolved with the high-gain system.

# Footnotes

USGS ANSS, Golden, CO